Copyright Information of the Article Published Online

| TITLE | Capsule endoscopy: the road ahead |
| --- | --- |
| AUTHOR(s) | Ana-Maria Singeap, Carol Stanciu, Anca Trifan |
| CITATION | Singeap AM, Stanciu C, Trifan A. Capsule endoscopy: the road ahead. World J Gastroenterol 2016; 22(1): 369-378 |
| URL | http://www.wjgnet.com/1007-9327/full/v22/i1/369.htm |
| DOI | http://dx.doi.org/10.3748/wjg.v22.i1.369 |
| OPEN-ACCESS | Open-Access: This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/ |
| CORE TIP | Since its introduction into clinical practice 15 years ago, small bowel capsule endoscopy (CE) has revolutionized direct endoscopic imaging of the gut. During this time, CE has gained tremendous popularity among gastroenterologists, and a vast research pertaining to its different aspects has been published. Dedicated esophageal and colon CE have expanded the field of application to upper and lower gastrointestinal disorders. However, besides its recognized advantages, CE also has several limitations such as the lack of therapeutic capabilities, the inability to obtain biopsies and control its locomotion. Active research is in progress to overcome the current limitations, while the latest advances in CE technology enable us to look forward to a next generation CE capable of performing both diagnostic and therapeutic procedures. This review summarizes the literature that prognosticates about the future of CE. |
| KEY WORDS | Capsule endoscopy; Biopsy; Drug delivery systems; Capsule endoscope locomotion; Capsule localization |
| COPYRIGHT | © The Author(s) 2016. Published by Baishideng Publishing Group Inc. All rights reserved. |
| NAME OF JOURNAL | World Journal of Gastroenterology |
| ISSN | 1007-9327 (print) and ISSN 2219-2840 (online) |
| PUBLISHER | Baishideng Publishing Group Inc, 8226 Regency Drive, Pleasanton, CA 94588, USA |
| WEBSITE | http://www.wjgnet.com |

REVIEW

Capsule endoscopy: the road ahead

Ana-Maria Singeap, Carol Stanciu, Anca Trifan

Ana-Maria Singeap, Anca Trifan, Gastroenterology and Hepatology Institute, “Gr. T. Popa” University of Medicine and Pharmacy, “St. Spiridon” Emergency Hospital, 700111 Iasi, Romania

Carol Stanciu, Gastroenterology and Hepatology Institute, “St. Spiridon” Emergency Hospital, 700111 Iasi, Romania

Author contributions: Singeap AM, Stanciu C and Trifan A contributed equally to the conception and design of the review; Singeap AM performed acquisition of data; Stanciu C analyzed the data, coordinated the manuscript drafting and revised it critically; Trifan A participated in the analysis and interpretation of data, and critically revised the manuscript for important intellectual content; all authors read and approved the final version of the manuscript.

Correspondence to: Carol Stanciu, MD, FRCP, Professor, Gastroenterology and Hepatology Institute, “St. Spiridon” Emergency Hospital, Independentei 1, 700111 Iasi, Romania. stanciucarol@yahoo.com

Telephone: +40-732-402860   Fax: +40-232-246611

Received: May 5, 2015   Revised: August 4, 2015  Accepted: September 30, 2015

Published online: January 7, 2016

Abstract

Since its introduction into clinical practice 15 years ago, capsule endoscopy (CE) has become the first-line investigation procedure in some small bowel pathologies, and more recently, dedicated esophageal and colon CE have expanded the fields of application to include the upper and lower gastrointestinal disorders. During this time, CE has become increasingly popular among gastroenterologists, with more than 2 million capsule examinations performed worldwide, and nearly 3000 PubMed-listed studies on its different aspects published. This huge interest in CE may be explained by its non-invasive nature, patient comfort, safety, and access to anatomical regions unattainable via conventional endoscopy. However, CE has several limitations which impede its wider clinical applications, including the lack of therapeutic capabilities, inability to obtain biopsies and control its locomotion. Several research groups are currently working to overcome these limitations, while novel devices able to control capsule movement, obtain high quality images, insufflate the gut lumen, perform chromoendoscopy, biopsy of suspect lesions, or even deliver targeted drugs directly to specific sites are under development. Overlooking current limitations, especially as some of them have already been successfully surmounted, and based on the tremendous progress in technology, it is expected that, by the end of next 15 years, CE able to perform both diagnostic and therapeutic procedures will remain the major form of digestive endoscopy. This review summarizes the literature that prognosticates about the future developments of CE.

Key words: Capsule endoscopy; Biopsy; Drug delivery systems; Capsule endoscope locomotion; Capsule localization

© The Author(s) 2016. Published by Baishideng Publishing Group Inc. All rights reserved.

Singeap AM, Stanciu C, Trifan A. Capsule endoscopy: the road ahead. World J Gastroenterol 2016; 22(1): 369-378  Available from: URL: http://www.wjgnet.com/1007-9327/full/v22/i1/369.htm  DOI: http://dx.doi.org/10.3748/wjg.v22.i1.369

Core tip: Since its introduction into clinical practice 15 years ago, small bowel capsule endoscopy (CE) has revolutionized direct endoscopic imaging of the gut. During this time, CE has gained tremendous popularity among gastroenterologists, and a vast research pertaining to its different aspects has been published. Dedicated esophageal and colon CE have expanded the field of application to upper and lower gastrointestinal disorders. However, besides its recognized advantages, CE also has several limitations such as the lack of therapeutic capabilities, the inability to obtain biopsies and control its locomotion. Active research is in progress to overcome the current limitations, while the latest advances in CE technology enable us to look forward to a next generation CE capable of performing both diagnostic and therapeutic procedures. This review summarizes the literature that prognosticates about the future of CE.

INTRODUCTION

Fifteen years have passed since small bowel capsule endoscopy (CE) was launched[1], revolutionizing noninvasive direct visualization of the small bowel, considered until then the “black box” of the gastrointestinal (GI) tract. During this time, CE has been used extensively, with more than 2 million capsules swallowed worldwide[2] and nearly 3000 PubMed-listed studies pertaining to its different aspects published[3]. Technical progress led to the introduction of some updated versions (2nd and 3rd generations) of CE for the small bowel and the manufacturing of the CE designed for esophagus and colon. In just a few years, CE has evolved very rapidly, becoming an invaluable tool for examination of almost the entire GI tract, and its diagnostic achievements have by far exceeded early expectations. Still, CE is not an ideal tool, as it has several limitations, including the lack of therapeutic capabilities, inability to control its locomotion and thus, to revisualize critical areas and obtain biopsies. The objective of many research groups worldwide is to overcome these limitations and develop a new generation of CE with higher diagnostic yield and therapeutic capabilities. Of course, it is very difficult to predict the future in medicine, and would be for CE. However, based on the extraordinary developments seen in just 15 years since its emergence, and the tremendous progresses of modern technology, it can be anticipated that, by the end of next 15 years, the new generation of CE able to perform both diagnostic and therapeutic procedures in a noninvasive, painless, and elegant manner will remain the major form of digestive endoscopy, covering the entire GI tract from mouth to anus, as its inventors have dreamed. This review summarizes available literature that prognosticates about the future developments in CE.

BRIEF LOOK BACK AND THE CURRENT STATUS

The first model of CE called M2A (meaning “mouth to anus”) was launched in 2000 by Given Diagnostic Imaging, Yoqneam, Israel[4], and the merits for its design belonged, in a similar degree, to the Israeli engineer Gavriel Iddan and the British gastroenterologist Paul Swain[1,4]. A year later, M2A was approved for clinical use in Europe and the United States, and after the advent of esophageal CE, M2A changed its name into PillCam SB (meaning “small bowel”). Several other companies have also developed small bowel endoscopic capsules: EndoCapsule (Olympus Corp., Tokyo, Japan)[5], OMOM capsule (Jinshan Science and Technology Company, Chongqing, China)[6], Mirocam (IntroMedic Co., Seoul, South Korea)[7], and CapsoCam SV1 (CapsoVision, Saratoga, CA, United States)[8], all having many similar characteristics and diagnostic performances to PillCam SB, but differing with regard to image acquisition rate, field of view, battery life, dimensions, and technology for transmission of images. Given Imaging has also developed PillCam ESO and PillCam COLON for the evaluation of esophageal and colonic diseases, respectively[9,10]. Improvements in technology have led to the development of 2nd and 3rd generation CEs which overcome some limitations of the 1st generation CE by increasing the view angle, extending the effective battery life, and including several others systems which offer superior image quality, tissue coverage, and interpretation efficiency[11-13].

In only 15 years since the introduction of CE into clinical practice, its achievements have exceeded what was previously thought as possible. Thus, CE has revolutionized the evaluation of obscure gastrointestinal bleeding (OGIB) and unexplained iron deficiency anemia (IDA)[13-15], becoming the first-line modality in the diagnosis of both. The role of CE in OGIB/IDA is supported mainly by its diagnostic performance, which is superior to other diagnostic modalities (push enteroscopy, intraoperative enteroscopy, small bowel barium radiography, CT-enterography, CT-angiography, MR-enterography), as well as by its positive impact on patient management and outcome[14,16-21]. When CE was compared to double-balloon enteroscopy, a similar diagnostic accuracy for OGIB was reported[22]. CE examination leads to therapeutic endoscopic or surgical interventions and, consequently, to bleeding being stopped and outcomes improved[23,24].

Thanks to its capacity to directly visualize mucosa of the entire small bowel, CE has undoubtedly contributed substantially to progress in diagnosis, therapeutic decision, and outcome in Crohn’s disease (CD). Reviews of existent literature on CE diagnostic yield, for both suspected and known small bowel CD, show it to be superior to other diagnostic techniques such as small bowel follow-through, enteroclysis, push-enteroscopy, ileo-colonoscopy, and CT-enterography[25-27]. CE is superior to MR-enterography in identifying small bowel mucosal lesions, while MR-enterography is superior to CE in diagnosing mural and extra-enteric lesions[28]. In patients with known CD, an important treatment goal is mucosal healing which can be reliably assessed by CE[29-31].

CE has an 8-fold magnification capacity and a minimum size of lesion detection of 0.1-0.2 mm, so that villi can be easily observed during a procedure; therefore, it may be a useful noninvasive diagnostic tool in patients with suspected or established celiac disease[32,33]. However, CE is actually an alternative to endoscopy with biopsy only in patients clinically suspected of celiac disease unable or unwilling to undergo conventional endoscopy.

CE has become the procedure of choice for detecting small bowel polyps in hereditary polyposis syndromes like Peutz-Jegher syndrome and familial adenomatous polyposis[34,35]. In addition, widespread use of CE has more than doubled the diagnosis rate in small bowel tumors[36-41].

Esophageal capsule endoscopy, although at 3rd generation, has limited role in clinical practice and it is still under evaluation[42]. Colon capsule is also under evaluation, and is currently recommended in case of incomplete colonoscopy and in patients unwilling or unable to perform colonoscopy[43,44].

Limitations of current capsule endoscopy

Although CE has seen tremendous advances in a short period of time since its introduction in clinical practice, it has several limitations. Thus, CE remains a purely visual technique with no ability to obtain biopsy specimens or perform therapeutic maneuvers. The most obvious drawback is the operator’s inability to control its locomotion through GI tract. The capsules presently on the market are unable to localize or mark the location of detected lesions. Visualization may be impaired by the presence of food materials or bubbles and, in contrast with conventional endoscopy, CE cannot perform flushing, suctioning, or air insufflation to obtain better images. All capsules for clinical use are powered by limited-life batteries which may be depleted before the examination is complete. The rate of missed lesions is still high for those located in the duodenum and proximal jejunum, where the transit is more rapid than in the distal segment of the small bowel. Reading time for interpretation is another shortcoming of CE, as it takes more than 1 h to read a full 8-h examination. Finally, the costs are still high.

FUTURE EXPECTATIONS IN CAPSULE ENDOSCOPY

The future of CE is difficult to predict (“Prediction is very difficult, especially about the future” - Niels Bohr, Nobel Prize winner, 1885-1962), although novel technologies may lead to developments which today seem almost unimaginable. Improvements achieved in just 15 years since the introduction of CE in clinical practice go beyond what was previously thought as possible. GI endoscopy has had a similar history: initially limited only to viewing the esophagus/gastric lumen, it has improved progressively over a few decades, developing into an accurate diagnostic and therapeutic technique. CE also started as a tool for visualizing only the “black box” (small bowel) which has long been the final frontier of the GI endoscopy, and it evolved very rapidly to become a non-invasive endoscopic tool in the examination of almost the entire GI tract.

Most likely, over the next 15 years, CE will slowly replace diagnostic standard endoscopy and take over most therapeutic procedures with no pain and no need for sedation. We know that several research groups throughout the world are working to develop new multifunctional capsules with diagnostic and therapeutic capabilities extending far beyond our imagination. What we do not yet know is whether the future CE will be “universal”, containing both diagnostic and therapeutic modules (an “ideal” CE)[45] or “specific”, for diagnosis or therapy[11].

Maneuverable capsules

In contrast to standard endoscopy, the movement of the current capsule endoscopes through the GI tract is passive, ensured by peristaltic motion, the operator being unable to control the endoscopic navigation (right and left, back and forth) in a given area. It is of the upmost importance to solving the CE’s maneuvering limitation so as to increase its diagnostic yield and allow targeted biopsy and even drug delivery. Besides enhancing diagnostic yield, a capsule whose locomotion can be controlled will reduce the amount of energy consumed, examination time, as well as the rate of capsule retention. Even more, an active control of the endoscopic capsule would allow us to examine the stomach, and finally, the entire GI tract[46].

Systems that can be used to propel or steer the capsule are under development. There are two locomotion systems: an internal one, integrated on-board the capsule, and an external one (outside the capsule), most frequently based on magnetic fields. Some proposed internal systems consist of legged-like mechanisms (propellers/paddles) that can be deployed by the capsule to resist peristaltic movements, while the external locomotion systems usually use a capsule covered with a magnetic shield which can interact with external magnetic fields created by an electromagnet or permanent magnet. Electromagnets require bulkier equipment by comparison to permanent magnets[47-50].

The legged-like device approach consists of providing the capsule with propellers/paddles which will start functioning on demand during capsule navigation through various segments of the GI tract. A four-legged capsule, two in the front and two in the rear, has been proposed, an eight-legged capsule was also suggested to be feasible, and even a twelve-legged locomotion capsule was designed to improve propulsion and reduce tissue injury[47,51-53]. However, several technical drawbacks such as insufficient space available within the capsule and high power consumption should be overcome. In addition, a failure in the synchronization of the legs may cause damage to the GI tissue.

Magnetic control appears to be the most attractive and promising approach. It is based on the principle that a large external magnetic field created by a permanent magnet or electromagnet near the patient interacts with a small internal magnet component integrated into the capsule to provide an active control of the endoscopic capsule[48]. Given Imaging has incorporated a magnet inside one of the domes of a standard PillCam colon capsule, which can be manipulated with an external handheld magnet moved on the patient’s abdomen[54,55]. Using such magnetically maneuverable capsule, one study reported > 75% of gastric mucosa visualized and no adverse events[55]. Siemens (Siemens Medical, Erlangen, Germany) and Olympus (Olympus America, Center Valley, PA, United States) have recently tested the prototype of a magnetically guided capsule endoscope that uses a three-dimensional, external magnetic field which interacts with the magnet inside the capsule, allowing the capsule to be moved forward or backward[56,57]. Rey et al[58,59] made the first blinded comparative clinical trial on gastric examination in humans, comparing a magnetically guided capsule endoscope with a conventional high-definition gastroscope, and found a similar diagnostic yield for both methods. Rahman et al[60], using the Intromedic MiroCam-Navi system, reported a high degree of visualization, control, and maneuverability with this system. A robotic magnetic navigation system used in cardiology (Niobe, Stereotaxis Inc., St. Louis, MO, United States) has been suggested for CE but has been tested only in plastic phantoms[61]. Several other versions of endoscopic capsules magnetically propelled by a robotic arm have been proposed[62].

Two research projects funded by the European Union aim to develop a self-propelling minirobot pill. One is VECTOR (Versatile Endoscopic Capsule for gastrointestinal Tumor Recognition and Therapy) for early diagnosis and treatment of GI cancer[63], and the other is NEMO (Nano-based capsule Endoscopy with Molecular imaging and Optical biopsy) which designed to combine optical, nano, and maneuvering technologies in a new capsule with different diagnostic and therapeutic capabilities[64].

A videocapsule endoscope called Compact Photonic Explorer (CPE), measuring 5 mm in size, has been developed at the City University and City College of New York. It can be manipulated externally by remote controlled radio signal and may be used in the future for diagnostic and therapeutic means[65]. Recently, a mathematical model of an electrically propelled capsule endoscopy has been proposed, using double pairs of electrodes, and which is able to move the capsule forward and backward at a speed of 2.91 mm/s and 2.23 mm/s, respectively[66].

To summarize, the development of propelled/steerable capsules will represent a major advance of capsule technology, which will open a myriad of possibilities, including a more detailed evaluation of affected areas and prelevation of biopsy specimens, endoscopic targeted therapy, examination of the stomach, thus the entire GI tract becoming virtually as accessible as the skin[67]!

Biopsy

Once a maneuverable capsule is developed, the next step is to obtain a tissue sample. Several biopsy devices have been developed and used on animal models. A spring-loaded device similar to the Crosby capsule, guided by real-time imaging and RF-controlled remote manipulation, and a capsule using Micro-Electro-Mechanical-Systems (MEMS) technology have been successfully tested[68]. Both NEMO and VECTOR projects develop capsules designed for virtual biopsies and drug delivery[63,64]. The rotational Micro Biopsy Capsule Device (Seoul, South Korea) which contains a triggering part with paraffin block and a rotational tissue-cutting razor (biopsy part) has been tested[69]. A tethered capsule endomicroscopy of the esophagus, which uses optical frequency domain imaging technology and enables 3D imaging of esophagus in microscopic detail, has also been developed[70]. This capsule endomicroscope is able to differentiate Barrett’s esophagus from normal esophageal mucosa. Other magnetic capsules using untethered microgrippers to grab tissue samples or magnetic torsion spring mechanism have been designed[71,72].

Optical enhancing techniques could lead to optical “biopsy”, which refers to a method of obtaining a morphological diagnosis without biopsy specimens, and prototype endoscopic capsules with such technology have been developed, including the wireless spectroscopic compact photonic explorer for diagnostic optical imaging to detect microscopic malignancy[65]. One research group integrated near-infrared fluorescent probe in CE to enhance optical diagnosis of neoplasia, which proved able to distinguish adenomatous tissue in experimental colitis in mice[73].

Power source

At present, available endoscopic capsules use two coin-shaped, silver-oxide batteries that can generate 20 mW of energy, far too little to accomplish the multiple diagnostic and therapeutic tasks of the future capsule, most of them requiring power consumption. In addition, batteries occupy most of the space in an endoscopic capsule. Therefore, increased power supply and reduced size of batteries, to leave enough space to incorporate diagnostic/therapeutic components into the capsule, are essential for further developments in CE technology. A solution may be lithium ion microbattery technology which could provide a power density up to 2000-times higher than other microbatteries[74]. Recently, Rathore et al[75] using Ultrascale FinFET 16 nm technology for manufacturing endoscopic capsules (instead of 18 µm used for conventional endoscopic capsules) have reported an increased battery life, reduced power consumption with up to 50%, and a reduced size of the capsule by 12% compared to traditional capsules. An alternative method to reduce battery consumption is to use low complexity video compression technology that saves radiofrequency (RF) transmitting power[76].

External rechargeable batteries (from an extracorporeal power supply) using RF, microwave or electric induction, and even “battery free” CE using wireless power transmission (WPT) technology are created. An excellent overview of the development of emergency WPT technique for application in CE has been recently published by Basar et al[77]. WPT system employs a transmitting coil positioned outside the human body and a receiving coil installed within the CE, thereby eliminating the need for an internal battery[78]. Thus, the RF System Lab (Nagano, Japan) was the first to use WPT technology in their Sayaka and Norika capsules[79], and several publications centered on WPT technology for the endoscopic capsule[77,80]. Jia et al[80], using WPT technology, have reported on its ability to transmit 500 mW of electricity, which is significantly higher than the amount generated by current batteries used for the endoscopic capsules available on the market.

An alternative solution will be the development of three-dimensional microbattery technology for geometrical energy and power density of battery[81,82], and many research groups are working in the field, still progressing in several laboratories.

Targeted drug delivery

Unfortunately, none of the current capsules is able to perform therapy. New capsule devices are under development in order to enable drug delivery in specific diseased areas of the GI tract. A number of clinical situations can benefit from targeted drug delivery such as the use of hemostatic spray to an active bleeding lesion or localized application of steroid/immunomodulation for CD. One capsule prototype is able to deliver an injection of 1 mL of targeted medication while using a holding mechanism[83]. To achieve this, an accurate control mechanism of capsule positioning and a drug release mechanism should be incorporated into a capsule endoscope. As future capsules will most likely be smaller, space limitation within the capsule is an important impediment when incorporating such mechanisms[84].

Philips company (Philips Research, Eindhoven, The Netherlands) has launched an “intelligent” pill (iPill) measuring 11 mm × 26 mm and incorporating a microprocessor, battery, pH sensor, temperature sensor, radiofrequency transceiver, fluid pump, and a drug reservoir[85]. Tracking of the iPill in the GI tract is based on information regarding pH change and gut transit time. Once tracked in the aimed area, the iPill will open and deliver the drug under the control of the microprocessor. The iPill is under trial in CD and colorectal cancer[68,85,86].

Several other wireless capsules, such as the Gastrotarget telemetric capsule (Gastrotarget, Tonawanda, NY, United States), High-frequency capsule (Battelle-Institute V, Frankfurt, Germany), Telemetric capsule (INSERM UG1, Strasbourg, Cedex, France), Enterion capsule (Pheaton Research, Nottingham, United Kingdom), and the IntelliSite capsule (Innovative Devices, Raleigh, NC, United States), have been developed for targeted drug delivery in specific areas of the GI tract[68]. However, capsule tracking is inaccurate due to lack of an anchoring mechanism and thus, drug release cannot be fully controlled. Two therapeutic capsule endoscopes have recently been proposed for bioadhesive patch release and targeted drug delivery, respectively, both capsules being controlled by an external permanent magnetic source[83,87]. A soft magnetically actuated capsule, capable of multimodal gradual or sudden drug release, has also been developed[88].

Even with a new CE designed for targeted drug delivery, several other problems should be taken into consideration. Thus, in some diseases of GI tract such as CD, drug delivery is required on a daily basis, for several days or even weeks. To overcome this problem, a pre-programmed non-viewing capsule for targeted drug delivery has been proposed[89].

Luminal insufflation

CE visualization, especially of the colonic mucosa, is limited as the capsule is unable to provide insufflation in order to distend the intestine and expose all mucosal surfaces for examination. This shortcoming is a potential cause for CE’s high false-negative diagnostic rate in the colon. Experiments and insufflation capsule prototypes show the feasibility of generating large volumes of gas from a small volume of liquid hydrogen peroxide, weak acids and bases in a capsule to provide wireless insufflation for enhancing visualization[90,91]. Recently, a method of controlled colonic insufflation (CO2) via an untethered capsule in vivo has been reported[67].

Shorter reading time

Future CE should allow shorter reading time for interpretation of images acquired by capsule, and this may be achieved by development of more efficient software[92]. A computer-aided lesion detection will significantly reduce reading time.

home procedure

In the near future, CE (small bowel capsule endoscopy and especially colon capsule endoscopy) will become a home procedure that could be done on weekends, thus avoiding work absence[93,94].

Accurate location of detected lesion

A tagging module consisting of a micro tag, compressed spring and thermal ignitor can be integrated within future CEs; when activated by an external signal, the micro tag is impaled into the mucosa to mark the precise location of a lesion for the following endoscopic therapy[95]. Location of lesion and estimation of its size is possible by using Rapid 6 system of software developed by Given Imaging[96].

Automated capsule localization

Automated capsule localization with a software using color image analysis to discriminate between different segments of GI tract (esophagus, stomach, small bowel, colon) identified CE passage across the pylorus in 93% of cases[97-99]. The next step will be development of the software program to increase the frame rate while CE is traversing the duodenum, in order to improve identification of the ampulla of Vater and detect more lesions in the periampullary region[100,101] which is poorly visualized by CE, CT- and MR-enterography[102,103].

Entire GI tract visualization

An ideal CE would be able to visualize the entire GI tract, from mouth to anus, during a single procedure. Currently available capsules cannot be used for this purpose because of the significant physiological differences of the various segments of the GI tract, and therefore, only specific esophageal, small bowel, and colon capsules are available. However, the colon capsule (PillCam COLON 2, Given Imaging) developed for evaluation of the colon, can also be used to visualize almost the entire GI tract. This capsule is provided with two cameras able to record video images from both ends, with an adaptive frame acquisition rate (between 4 and 35 frames per second). Thus, it may visualize the esophagus, examine the stomach and duodenum with an external maneuvering system to control capsule locomotion, then the small intestine and, finally, the colon. Preliminary studies have already concluded that GI tract evaluation with PillCam COLON 2 is feasible, especially for small bowel, although other segments (esophagus, stomach) need technical improvements to obtain a good visualization[46]. In the near future, a pan-endoscopy with CE may be a reality[29,104,105].

CONCLUSION

Undoubtedly, CE has opened a new era in endoscopic diagnosis for gastroenterologists and has set a milestone in the evolution of endoscopic examination of the GI tract without discomfort or need for sedation, or the risks implied by conventional endoscopy. During a relatively short period of time (15 years), CE has proven its high diagnostic yield in multiple pathologies of the GI tract such as obscure GI bleeding, CD, celiac disease, as well as in small and large bowel tumors. Nevertheless, the endoscopic capsules currently available are diagnostic tools only, and still have several limitations (passive locomotion, inability to perform biopsy or deliver therapy, etc). Modern technology continues to make tremendous progress in CE, helping it overcome the above mentioned limitations. Although it is difficult to make predictions about the future, we believe that in the next 15 years, our dreams of an efficient diagnostic and therapeutic CE for the diverse pathologies of the entire GI tract will become a reality.

REFERENCES

1	Iddan G, Meron G, Glukhovsky A, Swain P. Wireless capsule endoscopy. Nature 2000; 405: 417 [PMID: 10839527 DOI: 10.1038/35013140]

2	Eliakim R. Video capsule endoscopy of the small bowel. Curr Opin Gastroenterol 2013; 29: 133-139 [PMID: 23221650 DOI: 10.1097/MOG.0b013e32835bdc03]

3	Koulaouzidis A, Rondonotti E, Karargyris A. Small-bowel capsule endoscopy: a ten-point contemporary review. World J Gastroenterol 2013; 19: 3726-3746 [PMID: 23840112 DOI: 10.3748/wjg.v19.i24.3726]

4	Iddan GJ, Swain CP. History and development of capsule endoscopy. Gastrointest Endosc Clin N Am 2004; 14: 1-9 [PMID: 15062374 DOI: 10.1016/j.giec.2003.10.022]

5	Rey JF, Kuznetsov K, Vazquez-Ballesteros E. Olympus capsule endoscope for small and large bowel exploration. Gastrointest Endosc 2006; 63: AB176 [DOI: 10.1016/j.gie.2006.03.381]

6	Li CY, Zhang BL, Chen CX, Li YM. OMOM capsule endoscopy in diagnosis of small bowel disease. J Zhejiang Univ Sci B 2008; 9: 857-862 [PMID: 18988304 DOI: 10.1631/jzus.B0820034]

7	Bang S, Park JY, Jeong S, Kim YH, Shim HB, Kim TS, Lee DH, Song SY. First clinical trial of the “MiRo” capsule endoscope by using a novel transmission technology: electric-field propagation. Gastrointest Endosc 2009; 69: 253-259 [PMID: 18640676 DOI: 10.1016/j.gie.2008.04.033]

8	Friedrich K, Gehrke S, Stremmel W, Sieg A. First clinical trial of a newly developed capsule endoscope with panoramic side view for small bowel: a pilot study. J Gastroenterol Hepatol 2013; 28: 1496-1501 [PMID: 23701674]

9	Eliakim R, Yassin K, Shlomi I, Suissa A, Eisen GM. A novel diagnostic tool for detecting oesophageal pathology: the PillCam oesophageal video capsule. Aliment Pharmacol Ther 2004; 20: 1083-1089 [PMID: 15569110 DOI: 10.1111/j.1365-2036.2004.02206.x]

10	Eliakim R, Fireman Z, Gralnek IM, Yassin K, Waterman M, Kopelman Y, Lachter J, Koslowsky B, Adler SN. Evaluation of the PillCam Colon capsule in the detection of colonic pathology: results of the first multicenter, prospective, comparative study. Endoscopy 2006; 38: 963-970 [PMID: 17058158 DOI: 10.1055/s-2006-944832]

11	Eliakim R. Capsule endoscopy - Where are we at 2011 and where are we headed? Intest Res 2012; 10: 235-243 [DOI: 10.5217/jr.2012.10.3.235]

12	Spada C, Hassan C, Galmiche JP, Neuhaus H, Dumonceau JM, Adler S, Epstein O, Gay G, Pennazio M, Rex DK, Benamouzig R, de Franchis R, Delvaux M, Devière J, Eliakim R, Fraser C, Hagenmuller F, Herrerias JM, Keuchel M, Macrae F, Munoz-Navas M, Ponchon T, Quintero E, Riccioni ME, Rondonotti E, Marmo R, Sung JJ, Tajiri H, Toth E, Triantafyllou K, Van Gossum A, Costamagna G. Colon capsule endoscopy: European Society of Gastrointestinal Endoscopy (ESGE) Guideline. Endoscopy 2012; 44: 527-536 [PMID: 22389230 DOI: 10.1055/s-0031-1291717]

13	Wang A, Banerjee S, Barth BA, Bhat YM, Chauhan S, Gottlieb KT, Konda V, Maple JT, Murad F, Pfau PR, Pleskow DK, Siddiqui UD, Tokar JL, Rodriguez SA. Wireless capsule endoscopy. Gastrointest Endosc 2013; 78: 805-815 [PMID: 24119509 DOI: 10.1016/j.gie.2013.06.026]

14	Appleyard M, Glukhovsky A, Swain P. Wireless-capsule diagnostic endoscopy for recurrent small-bowel bleeding. N Engl J Med 2001; 344: 232-233 [PMID: 11188844 DOI: 10.1056/NEJM200101183440316]

15	Kopylov U, Seidman EG. Clinical applications of small bowel capsule endoscopy. Clin Exp Gastroenterol 2013; 6: 129-137 [PMID: 23983481 DOI: 10.2147/CEG.S48005]

16	Laine L, Sahota A, Shah A. Does capsule endoscopy improve outcomes in obscure gastrointestinal bleeding? Randomized trial versus dedicated small bowel radiography. Gastroenterology 2010; 138: 1673-1680.e1; quiz e11-12 [PMID: 20138043 DOI: 10.1053/j.gastro.2010.01.047]

17	Triester SL, Leighton JA, Leontiadis GI, Fleischer DE, Hara AK, Heigh RI, Shiff AD, Sharma VK. A meta-analysis of the yield of capsule endoscopy compared to other diagnostic modalities in patients with obscure gastrointestinal bleeding. Am J Gastroenterol 2005; 100: 2407-2418 [PMID: 16279893 DOI: 10.1111/j.1572-0241.2005.00274.x]

18	Saperas E, Dot J, Videla S, Alvarez-Castells A, Perez-Lafuente M, Armengol JR, Malagelada JR. Capsule endoscopy versus computed tomographic or standard angiography for the diagnosis of obscure gastrointestinal bleeding. Am J Gastroenterol 2007; 102: 731-737 [PMID: 17397406 DOI: 10.1111/j.1572-0241.2007.01058.x]

19	Koulaouzidis A, Rondonotti E, Giannakou A, Plevris JN. Diagnostic yield of small-bowel capsule endoscopy in patients with iron-deficiency anemia: a systematic review. Gastrointest Endosc 2012; 76: 983-992 [PMID: 23078923 DOI: 10.1016/j.gie.2012.07.035]

20	Hartmann D, Schmidt H, Bolz G, Schilling D, Kinzel F, Eickhoff A, Huschner W, Möller K, Jakobs R, Reitzig P, Weickert U, Gellert K, Schultz H, Guenther K, Hollerbuhl H, Schoenleben K, Schulz HJ, Riemann JF. A prospective two-center study comparing wireless capsule endoscopy with intraoperative enteroscopy in patients with obscure GI bleeding. Gastrointest Endosc 2005; 61: 826-832 [PMID: 15933683 DOI: 10.1016/S0016-5107(05)00372-X]

21	de Leusse A, Vahedi K, Edery J, Tiah D, Fery-Lemonnier E, Cellier C, Bouhnik Y, Jian R. Capsule endoscopy or push enteroscopy for first-line exploration of obscure gastrointestinal bleeding? Gastroenterology 2007; 132: 855-862; quiz 1164-1165 [PMID: 17324401 DOI: 10.1053/j.gastro.2006.12.002]

22	Pasha SF, Leighton JA, Das A, Harrison ME, Decker GA, Fleischer DE, Sharma VK. Double-balloon enteroscopy and capsule endoscopy have comparable diagnostic yield in small-bowel disease: a meta-analysis. Clin Gastroenterol Hepatol 2008; 6: 671-676 [PMID: 18356113 DOI: 10.1016/j.cgh.2008.01.005]

23	Cañas-Ventura A, Márquez L, Bessa X, Dedeu JM, Puigvehí M, Delgado-Aros S, Ibáñez IA, Seoane A, Barranco L, Bory F, Andreu M, González-Suárez B. Outcome in obscure gastrointestinal bleeding after capsule endoscopy. World J Gastrointest Endosc 2013; 5: 551-558 [PMID: 24255747 DOI: 10.4253/wjge.v5.i11.551]

24	Min YW, Kim JS, Jeon SW, Jeen YT, Im JP, Cheung DY, Choi MG, Kim JO, Lee KJ, Ye BD, Shim KN, Moon JS, Kim JH, Hong SP, Chang DK. Long-term outcome of capsule endoscopy in obscure gastrointestinal bleeding: a nationwide analysis. Endoscopy 2014; 46: 59-65 [PMID: 24254387 DOI: 10.1055/s-0033-1358803]

25	Leighton JA, Gralnek IM, Cohen SA, Toth E, Cave DR, Wolf DC, Mullin GE, Ketover SR, Legnani PE, Seidman EG, Crowell MD, Bergwerk AJ, Peled R, Eliakim R. Capsule endoscopy is superior to small-bowel follow-through and equivalent to ileocolonoscopy in suspected Crohn’s disease. Clin Gastroenterol Hepatol 2014; 12: 609-615 [PMID: 24075891 DOI: 10.1016/j.cgh.2013.09.028]

26	Triester SL, Leighton JA, Leontiadis GI, Gurudu SR, Fleischer DE, Hara AK, Heigh RI, Shiff AD, Sharma VK. A meta-analysis of the yield of capsule endoscopy compared to other diagnostic modalities in patients with non-stricturing small bowel Crohn’s disease. Am J Gastroenterol 2006; 101: 954-964 [PMID: 16696781 DOI: 10.1111/j.1572-0241.2006.00506.x]

27	Dionisio PM, Gurudu SR, Leighton JA, Leontiadis GI, Fleischer DE, Hara AK, Heigh RI, Shiff AD, Sharma VK. Capsule endoscopy has a significantly higher diagnostic yield in patients with suspected and established small-bowel Crohn’s disease: a meta-analysis. Am J Gastroenterol 2010; 105: 1240-1248; quiz 1249 [PMID: 20029412 DOI: 10.1038/ajg.2009.713]

28	Crook DW, Knuesel PR, Froehlich JM, Eigenmann F, Unterweger M, Beer HJ, Kubik-Huch RA. Comparison of magnetic resonance enterography and video capsule endoscopy in evaluating small bowel disease. Eur J Gastroenterol Hepatol 2009; 21: 54-65 [PMID: 19086147 DOI: 10.1097/MEG.0b013e32830ce7a7]

29	Hall B, Holleran G, McNamara D. Current applications and potential future role of wireless capsule technology in Crohn’s disease. Scand J Gastroenterol 2014; 49: 1275-1284 [PMID: 25260016 DOI: 10.3109/00365521.2014.962606]

30	Cotter J, Dias de Castro F, Moreira MJ, Rosa B. Tailoring Crohn’s disease treatment: the impact of small bowel capsule endoscopy. J Crohns Colitis 2014; 8: 1610-1615 [PMID: 24631311 DOI: 10.1016/j.crohns.2014.02.018]

31	Niv E, Fishman S, Kachman H, Arnon R, Dotan I. Sequential capsule endoscopy of the small bowel for follow-up of patients with known Crohn’s disease. J Crohns Colitis 2014; 8: 1616-1623 [PMID: 24666976 DOI: 10.1016/j.crohns.2014.03.003]

32	Rondonotti E, Spada C, Cave D, Pennazio M, Riccioni ME, De Vitis I, Schneider D, Sprujevnik T, Villa F, Langelier J, Arrigoni A, Costamagna G, de Franchis R. Video capsule enteroscopy in the diagnosis of celiac disease: a multicenter study. Am J Gastroenterol 2007; 102: 1624-1631 [PMID: 17459022 DOI: 10.1111/j.1572-0241.2007.01238.x]

33	Rokkas T, Niv Y. The role of video capsule endoscopy in the diagnosis of celiac disease: a meta-analysis. Eur J Gastroenterol Hepatol 2012; 24: 303-308 [PMID: 22266837 DOI: 10.1097/MEG.0b013e32834fa914]

34	Brown G, Fraser C, Schofield G, Taylor S, Bartram C, Phillips R, Saunders B. Video capsule endoscopy in peutz-jeghers syndrome: a blinded comparison with barium follow-through for detection of small-bowel polyps. Endoscopy 2006; 38: 385-390 [PMID: 16680639]

35	Koornstra JJ. Small bowel endoscopy in familial adenomatous polyposis and Lynch syndrome. Best Pract Res Clin Gastroenterol 2012; 26: 359-368 [PMID: 22704577 DOI: 10.1016/j.bpg.2012.01.022]

36	Schwartz GD, Barkin JS. Small-bowel tumors detected by wireless capsule endoscopy. Dig Dis Sci 2007; 52: 1026-1030 [PMID: 17380403 DOI: 10.1007/s10620-006-9483-8]

37	Urgesi R, Riccioni ME, Bizzotto A, Cianci R, Spada C, Pelecca G, Ricci R, Costamagna G. Increased diagnostic yield of small bowel tumors with PillCam: the role of capsule endoscopy in the diagnosis and treatment of gastrointestinal stromal tumors (GISTs). Italian single-center experience. Tumori 2012; 98: 357-363 [PMID: 22825512 DOI: 10.1700/1125.12405]

38	Girelli CM, Porta P, Colombo E, Lesinigo E, Bernasconi G. Development of a novel index to discriminate bulge from mass on small-bowel capsule endoscopy. Gastrointest Endosc 2011; 74: 1067-1074; quiz 1115.e1-5 [PMID: 21907982 DOI: 10.1016/j.gie.2011.07.022]

39	Rondonotti E, Pennazio M, Toth E, Menchen P, Riccioni ME, De Palma GD, Scotto F, De Looze D, Pachofsky T, Tacheci I, Havelund T, Couto G, Trifan A, Kofokotsios A, Cannizzaro R, Perez-Quadrado E, de Franchis R. Small-bowel neoplasms in patients undergoing video capsule endoscopy: a multicenter European study. Endoscopy 2008; 40: 488-495 [PMID: 18464193 DOI: 10.1055/s-2007-995783]

40	Cheung DY, Lee IS, Chang DK, Kim JO, Cheon JH, Jang BI, Kim YS, Park CH, Lee KJ, Shim KN, Ryu JK, Do JH, Moon JS, Ye BD, Kim KJ, Lim YJ, Choi MG, Chun HJ. Capsule endoscopy in small bowel tumors: a multicenter Korean study. J Gastroenterol Hepatol 2010; 25: 1079-1086 [PMID: 20594222 DOI: 10.1111/j.1440-1746.2010.06292.x]

41	Trifan A, Singeap AM, Cojocariu C, Sfarti C, Stanciu C. Small bowel tumors in patients undergoing capsule endoscopy: a single center experience. J Gastrointestin Liver Dis 2010; 19: 21-25 [PMID: 20361070]

42	Nakamura T, Terano A. Capsule endoscopy: past, present, and future. J Gastroenterol 2008; 43: 93-99 [PMID: 18306982 DOI: 10.1007/s00535-007-2153-6]

43	Spada C, Hassan C, Barbaro B, Iafrate F, Cesaro P, Petruzziello L, Minelli Grazioli L, Senore C, Brizi G, Costamagna I, Alvaro G, Iannitti M, Salsano M, Ciolina M, Laghi A, Bonomo L, Costamagna G. Colon capsule versus CT colonography in patients with incomplete colonoscopy: a prospective, comparative trial. Gut 2015; 64: 272-281 [PMID: 24964317 DOI: 10.1136/gutjnl-2013-306550]

44	Negreanu L, Babiuc R, Bengus A, Sadagurschi R. PillCam Colon 2 capsule in patients unable or unwilling to undergo colonoscopy. World J Gastrointest Endosc 2013; 5: 559-567 [PMID: 24255748 DOI: 10.4253/wjge.v5.i11.559]

45	Roberts-Thomson IC, Singh R, Teo E, Nguyen NQ, Lidums I. The future of endoscopy. J Gastroenterol Hepatol 2010; 25: 1051-1057 [PMID: 20594218 DOI: 10.1111/j.1440-1746.2010.06333.x]

46	Remes-Troche JM, Jiménez-García VA, García-Montes JM, Hergueta-Delgado P, Roesch-Dietlen F, Herrerías-Gutiérrez JM. Application of colon capsule endoscopy (CCE) to evaluate the whole gastrointestinal tract: a comparative study of single-camera and dual-camera analysis. Clin Exp Gastroenterol 2013; 6: 185-192 [PMID: 24068872 DOI: 10.2147/CEG.S45215]

47	Quirini M, Menciassi A, Scapellato S, Dario P, Rieber F, Ho CN, Schostek S, Schurr MO. Feasibility proof of a legged locomotion capsule for the GI tract. Gastrointest Endosc 2008; 67: 1153-1158 [PMID: 18513557 DOI: 10.1016/j.gie.2007.11.052]

48	Sliker LJ, Ciuti G. Flexible and capsule endoscopy for screening, diagnosis and treatment. Expert Rev Med Devices 2014; 11: 649-666 [PMID: 25148269 DOI: 10.1586/17434440.2014.941809]

49	Kim HM, Yang S, Kim J, Park S, Cho JH, Park JY, Kim TS, Yoon ES, Song SY, Bang S. Active locomotion of a paddling-based capsule endoscope in an in vitro and in vivo experiment (with videos). Gastrointest Endosc 2010; 72: 381-387 [PMID: 20497903 DOI: 10.1016/j.gie.2009.12.058]

50	Ciuti G, Menciassi A, Dario P. Capsule endoscopy: from current achievements to open challenges. IEEE Rev Biomed Eng 2011; 4: 59-72 [PMID: 22273791 DOI: 10.1109/RBME.2011.2171182]

51	Quirini M, Menciassi A, Scapellato S, Stefanini C, Dario P. Design and fabrication of a motor legged capsule for the active exploration of the gastrointestinal tract. IEEE/ASME Trans Mechatron 2008; 13: 169-179 [DOI: 10.11009/TMECH.2008.918491]

52	Valdastri P, Webster RJ, Quaglia C, Quirini M, Menciassi A, Dario P. A new mechanism for mesoscale legged locomotion in compliant tubular environments. IEEE Trans Robot 2009; 25: 1047-1057 [DOI: 10.1109/TRO.2009.2014127]

53	Quaglia C, Buselli E, Webster RJ, Valdastri P, Menciassi A, Dario P. An endoscopic capsule robot: A meso-scale engineering case study. J Micromech Microeng 2009; 19: 1-11 [DOI: 10.1088/0960-1317/19/10/105007]

54	Swain P, Toor A, Volke F, Keller J, Gerber J, Rabinovitz E, Rothstein RI. Remote magnetic manipulation of a wireless capsule endoscope in the esophagus and stomach of humans (with videos). Gastrointest Endosc 2010; 71: 1290-1293 [PMID: 20417507 DOI: 10.1016/j.gie.2010.01.064]

55	Keller J, Fibbe C, Volke F, Gerber J, Mosse AC, Reimann-Zawadzki M, Rabinovitz E, Layer P, Schmitt D, Andresen V, Rosien U, Swain P. Inspection of the human stomach using remote-controlled capsule endoscopy: a feasibility study in healthy volunteers (with videos). Gastrointest Endosc 2011; 73: 22-28 [PMID: 21067740 DOI: 10.1016/j.gie.2010.08.053]

56	Rey JF, Ogata H, Hosoe N, Ohtsuka K, Ogata N, Ikeda K, Aihara H, Pangtay I, Hibi T, Kudo S, Tajiri H. Feasibility of stomach exploration with a guided capsule endoscope. Endoscopy 2010; 42: 541-545 [PMID: 20593331 DOI: 10.1055/s-0030-1255521]

57	Siemens and Olympus Healthcare. Magnetically guided capsule endoscope system for comfortable examination of the stomach - More than 50 participants in the fisrt successful study. Available from: URL: http://www.siemens.com/press/en/materials/healthcare/2010-10-Kapselendoskopie.php

58	Rey JF. The future of capsule endoscopy. Keio J Med 2013; 62: 41-46 [PMID: 23708295 DOI: 10.2302/kjm.2012-0011-RE]

59	Rey JF, Ogata H, Hosoe N, Ohtsuka K, Ogata N, Ikeda K, Aihara H, Pangtay I, Hibi T, Kudo SE, Tajiri H. Blinded nonrandomized comparative study of gastric examination with a magnetically guided capsule endoscope and standard videoendoscope. Gastrointest Endosc 2012; 75: 373-381 [PMID: 22154417 DOI: 10.1016/j.gie.2011.09.030]

60	Rahman I, Pioche M, Shim CS, Sung IK, Saurin JC, Patel P. Magnet assisted capsule endoscopy (MACE) in the upper GI tract is feasible: first human series using the novel Mirocam-Navi System. Gastrointest Endosc 2014; 79: AB122 [DOI: 10.1016/j.gie.2014.02.059]

61	Carpi F, Pappone C. Stereotaxis Niobe magnetic navigation system for endocardial catheter ablation and gastrointestinal capsule endoscopy. Expert Rev Med Devices 2009; 6: 487-498 [PMID: 19751121 DOI: 10.1586/erd.09.32]

62	Ciuti G, Donlin R, Valdastri P, Arezzo A, Menciassi A, Morino M, Dario P. Robotic versus manual control in magnetic steering of an endoscopic capsule. Endoscopy 2010; 42: 148-152 [PMID: 20017088 DOI: 10.1055/s-0029-1243808]

63	Schostek S, Schurr MO. European research on wireless endoscopy--the VECTOR project. Stud Health Technol Inform 2013; 189: 193-199 [PMID: 23739381]

64	Community Research and Development Information Service. Nano based capsule-Endoscopy with molecular imaging and optical biopsy. Available from: URL: http://www.cordis.europa.eu/project/rcn/84969_en.html

65	Wang L, Zhang G, Luo JC, Zeng F, Wang QZ, Alfano SA, Katz A, Zevallos M, Alfano RR. Wireless spectroscopic compact photonic explorer for diagnostic optical imaging. Biomed Microdevices 2005; 7: 111-115 [PMID: 15940423 DOI: 10.1007/s10544-005-1588-x]

66	Woo SH, Kim TW, Mohy-Ud-Din Z, Park IY, Cho JH. Small intestinal model for electrically propelled capsule endoscopy. Biomed Eng Online 2011; 10: 108 [PMID: 22177218 DOI: 10.1186/1475-925X-10-108]

67	Pasricha T, Smith BF, Mitchell VR, Fang B, Brooks ER, Gerding JS, Washington MK, Valdastri P, Obstein KL. Controlled colonic insufflation by a remotely triggered capsule for improved mucosal visualization. Endoscopy 2014; 46: 614-618 [PMID: 24845802 DOI: 10.1055/s-0034-1365497]

68	Sharma VK. The future is wireless: advances in wireless diagnostic and therapeutic technologies in gastroenterology. Gastroenterology 2009; 137: 434-439 [PMID: 19545570 DOI: 10.1053/j.gastro.2009.06.029]

69	Kong K, Cha J, Jeon D, Cho D. A rotational micro biopsy device for the capsule endoscope. In: IROS 2005: IEEE/RSJ International Conference on Intelligent Robots and Systems. Alberta, Canada: Edmonton, 2005: 1839-1843

70	Ray K. Endoscopy: Tethered capsule endomicroscopy of the oesophagus--an easy pill to swallow. Nat Rev Gastroenterol Hepatol 2013; 10: 129 [PMID: 23358395 DOI: 10.1038/nrgastro.2013.15]

71	Yim S, Gultepe E, Gracias DH, Sitti M. Biopsy using a magnetic capsule endoscope carrying, releasing, and retrieving untethered microgrippers. IEEE Trans Biomed Eng 2014; 61: 513-521 [PMID: 24108454]

72	Simi M, Gerboni G, Menciassi A, Valdastri P. Magnetic torsion spring mechanism for a wireless biopsy capsule. J Med Device 2013; 7: 041009-9 [DOI: 10.1115/1.4025185]

73	Zhang H, Morgan D, Cecil G, Burkholder A, Ramocki N, Scull B, Lund PK. Biochromoendoscopy: molecular imaging with capsule endoscopy for detection of adenomas of the GI tract. Gastrointest Endosc 2008; 68: 520-527 [PMID: 18499106 DOI: 10.1016/j.gie.2008.02.023]

74	Pikul JH, Gang Zhang H, Cho J, Braun PV, King WP. High-power lithium ion microbatteries from interdigitated three-dimensional bicontinuous nanoporous electrodes. Nat Commun 2013; 4: 1732 [PMID: 23591899 DOI: 10.1038/ncomms2747]

75	Rathore P, Tiwary D, Kumar N. A methodology for very low power and small size capsule manufacturing for wireless capsule endoscopy. Int J Comput Appl T 2014; 95: 10-14 [DOI: 10.5120/16725-6905]

76	Wu J, Li Y. Low-complexity video compression for capsule endoscope based on compressed sensing theory. Conf Proc IEEE Eng Med Biol Soc 2009; 2009: 3727-3730 [PMID: 19965013 DOI: 10.1109/IEMBS.2009.5334819]

77	Basar MR, Ahmad MY, Cho J, Ibrahim F. Application of wireless power transmission systems in wireless capsule endoscopy: an overview. Sensors (Basel) 2014; 14: 10929-10951 [PMID: 24949645 DOI: 10.3390/s140610929]

78	Goodman RR, Torres RA, McMurtry JG. Acoustic schwannoma and epidermoid cyst occurring as a single cerebellopontine angle mass. Neurosurgery 1991; 28: 433-436 [PMID: 2011228 DOI: 10.1002/rcs.298]

79	Donnelly PK, Simpson AR. Cytotoxic cross-matching for organ transplantation. Lancet 1991; 337: 1553-1554 [PMID: 1675410 DOI: 10.1080/13645700310014933]

80	Jia Z, Yan G, Liu H, Wang Z, Jiang P, Shi Y. The optimization of wireless power transmission: design and realization. Int J Med Robot 2012; 8: 337-347 [PMID: 22508580 DOI: 10.1002/rcs.1428]

81	Basar MR, Malek F, Juni Khairudi M, Shaharom Idris M, Iskandar M. Saleh M. Ingestible Wireless Capsule Technology: A Review of Development and Future Indication. Inter J of Antenna and Propagation 2012; 2012: 807165 [DOI: 10.1155/2012/807165]

82	Roberts M, Johns P, Owen J, Brandell D, Edstrom K, El Enany G, Guery C, Golodnitsky D, Lacey M, Lecoeur C, Mazor H, Peled E, Perre E, Shaijumon M, Simon P, Taberna PL. 3D lithium ion batteries - from fundamentals to fabrication. J Mater Chem 2011; 21: 9876-9890

83	Woods SP, Constandinou TG. Wireless capsule endoscope for targeted drug delivery: mechanics and design considerations. IEEE Trans Biomed Eng 2013; 60: 945-953 [PMID: 23192476 DOI: 10.1109/TBME.2012.2228647]

84	Munoz F, Alici G, Li W. A review of drug delivery systems for capsule endoscopy. Adv Drug Deliv Rev 2014; 71: 77-85 [PMID: 24384373 DOI: 10.1016/j.addr.2013.12.007]

85	Phillips Technology. Phillips Intelligent Pill Technology. Available from: URL: http://www.research.philips.com/newscenter/backgrounders/081111-ipill.html

86	Hale MF, Sidhu R, McAlindon ME. Capsule endoscopy: current practice and future directions. World J Gastroenterol 2014; 20: 7752-7759 [PMID: 24976712 DOI: 10.3748/wjg.v20.i24.7752]

87	Quaglia C, Tognarelli S, Sinibaldi E, Funaro N, Dario P, Menciassi A. Wireless robotic capsule for releasing bioadhesive patches in the gastrointestinal tract. J Med Device 2013; 8: 014503-3 [DOI: 10.1115/1.4025450]

88	Yim S, Goyal K, Sitti M. Magnetically Actuated Soft Capsule With the Multimodal Drug Release Function. IEEE ASME Trans Mechatron 2013; 18: 1413-1418 [PMID: 25378896 DOI: 10.1109/TMECH.2012.2235077]

89	Fireman Z. Capsule endoscopy: Future horizons. World J Gastrointest Endosc 2010; 2: 305-307 [PMID: 21160761 DOI: 10.4253/wjge.v2.i9.305]

90	Toennies JL, Ciuti G, Smith BF, Menciassi A, Valdastri P, Webster RJ. Toward tetherless insufflation of the GI Tract. Conf Proc IEEE Eng Med Biol Soc 2010; 2010: 1946-1949 [PMID: 21097004 DOI: 10.1109/IEMBS.2010.5627793]

91	Gorlewicz JL, Battaglia S, Smith BF, Ciuti G, Gerding J, Menciassi A, Obstein KL, Valdastri P, Webster RJ. Wireless insufflation of the gastrointestinal tract. IEEE Trans Biomed Eng 2013; 60: 1225-1233 [PMID: 23212312 DOI: 10.1109/TBME.2012.2230631]

92	Koulaouzidis A, Iakovidis DK, Karargyris A, Plevris JN. Optimizing lesion detection in small-bowel capsule endoscopy: from present problems to future solutions. Expert Rev Gastroenterol Hepatol 2015; 9: 217-235 [PMID: 25169106 DOI: 10.1586/17474124.2014.952281]

93	Adler SN, Hassan C, Metzger Y, Sompolinsky Y, Spada C. Second-generation colon capsule endoscopy is feasible in the out-of-clinic setting. Surg Endosc 2014; 28: 570-575 [PMID: 24043646 DOI: 10.1007/s00464-013-3206-y]

94	Bouchard S, Ibrahim M, Van Gossum A. Video capsule endoscopy: perspectives of a revolutionary technique. World J Gastroenterol 2014; 20: 17330-17344 [PMID: 25516644 DOI: 10.3748/wjg.v20.i46.17330]

95	Chandrappan J, Ruiqi L, Su N, Qiang TS, Vaidyanathan K. Tagging module for lesion localization in capsule endoscopy. Conf Proc IEEE Eng Med Biol Soc 2010; 2010: 1890-1893 [PMID: 21096425 DOI: 10.1109/IEMBS.2010.5627090]

96	Goenka MK, Majumder S, Goenka U. Capsule endoscopy: Present status and future expectation. World J Gastroenterol 2014; 20: 10024-10037 [PMID: 25110430 DOI: 10.3748/wjg.v20.i29.10024]

97	Woo SH, Mohy-Ud-Din Z, Cho JH. Duodenum identification mechanism for capsule endoscopy. IEEE Trans Biomed Eng 2011; 58: 905-912 [PMID: 21134813 DOI: 10.1109/TBME.2010.2095849]

98	Fisher LR, Hasler WL. New vision in video capsule endoscopy: current status and future directions. Nat Rev Gastroenterol Hepatol 2012; 9: 392-405 [PMID: 22565098 DOI: 10.1038/nrgastro.2012.88]

99	Ibrahim M, Van Gossum A. Novel imaging enhancements in capsule endoscopy. Gastroenterol Res Pract 2013; 2013: 304723 [PMID: 23878532 DOI: 10.1155/2013/304723]

100	Selby WS, Prakoso E. The inability to visualize the ampulla of Vater is an inherent limitation of capsule endoscopy. Eur J Gastroenterol Hepatol 2011; 23: 101-103 [PMID: 21030868 DOI: 10.1097/MEG.0b013e3283410210]

101	Koulaouzidis A, Douglas S, Plevris JN. Identification of the ampulla of Vater during oesophageal capsule endoscopy: two heads and viewing speed make a difference. Eur J Gastroenterol Hepatol 2011; 23: 361; author reply 362 [PMID: 21430448 DOI: 10.1097/MEG.0b013e3283440443]

102	Gupta A, Postgate AJ, Burling D, Ilangovan R, Marshall M, Phillips RK, Clark SK, Fraser CH. A prospective study of MR enterography versus capsule endoscopy for the surveillance of adult patients with Peutz-Jeghers syndrome. AJR Am J Roentgenol 2010; 195: 108-116 [PMID: 20566803 DOI: 10.2214/AJR.09.3174]

103	Clarke JO, Giday SA, Magno P, Shin EJ, Buscaglia JM, Jagannath SB, Mullin GE. How good is capsule endoscopy for detection of periampullary lesions? Results of a tertiary-referral center. Gastrointest Endosc 2008; 68: 267-272 [PMID: 18378233 DOI: 10.1016/j.gie.2007.11.055]

104	Romero-Vázquez J, Argüelles-Arias F, García-Montes JM, Caunedo-Álvarez Á, Pellicer-Bautista FJ, Herrerías-Gutiérrez JM. Capsule endoscopy in patients refusing conventional endoscopy. World J Gastroenterol 2014; 20: 7424-7433 [PMID: 24966612 DOI: 10.3748/wjg.v20.i23.7424]

105	Hosoe N, Naganuma M, Ogata H. Current status of capsule endoscopy through a whole digestive tract. Dig Endosc 2015; 27: 205-215 [PMID: 25208463 DOI: 10.1111/den.12380]

Footnotes

Conflict-of-interest statement: There is no conflict of interest associated with any of the senior author or other coauthors contributed their efforts in this manuscript.

Open-Access: This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

Peer-review started: May 9, 2015

First decision: July 19, 2015

Article in press: September 30, 2015

P- Reviewer: Cerwenka HR, de’Angelis GL, Inamori M, Naito Y    S- Editor: Ma YJ    L- Editor: Filopodia    E- Editor: Ma S