Copyright Information of the Article Published Online

| TITLE | Methods, units and quality requirements for the analysis of haemoglobin A1c in diabetes mellitus |
| --- | --- |
| AUTHOR(s) | Ilkka Penttilä, Karri Penttilä, Päivi Holm, Harri Laitinen, Päivi Ranta, Jukka Törrönen, Rainer Rauramaa |
| CITATION | Penttilä I, Penttilä K, Holm P, Laitinen H, Ranta P, Törrönen J, Rauramaa R. Methods, units and quality requirements for the analysis of haemoglobin A1c in diabetes mellitus. World J Methodol 2016; 6(2): 133-142 |
| URL | http://www.wjgnet.com/2222-0682/full/v6/i2/133.htm |
| DOI | http://dx.doi.org/10.5662/wjm.v6.i2.133 |
| OPEN ACCESS | This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/ |
| CORE TIP | The aim of this review is to clarify methods, units, quality requirements, reference and cutoff limits for hemoglobin A1c (HbA1c) and ratio of blood glucose/HbA1c on the basis of the results from Finnish quality control surveys by comparing them to the literature. The HbA1c surveys of Labquality Ltd. (Helsinki, Finland) were started in 1986 by using two fresh EDTA-blood samples. From 1994, the number of the participating laboratories had risen to 139, of which 75 were Finnish and 64 from five other countries. In 2014, the number of the participating laboratories was total 214, 141 were Finnish and 73 from 13 other countries. |
| KEY WORDS | Diabetes; Hemoglobin A1c; Glycohemoglobin; Glucose; International Federation of Clinical Chemistry and Laboratory Medicine; Reference values; Quality assurance; Recommendation; Target limits |
| COPYRIGHT | © The Author(s) 2016. Published by Baishideng Publishing Group Inc. All rights reserved. |
| NAME OF JOURNAL | World Journal of Methodology |
| ISSN | 2222-0682 (online) |
| PUBLISHER | Baishideng Publishing Group Inc, 8226 Regency Drive, Pleasanton, CA 94588, USA |
| WEBSITE | http://www.wjgnet.com |

MINIREVIEWS

Methods, units and quality requirements for the analysis of haemoglobin A1c in diabetes mellitus

Ilkka Penttilä, Karri Penttilä, Päivi Holm, Harri Laitinen, Päivi Ranta, Jukka Törrönen, Rainer Rauramaa

Ilkka Penttilä, Jukka Törrönen, Rainer Rauramaa, Kuopio Research Institute of Exercise Medicine, University of Eastern Finland, 70100 Kuopio, Finland

Karri Penttilä, Finnish Medicines Agency FIMEA, 70210 Kuopio, Finland

Päivi Holm, Fimlab Laboratories Ltd., 30430 Tampere, Finland

Harri Laitinen, Päivi Ranta, Labquality Ltd., 00530 Helsinki, Finland

Author contributions: All authors equally contributed to this paper with conception and design of the study, literature review and analysis, drafting and critical revision and editing, and final approval of the final version.

Correspondence to: Ilkka Penttilä, MD, PhD, Kuopio Research Institute of Exercise Medicine, University of Eastern Finland, Haapaniementie 16, 70100 Kuopio, Finland. ilkka.penttila@uef.fi

Telephone: +358-40-5825564

Received: September 19, 2015          Revised: March 1, 2016        Accepted: March 24, 2016

Published online: June 26, 2016

Abstract

The formation of glycohemoglobin, especially the hemoglobin A1c (HbA1c) fraction, occurs when glucose becomes coupled with the amino acid valine in the -chain of Hb; this reaction is dependent on the plasma concentration of glucose. Since the early 1970s it has been known that diabetics display higher values of HbA1c because they have elevated blood glucose concentrations. Thus HbA1c has acquired a very important role in the treatment and diagnosis of diabetes mellitus. After the introduction of the first quantitative measurement of HbA1c, numerous methods for glycohemoglobin have been introduced with different assay principles: From a simple mini-column technique to the very accurate automated high-pressure chromatography and lastly to many automated immunochemical or enzymatic assays. In early days, the results of the quality control reports for HbA1c varied extensively between laboratories, therefore in United States and Canada working groups (WG) of the Diabetes Controls and Complications Trial (DCCT) were set up to standardize the HbA1c assays against the DCCT/National Glycohemoglobin Standardization Program reference method based on liquid chromatography. In the 1990s, the International Federation of Clinical Chemistry and Laboratory Medicine (IFCC) appointed a new WG to plan a reference preparation and method for the HBA1c measurement. When the reference procedures were established, in 2004 IFCC recommended that all manufacturers for equipment used in HbA1c assays should calibrate their methods to their proposals. This led to an improvement in the coefficient of variation (CV%) associated with the assay. In this review, we describe the glycation of Hb, methods, standardization of the HbA1c assays, analytical problems, problems with the units in which HbA1c values are expressed, reference values, quality control aspects, target requirements for HbA1c, and the relationship of the plasma glucose values to HbA1c concentrations. We also note that the acceptance of the mmol/mol system for HbA1c as recommended by IFCC, i.e., the new unit and reference ranges, are becoming only slowly accepted outside of Europe where it seems that expressing HbA1c values either only in per cent units or with parallel reporting of percent and mmol/mol will continue. We believe that these issues should be resolved in the future and that it would avoid confusion if mmol/mol unit for HbA1c were to gain worldwide acceptance.

Key words: Diabetes; Hemoglobin A1c; Glycohemoglobin; Glucose; International Federation of Clinical Chemistry and Laboratory Medicine; Reference values; Quality assurance; Recommendation; Target limits

Penttilä I, Penttilä K, Holm P, Laitinen H, Ranta P, Törrönen J, Rauramaa R. Methods, units and quality requirements for the analysis of haemoglobin A1c in diabetes mellitus. World J Methodol 2016; 6(2): 133-142  Available from: URL: http://www.wjgnet.com/2222-0682/full/v6/i2/133.htm  DOI: http://dx.doi.org/10.5662/wjm.v6.i2.133

Core tip: The aim of this review is to clarify methods, units, quality requirements, reference and cutoff limits for hemoglobin A1c (HbA1c) and ratio of blood glucose/HbA1c on the basis of the results from Finnish quality control surveys by comparing them to the literature. The HbA1c surveys of Labquality Ltd. (Helsinki, Finland) were started in 1986 by using two fresh EDTA-blood samples. From 1994, the number of the participating laboratories had risen to 139, of which 75 were Finnish and 64 from five other countries. In 2014, the number of the participating laboratories was total 214, 141 were Finnish and 73 from 13 other countries.

INTRODUCTION

Fisher[1] synthetized a molecule named fructosamine in 1889. In the 1950s and 60s, it was reported that carbohydrate residues could become attached to hemoglobin (Hb)[2,3]. In 1967 Holmquist and Schroeder[4] observed that glucose bound to Hb and Rahbar[5] revealed that the electrophoretic fraction containing glucose was higher in blood samples from diabetic subjects than in healthy subjects. Subsequently, Shapiro et al[6] described that many carbohydrate components including glucose could be bound to Hb and that glucose was most efficiently bound to the β-chain of Hb. After Trivelli et al[7] published their quantitative measurement of Hb fractions in human blood, the analysis of HbA1c was recognized as being a very important parameter in the assessment of diabetic patients.

In the early years, the results of HbA1c analysis varied extensively between methods and laboratories[8-12]. Therefore during the 1980s, in United States and Canada working groups (WG) of the Diabetes Controls and Complications Trial (DCCT) were set up; these were originally incorporated within a multicenter, randomized clinical trial designed to compare treatments of insulin dependent diabetes mellitus in the National Glycohemoglobin Standardization Program (NGSP), but subsequently this activity was expanded to standardize the HbA1c assays to the DCCT/NGSP reference method of liquid chromatography[9].

Later in the 1990s, the International Federation of Clinical Chemistry (IFCC) decided to develop a reference system for the international standardization of HbA1c/glycohemoglobin measurements[13], which became the basis of the reference laboratory network for HbA1c. In addition, IFCC set up another WG on Standardization for HbA1c (1994/1995) in order to develop a primary standard[14] and reference method[15] for worldwide use in the HbA1c measurement. These were the reasons why from the 1990s, IFCC organized WG to promote the standardization of all types of assays for HbA1c. The adoption of these recommendations achieved substantial improvements in the analytical aspects and clinical significance of HbA1c[16,17] so that today reference standards and standardized methods are in widespread use.

Further in 2010, the American Diabetes Association (ADA) emphasized the important role of HbA1c in the diagnosis of diabetes, setting an analytical cutoff limit value of 6.5% for HbA1c values[18] if they were expressed in percentage terms. Recently Hanas and John[19] reported the conclusions of the International Consensus Committee 2013 Update that HbA1c results should be reported by clinical laboratories worldwide in Système Internationale (SI) units (mmol/mol - no decimals) and the corresponding NGSP units (% - one decimal) and recommended strongly to the editors of scientific journals that submitted manuscripts should report HbA1c values in both SI (IFCC) and NGSP/DCCT units.

This review provides a brief summary of the reaction of glucose with Hb and the possible techniques available for analyzing HbA1c. It also reviews the quality control, requirements of the target limits and considers issues related to the units and cutoff limits for HbA1c in relation to the recommendations of IFCC and ADA.

GLYCATION OF HB

Glucose is the major soluble carbohydrate and it can combine with different protein molecules in blood and tissues depending on the glucose concentration. In blood, Hb is one of these proteins and the -chain of Hb is the main target of glucose[6]. Glucose is initially non-enzymatically bound to the amino acid valine on the -chain of Hb via the formation of a reversible aldimine moiety, which then becomes rearranged into the irreversible ketoamine form. The structure of this ketoamine is similar to that of fructosamine[6]. This phenomenon occurs during the whole lifetime of erythrocytes (120 d) and thus the content of the ketoamine in Hb correlates with the age of the erythrocytes and the value of HbA1c normally represents a mean value reflecting a time period starting from about three months before sampling.

By analyzing Hb components in blood, Trivelli et al[7] demonstrated that the small fraction A1c was glycated more than the other small fractions A1a and A1b and that the amounts of this A1c fraction were clearly higher in the blood of diabetic patients than in the normal subjects. The fractionation of Hb by a Mono S™ HR5/5 cation exchange column from Pharmacia with 0.01 mol/L malonate buffer for normal and diabetic human blood samples is illustrated in Figure 1[20].

METHODS OF HBA1C ASSAYS

The first methods for the HbA1c measurements were simple electrophoretic or ion-exchange mini-column chromatography assays, but they were rather soon replaced by many automated techniques, e.g., different liquid chromatography techniques [ion-exchange chromatography (HPLC or FPLC) or affinity chromatography][8-12,20-22] as seen in Tables 1 and 2. The NGSP/DCCT organizations in United States and Canada selected the automated Bio-Rex 70 liquid chromatography assay as the reference method for the HbA1c measurement[9]. It should be stressed that this equipment has not been available for many years. However, in addition to the NGSP/DCCT system, there were other different assay techniques used for standardization and these had their own standards, e.g., in Japan (NGH) and in Sweden (FPLC with a Mono S column). These methods could be compared to each other and to the new IFCC procedure by using the appropriate master equations[23].

In 1988, an automated immunoassay method for the epitope assays of proteins was developed[24] and then utilized in the assay of HbA1c. These immunoassay or enzymatic methods have replaced the chromatography methods previously used in clinical laboratories so that today only about 20% to 40% of methods for HbA1c are based on liquid chromatographic techniques[16,25]. The numbers of immunoassays for HbA1c based on many different principles are continuously increasing but these techniques usually have a higher inter-laboratory CV% than can be attained with liquid chromatography or enzymatic methods[16,22,25-27]; some of the most important immunoassay methods are listed in Tables 1 and 2.

STANDARDIZATION OF HbA1C

As noted before, there were marked inter-laboratory differences in the quality assurance results before standardization procedures were adopted[8-12]. It was also reported that by incorporating an extra sample, against which the primary results could be recalculated, and significantly smaller CV% values could be obtained[9,16,17].

After the appearance of the IFCC WG for HbA1c, a reference system was soon organized in 1996[13], with the goal of achieving standardization of HbA1c analysis; this formed the basis for the worldwide reference laboratory network to help clinical laboratories in their HbA1c measurements. The reference preparations which represented the primary and secondary standards of HbA1c and HbA0 were produced in 1998[14] followed in 2002 by the reference method for the specific measurement of HbA1c[15]. The final measurements in the reference method were based on the assay of N1-deoxyfructosyl-hemoglobin[28]. The assay consisted of a primary fractionation of the sample by affinity chromatography followed by analysis utilizing either HPLC/electrospray mass spectrometry or HPLC/capillary electrophoresis. The primary and secondary standards and the reference methods and the guidelines about names and units were then introduced to be adopted worldwide by all manufacturers making equipment for analysis of HbA1c and also for clinical laboratories[13,23,25,29].

The standardization protocols about HbA1c have been published by many societies of laboratory medicine in their own languages, e.g., ADA in United States[18,29], DGKL in Germany[30], NEQAS in United Kingdom[31], SiBioC in Italy[32], EQUALIS in Sweden[33], Finnish Society of Clinical Chemistry (FSCC) in Finland[34], etc. With respect to the standardization procedures, it should be noted that according to Weykamp et al[35] the units and standardization protocols for the NGSP/DCCT and IFCC procedures are different for NGSP/DCCT and IFCC results and thus the reference values and quality control requirements cannot be of the same size.

ANALYTICAL PROBLEMS OF HBA1C ASSAYS

The typical problems[36] which interfere with the HbA1c assays, are attributable to hyperbilirubinemia, hypertriglyceridemia, leukocytosis and many Hb variants. In addition, certain physiological and pathological characteristics such as gestational stage, age, race, pre-symptomatic type 1 diabetes, malaria, iron deficiency, bleedings, transfusions, splenectomy, kidney failure, alcohol abuse, and some drugs may affect the HbA1c results. The HbA1c results may be too high in some cases and too low in others (hemolysis, pregnancy).

If erythrocytes have a short life-time, e.g., in hemolytic anemias, this may decrease the HbA1c results. Some abnormal forms of Hb cause erroneous results (HbF, HbS, HbC, HbD, HbE, etc.) depending on the assay type[37-39]. Furthermore, iron deficiency anemia as well as iron deficiency without anemia may induce elevated HbA1c values compared with controls even though blood glucose levels are normal[39].

Many of these errors can be very difficult to detect, especially when using immunochemical assays. With some systems such as with liquid chromatography, the erroneous results can be visualized in the chromatograms, and these assays are commonly used for comparison[23,35-40]. However, errors can also be due to the problems in the action or response of insulin which are not directly related to the methods being used to assay HbA1c[36].

UNITS OF HBA1C IN LABORATORY PRACTICE

In 2009, prior to the use of the new IFCC system with the accepted name and unit for HbA1c, questionnaires about which units should be used were sent to the European societies of laboratory medicine and some other societies outside of Europe (mainly in clinical chemistry)[40]. Germany was the first country to adopt mmol/mol exclusively in the daily laboratory practice (Deutsche Vereinte Gesellschaft für Klinische Chemie und Laboratoriums-medizin); this was inaugurated at the start of 1.1.2010. Germany was followed later in 2011 by The Netherlands, Sweden and the United Kingdom. Then gradually the number of mmol/mol reports of HbA1c increased up to 13 in 2014, representing 25% of all replies from 51 queries. During these years, there was also an increase in the numbers of laboratories (from 9 to 12) reporting HbA1c values in parallel units, i.e., in both % and mmol/mol. By 2014, a minority of responders (24%) stated that they were using only the mmol/mol whereas nearly every other respondent (49%) stated that values were being expressed not only in mmol/mol but also as % values (Table 3). However, by the end of 2014, ten societies had not responded to the questionnaire, although the e-mail, telefax and mail addresses were taken from the annual catalogues of IFCC.

On the other hand, in the period from 2010 to 2015, the users of % values in the HbA1c surveys conducted by Labquality Ltd.[41] have gradually but significantly decreased (from 54% to 40%).

The Finnish experience is an example of the difficulty to obtain acceptance of the molar unit recommended by IFCC for expressing HbA1c. In 2009, the FSCC recommended that clinical laboratories should report the HbA1c results in parallel, i.e., as mmol/mol and % values according to the recommendations[23]. In 2011, the FSCC proposed that the HbA1c values should be provided only in mmol/mol. Nonetheless, the issue has not fully resolved due to the opposition from physicians treating diabetics. Thus the following compromise was agreed: In the future, the laboratories will report the HbA1c values in parallel units, i.e., as mmol/mol and % although the measurements in many laboratories have originally been calculated in mmol/mol[40].

In Finland, one out of five university hospital districts (Pirkanmaa and Kanta-Häme) stopped reporting the % values for HbA1c from April 2014; in that hospital the HbA1c values are only being reported in mmol/mol[40]. It is possible that some other Finnish hospital districts will come to the same decision in the near future. In September 2014, the board of FSCC established a new working group to discuss with physicians and with non-laboratory societies and as well as with other interested parties about the issues related to the determination of HbA1c. During the summer of 2015, FSCC decided to recommend that from 1.1.2016 the HbA1c values should no longer be expressed as % units; hopefully this recommendation will prove acceptable.

QUALITY ASSYRANCE AND TARGET LIMITS OF HbA1C

The QC surveys utilize common statistical methods such as the coefficient of variation (CV%)[16-18]. On the basis of these and the standard deviation (SD) values, the new target limits have been calculated for the QC-surveys of HbA1c around the mean and the reference values expressed as either % or mmol/mol values. This means that within the range mean ± 2SD 95.6% of all results are within the acceptable limits and these correspond to the recent findings of Hyltoft Petersen and Lee[42]. Furthermore, the new report of Nielsen et al[43] about the value of HbA1c in the classification of diabetes highlights the importance of the exact and precise measurement of HbA1c.

As mentioned earlier, there was a remarkably high variation in the values being obtained with the different methods being used in laboratories all around the world[8-12]. Weykamp et al[11] were one of the first groups who described the dramatic improvement that could be obtained in the CV% of the HbA1c results of either normal or diabetic subjects when the primary results were recalculated by incorporating an extra sample with a known HbA1c value. They had reviewed the corrected results from 110 laboratories using 21 different methods.

In Finland, the glycohemoglobin quality assurance surveys of Labquality started in 1986 with two fresh native EDTA-blood samples, one at the normal level and the second at the diabetic level, which were sent in each survey to all participants[41]. The mean values of the surveys were used as the target values and the acceptable ranges were ± 10% around the mean values. Figure 2A displays the significant improvement in the HbA1c % values (r = 0.896) from 1994 to 2015 expressed as CV% values, they correspond well to the earlier publications[16,17]. Similarly in Figure 2B from 2009 when the results were reported in mmol/mol units, a further improvement in the CV%s was observed (for results expressed as mmol/mol) (r = 0.642). The mean ± SD values and CV% have been calculated from the surveys of Labquality.

In the most recent survey (2015) the range mean ± 2SD indicates that 95% of all results are within these limits[9,44,45]. For example, in December 2014, at the HbA1c level of 6.77%, near the cut-off limit of ADA of 6.5%[18], the SD value was 0.21 for % units and that for mmol/mol units at the level of 50.8 mmol/mol the corresponding SD value was 2.1 mmol/mol. Thus the calculated acceptable limits with the old range of mean HbA1c 6.77% ± 10% were from 6.09% to 7.48% and with the new narrower limit dating from 1.1.2015 the limits with mean ± 6% were from 6.36% to 7.18%. The latter range corresponds well with that of 6.77% ± 0.21% with the CV of 3.26% for the survey (0.21/6.77 × 100%). If one takes the mean ± 2SD of all results, then the range would be 6.35% to 7.19% in that survey and these would be clearly inside the old range of 10% and similar to the new limits from Labquality. Correspondingly in the same survey of December 2014 for mmol/mol values, the mean HbA1c value was 50.8 ± 2.1 mmol/mol and the range for mean ± 2SD extended from 46.6 to 55.0 mmol/mol being very close to the new mmol/mol limit from Labquality as mean ± 8%, i.e., from 46.7 to 54.9 mmol/mol with the CV% of 4.13% (2.1/50.8 × 100%). In the future, the target limits for HbA1c of Labquality may be made more demanding. The latest Finnish results that in 2015 the CV% values for HbA1c have reached a sufficiently low level (Figure 2) to be comparable to others in the literature both for % and mmol/mol results[9,16-18,46,47].

From the survey conducted in 2002, the HbA1c values of the EDTA blood samples for the % results were also analyzed by the European Reference Laboratory for Glycohemoglobin (ERLGH)[48]. It was found that there was an almost perfect correlation between the mean HbA1c % values of Labquality and the values HbA1c % values of ERLGH (r = 0.997). The same phenomenon was also seen for the mmol/mol results from 2010 when comparing the mean values of Labquality and those of ERLGH (r = 0.973). Since the mean values of the surveys conducted by Labquality are in practice around the same size as the ERLGH values, thus one may be utilized as target values for mmol/mol The findings correspond well with the earlier reports[43,44] despite the fact that there are differences in both units and standardization programs between the NGSP/DCCT and IFCC procedures[35]. The new target limits of the HbA1c results in 2015 for % results are ± 6% and for mmol/mol results ± 8% around the target value and agree well to those reported by Little et al[17], Weykamp et al[49] and Lindblad and Nordin[50]. In addition, on the basis of CAP surveys, Little et al[17] have reported that at the normal HbA1c levels (6%-7%) the target value described in % units and with ± 6% around the reference value may be good enough for the diagnosis and follow-up of the treatment of diabetes. Lindblad and Nordin[50] reported that when near to the critical level of HbA1c of 48 mmol/mol, the maximum allowable difference from the target value should be less than 3.5 mmol/mol, which corresponds to a CV% of 7.3%.

The Uppsala Meeting for Quality Specifications in Clinical Laboratories in early 1990s had one session devoted to the quality of HbA1c measurements. The Danish clinical chemists[51,52] proposed that a change of 1.0% from the measured % HbA1c value might suggest that the treatment could be necessary whereas a change of 2.0% would demand that treatment should be initiated. This degree of accuracy is necessary that clinicians could feel confident with their decisions to initiate what could well be life-time therapy. These findings are in accordance with the earlier reports[8,16,20] and with our present findings. On the other hand, recently in 2015, Weykamp et al[45] reported that when using mmol/mol units, the calculated total allowable error could be 4.2 mmol/mol. The HbA1c results of current Labquality surveys conducted in 2015, i.e., the mean values of 50.8 ± 2.1 mmol/mol are within this value.

On the other hand, a single laboratory should be able to perform HbA1c % measurements with a total analytical variation (within and between series) from between 1.4% to 3.0% when using liquid chromatography[10,20,21,26], this variation due to the analytical procedure is about 4.5 times lower than the measured biological variation of this parameter[42,53]. Furthermore, also the differences (errors) between frequent measurements should be as small as possible to ensure reliable follow-up of the treatment of those patients in a stable diabetic state[16,17,23,31]. After the proposal from ADA[18] to use a diagnostic cutoff limit of HbA1c for diabetes, this change demanded that the accuracy of the analytical performance of HbA1c had to be better than before the setting of this fixed limit, a fact that has also been criticized.

REFERENCE VALUES OF HBA1C

The commonly used reference intervals for HbA1c in % DCCT/NGHP units have been from 4.0% to 6.0%[16,17,20,49,54]; the corresponding values for mmol/mol have been from 20 to 42 mmol/mol[18,50] as presented in Table 4. The exact values somewhat vary around those in the table but nonetheless are rather close to those published earlier[16,17,20,49,50,54]. In addition, the commonly applied reference values of HbA1c in published reports expressed as per cent and mmmol/mol are normally stated in the laboratory manuals all around the world as are also the important limits for diagnosis and treatment of diabetes; these manuals also provide the master equations to convert between the system in use and the procedure of IFCC[23,30-35].

HBA1C AND PLASMA GLUCOSE

The commonly used limit of plasma glucose of 7.0 mmol/L[8,9,17-19] has been widely accepted when there is a need to diagnose diabetes or to monitor the treatment of the patients. Many physicians treating diabetics like to compare the actual plasma glucose values to the HbA1c values and consider that they are assessing some kind of balance of diabetes. However, it was not until 2002 when Rohlfing et al[55] collected enough published data to be able to devise an equation describing the well-established ratio between blood mean HbA1c and plasma glucose, i.e., (PG/HbA1c): Plasma glucose (mmol/L) = [1.98 × HbA1c(%)] - 4.29 (n = 1439, r = 0.82). However, when inspecting their proposal, it is quite evident that at the same HbA1c value, e.g., 8.0%, the plasma glucose concentration may maximally vary from 6 to 15 mmol/L and thus the results are not precise, especially at the higher blood HbA1c levels[49]. In addition, in subjects with impaired glucose tolerance (plasma glucose value from 5.6 to 6.4 mmol/L or blood HbA1c from 5.7 to 6.4 mmol/mol) the 2 h oral glucose tolerance test may be as good or better at revealing both the insulin sensitivity and disturbance in glucose metabolism than can be achieved with a single value of fasting plasma glucose[56-60]. These are the reasons why many laboratories do not calculate the mean PG/HbA1c ratios but report glucose and HbA1c results separately[41]. It has also argued that in fulminant (acute) type Ⅰ diabetes, the specificity of the analysis of HbA1c may be doubtful low when compared to the analysis of glycated albumin with a significantly higher specificity[59]. This may have a significant role to make a selection of the analyze type in diagnosing of acute diabetes or prediabetes. It is also important that plasma glucose and blood HbA1c levels should be assessed in an accredited laboratory if the values are to be used for diagnosis or screening of diabetes[18,19,23,58-62].

HBA1C AND SOME OTHER ASPECTS

Finally, it has been pointed out that in both men and women elevated levels of blood HbA1c increase the risk of developing cardiovascular disease[63]. In addition, as described in the chapter of analytical problems, many other diseases may affect the HbA1c results and cause difficulties for the clinicians. Some diseases other than cardiovascular are also associated with HbA1c[62,63], e.g., diseases altering iron metabolism.

The point-of care instruments (POC) are being continuously introduced for the analysis of HAb1c[25,64] but there are still many issues associated with their use[64-66]: (1) they have not been universally recommended for the diagnosis of diabetes according to the guidelines[17-19,30]; (2) many POC users are not participating in the quality control programs of their home country; and (3) the POC analyses are difficult to standardize. In the future, in conjunction with the further development of these methods, especially with the adoption of reliable quality systems, POC analyses may achieve real breakthroughs[25].

CONCLUSION

This assessment of HbA1c analytical procedures and values indicates that a considerable improvement has occurred during the past 30 years with respect to both their the precision and accuracy and these improvements are still on-going as reflected in reduced assay CV% values. The immunoassay techniques have replaced many chromatographic procedures during this time period.

During 2014/2015, the reports from quality assurance systems have confirmed the marked improvement for the quality of HbA1c measurements irrespective of whether the results have been expressed in % or mmol/mol units. The acceptance of the mmol/mol system recommended by IFCC for HbA1c and the new unit and reference ranges are only becoming slowly accepted outside of Europe where it seems that the parallel reporting for HbA1c will continue. The use of a diagnostic cutoff limit for the HbA1c value is still not finalized.

The reference analyses from the mean values of the survey results may be used as target values in both % and mmol/mol units assuming that the number of the participating laboratories is high enough to be statistically satisfactory.

The authors also hope that the use of the mmol/mol unit for HbA1c can gain worldwide acceptance as this would make it much easier to compare results from different studies and remove the possibility of confusion when units are converted from one form to the other.

ACKNOWLEDGMENTS

The authors cordially than all 41 societies which sent replies to the queries in 2009 to 2014. We also thank Labqualty Ltd. for the survey results of HbA1c for our disposal and Dr. Ewen McDonald for correcting the English language.

REFERENCES

1	Fisher E. Über Isoglucosamin. Chemische Berichte 1886; 19: 1920-1924

2	Kunkel HG, Wallenius G. New hemoglobin in normal adult blood. Science 1955; 122: 288 [PMID: 13246634]

3	Schnek AG, Schroeder WA. A chromatographic study of the minor components of normal adult human hemoglobin as isolated by chromatography and starch block electrophoresis. J Amer Chem Soc 1961; 83: 1472-1478 [DOI: 10.1021/ja01467a046]

4	Holmquist WR, Schroeder WA. A new N-terminal blocking group involving a Schiff base in hemoglobin AIc. Biochemistry 1966; 5: 2489-2503 [PMID: 5968556]

5	Rahbar S. An abnormal hemoglobin in red cells of diabetics. Clin Chim Acta 1968; 22: 296-298 [PMID: 5687098]

6	Shapiro R, McManus MJ, Zalut C, Bunn HF. Sites of nonenzymatic glycosylation of human hemoglobin A. J Biol Chem 1980; 255: 3120-3127 [PMID: 7358733]

7	Trivelli LA, Ranney HM, Lai HT. Hemoglobin components in patients with diabetes mellitus. N Engl J Med 1971; 284: 353-357 [PMID: 5539916]

8	Little RR, Wiedmeyer HM, England JD, Wilke AL, Rohlfing CL, Wians FH, Jacobson JM, Zellmer V, Goldstein DE. Interlaboratory standardization of measurements of glycohemoglobins. Clin Chem 1992; 38: 2472-2478 [PMID: 1458587]

9	Little RR, Rohlfing CL, Wiedmeyer HM, Myers GL, Sacks DB, Goldstein DE. The national glycohemoglobin standardization program: a five-year progress report. Clin Chem 2001; 47: 1985-1992 [PMID: 11673367]

10	Penttilä I, Gävert J, Julkunen A, Tantanen T. Quality control and quality requirements for the measurement of glycated hemoglobin. Ups J Med Sci 1990; 95: 291-297 [PMID: 2100408]

11	Weykamp CW, Penders TJ, Muskiet FA, van der Slik W. Effect of calibration on dispersion of glycohemoglobin values determined by 111 laboratories using 21 methods. Clin Chem 1994; 40: 138-144 [PMID: 8287521]

12	Jensen ON, de Fine Olivarius N, Petersen PH, Klitgaard NA, Blaabjerg O, Hørder M. Discrepancy in HbA1c measurements performed at different local laboratories and at a selected central reference laboratory. Ups J Med Sci 1993; 98: 275-282 [PMID: 7974856]

13	Hoelzel W, Miedema K. Development of a reference system for the international standardization of HbA1c/glycohemoglobin determinations. J Int Fed Clin Chem 1996; 8: 62-64, 66-67 [PMID: 10163516]

14	Finke A, Kobold U, Hoelzel W, Weykamp C, Miedema K, Jeppsson JO. Preparation of a candidate primary reference material for the international standardisation of HbA1c determinations. Clin Chem Lab Med 1998; 36: 299-308 [PMID: 9676387]

15	Jeppsson JO, Kobold U, Barr J, Finke A, Hoelzel W, Hoshino T, Miedema K, Mosca A, Mauri P, Paroni R, Thienpont L, Umemoto M, Weykamp C. Approved IFCC reference method for the measurement of HbA1c in human blood. Clin Chem Lab Med 2002; 40: 78-89 [PMID: 11916276]

16	Penttilä IM, Halonen T, Punnonen K, Tiikkainen U. Best use of the recommended IFCC reference method, material and values in HbA1C analyses. Scand J Clin Lab Invest 2005; 65: 453-462 [PMID: 16179278]

17	Little RR, Rohlfing CL, Sacks DB. Status of hemoglobin A1c measurement and goals for improvement: from chaos to order for improving diabetes care. Clin Chem 2011; 57: 205-214 [PMID: 21148304 DOI: 10.1373/clinchem.2010.148841]

18	American Diabetes Association. Diagnosis and classification of diabetes mellitus. Diabetes Care 2010; 33 (Suppl 1): S62-S69 [DOI: 10.2337/dc10-S062]

19	Hanas R, John WG. 2013 Update on the worldwide standardization of the hemoglobin A1c measurement. Pediatr Diabetes 2014; 15: e1-e2 [PMID: 23691962 DOI: 10.1111/pedi.1204]

20	Penttilä IM, Gävert J, Julkunen A, Rantanen T. Quality specifications for glycated hemoglobin A1c. Ups J Med Sci 1993; 98: 395-399 [PMID: 7974871]

21	Stenman UH, Pesonen K, Ylinen K, Huhtala ML, Teramo K. Rapid chromatographic quantitation of glycosylated haemoglobins. J Chromatogr 1984; 297: 327-332 [PMID: 6490764]

22	Little RR, Roberts WL. A review of variant hemoglobins interfering with hemoglobin A1c measurement. J Diabetes Sci Technol 2009; 3: 446-451 [PMID: 20144281]

23	Hoelzel W, Weykamp C, Jeppsson JO, Miedema K, Barr JR, Goodall I, Hoshino T, John WG, Kobold U, Little R, Mosca A, Mauri P, Paroni R, Susanto F, Takei I, Thienpont L, Umemoto M, Wiedmeyer HM. IFCC reference system for measurement of hemoglobin A1c in human blood and the national standardization schemes in the United States, Japan, and Sweden: a method-comparison study. Clin Chem 2004; 50: 166-174 [PMID: 14709644]

24	Fiore M, Mitchell J, Doan T, Nelson R, Winter G, Grandone C, Zeng K, Haraden R, Smith J, Harris K. The Abbott IMx automated benchtop immunochemistry analyzer system. Clin Chem 1988; 34: 1726-1732 [PMID: 2458201]

25	Weykamp C, John WG, Mosca A. A review of the challenge in measuring hemoglobin A1c. J Diabetes Sci Technol 2009; 3: 439-445 [PMID: 20144280]

26	Mosca A, Goodall I, Hoshino T, Jeppsson JO, John WG, Little RR, Miedema K, Myers GL, Reinauer H, Sacks DB, Weykamp CW. Global standardization of glycated hemoglobin measurement: the position of the IFCC Working Group. Clin Chem Lab Med 2007; 45: 1077-1080 [PMID: 17867998]

27	Penttilä I, Penttilä K, Halonen T, Pulkki K, Törrönen J, Rauramaa R. Adaptation of the Diazyme Direct Enzymatic HbA1c Assay for a microplate reader at room temperature. Clin Chem Lab Med 2011; 49: 1221-1223 [PMID: 21574882 DOI: 10.1515/CCLM.2011.186]

28	Nordin G, Dybkaer R. Recommendation for term and measurement unit for “HbA1c”. Clin Chem Lab Med 2007; 45: 1081-1082 [PMID: 17867999]

29	Nathan DM, Buse JB, Davidson MB, Ferrannini E, Holman RR, Sherwin R, Zinman B. Medical management of hyperglycaemia in type 2 diabetes mellitus: a consensus algorithm for the initiation and adjustment of therapy: a consensus statement from the American Diabetes Association and the European Association for the Study of Diabetes. Diabetologia 2009; 52: 17-30 [PMID: 18941734 DOI: 10.1007/s00125-008-1157-y]

30	Deutsche Vereinte Gesellschaft für Klinische Chemie und Laboratoriums-medizin. Stellungnahme der Deutschen Diabetes Gesellschaft, diabetes und des Kompetenznetzes Diabetes mellitus zur Verwendung des HbA1c-Wertes als Biomarker zur Diabetesdiagnose. Available from: URL: www.diabetesde.org/

31	The Association for Clinical Biochemistry & Laboratory Medicine. United Kingdom/Birmingham Quality, HbA1c - UK NEQAS Guide to Dual reporting. Available from: URL: www.birminhamquality.org/

32	Mosca A, Branca MT, Carta M, Genna ML, Giorda CB, Ghidelli R, Ghislandi G, Iafusco D, Lapolla A, Lombardi VB, Lovagnini Scher CA, Marra M, Medea G, Meschi F, Pizzini A, Rossi F, Scalpone R, Tofini G, Trovati M, Zaninotto M. Recommendations for the implementation of international standardization of glycated hemoglobin in Italy. Clin Chem Lab Med 2010; 48: 623-626 [PMID: 20201745 DOI: 10.1515/CCLM.2010.140]

33	Landin-Olsson M, Jeppsson JO, Nordin G. [HbA1c--new standardization introduced in Sweden. The new unit is mmol/mol]. Lakartidningen 2010; 107: 3282-3285 [PMID: 21294350]

34	Finnish Society of Clinical Chemistry. Available from: URL: www.skky.fi/english

35	Weykamp CW, Mosca A, Gillery P, Panteghini M. The analytical goals for hemoglobin A(1c) measurement in IFCC units and National Glycohemoglobin Standardization Program Units are different. Clin Chem 2011; 57: 1204-1206 [PMID: 21571810 DOI: 10.1373/clinchem.2011.162719]

36	Lapolla A, Mosca A, Fedele D. The general use of glycated haemoglobin for the diagnosis of diabetes and other categories of glucose intolerance: still a long way to go. Nutr Metab Cardiovasc Dis 2011; 21: 467-475 [PMID: 21641782 DOI: 10.1016/j.numecd.20 11.02.006]

37	Mongia SK, Little RR, Rohlfing CL, Hanson S, Roberts RF, Owen WE, D’Costa MA, Reyes CA, Luzzi VI, Roberts WL. Effects of hemoglobin C and S traits on the results of 14 commercial glycated hemoglobin assays. Am J Clin Pathol 2008; 130: 136-140 [PMID: 18550483 DOI: 10.1309/1YU0D34VJKNUCGT1]

38	Puukka R, Puukka M. Effect of hemoglobin F on measurements of hemoglobin A1c with physicians’ office analyzers. Clin Chem 1994; 40: 342-343 [PMID: 7508831]

39	English E, Idris I, Smith G, Dhatariya K, Kilpatrick ES, John WG. The effect of anaemia and abnormalities of erythrocyte indices on HbA1c analysis: a systematic review. Diabetologia 2015; 58: 1409-1421 [PMID: 25994072 DOI: 10.1007/s00125-015-3599-3]

40	Penttilä I, Penttilä K, Holm P, Laitinen H, Rauramaa R. Hemoglobin A1c reported in units and diagnostic cut-offs in relation to the international recommendations. Clin Chem Lab Med 2015; 53: e215-e217 [PMID: 26136301 DOI: 10.1515/cclm-2015-0247]

41	Labquality Ltd. Finland. Available from: URL: www.labquality.fi

42	Hyltoft Petersen P, Klee GG. Influence of analytical bias and imprecision on the number of false positive results using Guideline-Driven Medical Decision Limits. Clin Chim Acta 2014; 430: 1-8 [PMID: 24362233 DOI: 10.1016/j.cca.2013.12.014]

43	Nielsen AA, Petersen PH, Green A, Christensen C, Christensen H, Brandslund I. Changing from glucose to HbA1c for diabetes diagnosis: predictive values of one test and importance of analytical bias and imprecision. Clin Chem Lab Med 2014; 52: 1069-1077 [PMID: 24659606 DOI: 10.1515/cclm-2013-0337]

44	John WG, Little R, Sacks DB, Weykamp C, Lenters-Westra E, Hornsby T, Zhao Z, Siebelder C, Tennill A, English E. Multicentre evaluation of the Premier Hb9210 HbA1c analyser. Clin Chem Lab Med 2015; 53: 319-327 [PMID: 25274956 DOI: 10.1515/cclm-2014-0589]

45	Weykamp C, John G, Gillery P, English E, Ji L, Lenters-Westra E, Little RR, Roglic G, Sacks DB, Takei I. Investigation of 2 models to set and evaluate quality targets for hb a1c: biological variation and sigma-metrics. Clin Chem 2015; 61: 752-759 [PMID: 25737535 DOI: 10.1373/clinchem.2014.235333]

46	Sacks DB, Arnold M, Bakris GL, Bruns DE, Horvath AR, Kirkman MS, Lernmark A, Metzger BE, Nathan DM. Guidelines and recommendations for laboratory analysis in the diagnosis and management of diabetes mellitus. Clin Chem 2011; 57: e1-e47 [PMID: 21617152 DOI: 10.1373/clinchem.2010.161598]

47	Lenters-Westra E, Weykamp C, Schindhelm RK, Siebelder C, Bilo HJ, Slingerland RJ. One in five laboratories using various hemoglobin A1c methods do not meet the criteria for optimal diabetes care management. Diabetes Technol Ther 2011; 13: 429-433 [PMID: 21355726 DOI: 10.1089/dia.2010.0148]

48	Weykamp C, Lenters-Westra E, van der Vuurst H, Slingerland R, Siebelder C, Visser-Dekkers W. Evaluation of the Menarini/ARKRAY ADAMS A1c HA-8180V analyser for HbA1c. Clin Chem Lab Med 2011; 49: 647-651 [PMID: 21323625 DOI: 10.1515/CCLM.2011.096]

49	Weykamp CW, Penders TJ, Miedema K, Muskiet FA, van der Slik W. Standardization of glycohemoglobin results and reference values in whole blood studied in 103 laboratories using 20 methods. Clin Chem 1995; 41: 82-86 [PMID: 7813086]

50	Lindblad B, Nordin G. External quality assessment of HbA1c and its effect on comparison between Swedish pediatric diabetes clinics. Experiences from the Swedish pediatric diabetes quality register (Swediabkids) and Equalis. Clin Chem Lab Med 2013; 51: 2045-2052 [PMID: 23773967]

51	Hyltoft Petersen P, Lytken Larsen M, Fraser CG. The quality needed for measuring glycated haemoglobin. An application. Ups J Med Sci 1990; 95: 185-190 [PMID: 2100393]

52	Larsen ML, Petersen PH, Fraser CG. Quality specifications for haemoglobin A1c assays in the monitoring of diabetes. Ups J Med Sci 1993; 98: 335-338 [PMID: 7974864]

53	Lenters-Westra E, Røraas T, Schindhelm RK, Slingerland RJ, Sandberg S. Biological variation of hemoglobin A1c: consequences for diagnosing diabetes mellitus. Clin Chem 2014; 60: 1570-1572 [PMID: 25248570 DOI: 10.1373/clinchem.2014.227983]

54	Braga F, Dolci A, Montagnana M, Pagani F, Paleari R, Guidi GC, Mosca A, Panteghini M. Revaluation of biological variation of glycated hemoglobin (HbA(1c)) using an accurately designed protocol and an assay traceable to the IFCC reference system. Clin Chim Acta 2011; 412: 1412-1416 [PMID: 21524644 DOI: 10.1016/j.cca.2011.04.014]

55	Rohlfing CL, Wiedmeyer HM, Little RR, England JD, Tennill A, Goldstein DE. Defining the relationship between plasma glucose and HbA(1c): analysis of glucose profiles and HbA(1c) in the Diabetes Control and Complications Trial. Diabetes Care 2002; 25: 275-278 [PMID: 11815495]

56	Matsuda M, DeFronzo RA. Insulin sensitivity indices obtained from oral glucose tolerance testing: comparison with the euglycemic insulin clamp. Diabetes Care 1999; 22: 1462-1470 [PMID: 10480510]

57	Cavagnolli G, Comerlato J, Comerlato C, Renz PB, Gross JL, Camargo JL. HbA(1c) measurement for the diagnosis of diabetes: is it enough? Diabet Med 2011; 28: 31-35 [PMID: 21210540]

58	Gordon-Dseagu VL, Mindell JS, Steptoe A, Moody A, Wardle J, Demakakos P, Shelton NJ. Impaired glucose metabolism among those with and without diagnosed diabetes and mortality: a cohort study using Health Survey for England data. PLoS One 2015; 10: e0119882 [PMID: 25785731 DOI: 10.1371/journal.pone.0119882]

59	Koga M, Murai J, Saito H, Mukai M, Kasayama S. Serum glycated albumin, but not glycated hemoglobin, is low in relation to glycemia in men with hypertriglyceridemia. J Diabetes Investig 2010; 1: 202-207 [PMID: 24843433 DOI: 10.1111/j.2040-1124.2010.00049.x]

60	Fujibayashi K, Yokokawa H, Gunji T, Sasabe N, Okumura M, Iijima K, Haniu T, Hisaoka T, Fukuda H. Utility of 75-g oral glucose tolerance test results and hemoglobin a1c values for predicting the incidence of diabetes mellitus among middle-aged Japanese men -a large-scale retrospective cohort study performed at a single hospital. Intern Med 2015; 54: 717-723 [PMID: 25832931 DOI: 10.2169/internalmedicine.54.2839]

61	American Diabetes Association. Standards of medical care in diabetes - 2011. Diabetes Care 2011; 34 (Suppl 1): S11-S61 [DOI: 10.2337/cd11-5011]

62	Weykamp C. HbA1c: a review of analytical and clinical aspects. Ann Lab Med 2013; 33: 393-400 [PMID: 24205486 DOI: 10.3343/alm.2013.33.6.393]

63	Gerstein HC, Swedberg K, Carlsson J, McMurray JJ, Michelson EL, Olofsson B, Pfeffer MA, Yusuf S. The hemoglobin A1c level as a progressive risk factor for cardiovascular death, hospitalization for heart failure, or death in patients with chronic heart failure: an analysis of the Candesartan in Heart failure: Assessment of Reduction in Mortality and Morbidity (CHARM) program. Arch Intern Med 2008; 168: 1699-1704 [PMID: 18695086 DOI: 10.1001/archinte.168.15.1699]

64	Lenters-Westra E, Slingerland RJ. Six of eight hemoglobin A1c point-of-care instruments do not meet the general accepted analytical performance criteria. Clin Chem 2010; 56: 44-52 [PMID: 19926777 DOI: 10.1373/clinchem.2009.130641]

65	Al-Ansary L, Farmer A, Hirst J, Roberts N, Glasziou P, Perera R, Price CP. Point-of-care testing for Hb A1c in the management of diabetes: a systematic review and metaanalysis. Clin Chem 2011; 57: 568-576 [PMID: 21368238 DOI: 10.1373/clinchem.2010.157586]

66	Lenters-Westra E, Slingerland RJ. Hemoglobin A1c point-of-care assays; a new world with a lot of consequences! J Diabetes Sci Technol 2009; 3: 418-423 [PMID: 20144277]

Figure Legends

Figure 1  Fractionation of blood hemoglobin by an automated Mono S™ HR5/5 cation exchange column from Pharmacia with malonate buffer (0.01 mol/L, pH 5.7, 11 min, room temperature) for normal and diabetic human blood samples is illustrated as a chromatogram (Penttilä et al[10]).

Figure 2  The total variations of the glycohemoglobin results have been presented both in % and in mmol/mol units from the surveys of Labquallity Ltd.  from 1994 to 2015. In the part A have been presented the total variations of the glycohemoglobin in % units and in the part B the total variations of HbA1c in mmol/mol units.

Footnotes

Conflict-of-interest statement: The authors indicate that none of the authors have no potential conflicts of interest related to the manuscript and that they have had not any financial support. Ilkka Penttilä, MD, PhD, Emeritus Professor is the survey expert for glycohemoglobin of Labquality Ltd, Harri Laitinen is PhM and the International relationship manager of Labquality Ltd., and Päivi Ranta is specialist of clinical biochemistry for HbA1c surveys of Labquality Ltd, Päivi Holm is the specialist in clinical biochemistry for mpere university hospital, Karri Penttilä, MD, PhD is specialist in internal medicine and hematology acting as the clinician for this report, and Rainer Rauramaa, MD, PhD professor as the chief of the Kuopio Research Institute of Exercise Medicine. The numerical data from Labquality Ltd. has been received with permission for this report.

Open-Access: This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

Peer-review started: September 22, 2015

First decision: October 30, 2015

Article in press: March 25, 2016

P- Reviewer: Gómez-Sáez JM, Haidara MAA, Moeintaghavi A    S- Editor: Ji FF    L- Editor: A    E- Editor: Jiao XK

# Text boxes

Table 1  The methods commonly used for analyzing hemoglobin A1c according to the surveys conducted by Labquality Ltd. in 1994

| Numerical summary from a diabetic EDTA-blood sample 1/1994 |
| --- |
| Analyte/method group | Mean | SD | CV% | Number |
| GHbA1c % |  |  |  |  |
| Ion exchange Ciba Corning | 8.7 | 0.4 | 4.3 | 4 |
| Ion exchange Diamat Biorad | 9.7 | 0.8 | 8.4 | 15 |
| Ion exchange Kyoto Daichii | 8.4 | 0.3 | 3.4 | 3 |
| Ion exchange Pharmacia MonoS | 9.0 | 0.8 | 8.7 | 17 |
| Ion exchange Shimadzu | 10.7 | 0.1 | 0.7 | 2 |
| Ion exchange others | 9.5 | 0.8 | 7.9 | 5 |
| Ion exchange minicolumn Biorad | 9.7 | 0.8 | 8.1 | 4 |
| Electrophoresis Beckman | 9.7 | 0.8 | 7.8 | 8 |
| Affininity method Abbott Imx | 8.4 | 0.6 | 7.5 | 15 |
| Affininity method Abbott Vision | 8.8 | 0.4 | 4.8 | 8 |
| LaInIA Ames DCA 2000 | 9.1 | 0.2 | 2.6 | 29 |
| TinIA Boehringer | 9.1 | 1.0 | 10.4 | 15 |
| Immunochemical others | 8.8 | 1.7 | 19.3 | 2 |
| All method groups | 9.2 | 0.8 | 8.9 | 127 |
| GHbA1 |  |  |  |  |
| Ion exchange Ciba Corning | 11.8 | 0.3 | 2.4 | 2 |
| Ion exchange Diamat Bio-Rad | 12.1 | 0.4 | 3.2 | 7 |
| Ion exchange others | 11.5 | - | - | 1 |
| All methods groups | 12.0 | 0.4 | 3.3 | 10 |
| GHb |  |  |  |  |
| Affinity method Abbott Imx | 13.4 | 0.4 | 2.7 | 2 |
| Affinity method mini col Isolab | 10.3 | 1.4 | 13.7 | 2 |
| All method groups | 11.8 | 2.0 | 16.5 | 4 |

HbA1c: Hemoglobin A1c.

Table 2  The methods commonly used for analyzing hemoglobin A1c according to the surveys conducted by Labquality Ltd. in 2015

| Numerical summary from a diabetic EDTA-blood sample 5/2015 |
| --- |
| Analyte/method group | Mean | SD | CV% | Min | Max | Number |
| HbA1c, mmol/mol |  |  |  |  |  |  |
| Abbott Arhitect enzymatic | 53.64 | 2.45 | 4.6 | 48.4 | 32.1 | 6 |
| Axis-Shield Afinion | 54.70 | 2.17 | 4.0 | 51.0 | 58.0 | 44 |
| Beckman Coulter | 55.2 |  |  |  |  | 1 |
| Hemoque | 51.4 | 3.98 | 7.7 | 46.0 | 55.0 | 5 |
| HPLC Bio-Rad D-10 | 55.00 |  |  |  |  | 1 |
| HPLC Bio-Rad Variant | 56.00 | 3.21 | 5.7 | 54.0 | 58.0 | 2 |
| HPLC Tosoh | 56.42 | 1.23 | 2.2 | 35.6 | 59.8 | 33 |
| Roche Cobas c Tina-Quant | 54.76 | 2.56 | 4.7 | 33.0 | 57.0 | 13 |
| Roche Cobas Integra | 55.0 |  |  |  |  | 1 |
| Rocgw Tina Quant | 56.00 |  |  |  |  | 1 |
| Siemens Advia Centaur | 57.93 |  |  |  |  | 1 |
| Siemens DCA 2000+ and Vant | 55.05 | 1.96 | 3.6 | 51.0 | 58.0 | 33 |
| Thermo Scientific Konelab | 54.44 | 2.63 | 4.8 | 48.0 | 58.4 | 24 |
| All | 55.05 | 2.17 | 3.9 | 32.1 | 59.8 | 165 |
| HbA1c |  |  |  |  |  |  |
| Abbott Arhitect enzymatic | 7.20 | 0.00 | 0.0 | 7.2 | 7.2 | 3 |
| Axis-Shield Afinion | 7.17 | 0.21 | 2.9 | 6.9 | 7.5 | 44 |
| Beckman Coulter | 7.2 |  |  |  |  | 1 |
| Hemoqe | 6.88 | 0.25 | 3.7 | 6.6 | 7.2 | 6 |
| HPLC Bio-Rad D-10 | 7.2 |  |  |  |  | 1 |
| HPLC Bio-Rad Variant | 7.5 |  |  |  |  | 1 |
| HPLC Tosoh | 7.34 | 0.13 | 1.8 | 7.0 | 8.8 | 23 |
| Roche Cobas c Tina-Quant | 7.29 | 0.12 | 1.6 | 6.7 | 7.4 | 7 |
| Siemens Advia Centaur | 7.45 |  |  | 5.8 | 7.0 | 1 |
| Siemens DCA 2000+ and Vant | 7.19 | 0.18 | 2.6 | 6.8 | 7.5 | 31 |
| Thermo Scientific Konelab | 7.13 | 0.25 | 3.5 | 6.6 | 7.5 | 26 |
| All | 7.20 | 0.21 | 2.9 | 6.6 | 8.8 | 143 |

HbA1c: Hemoglobin A1c.

Table 3  Summary of the questionnaires about the use of the hemoglobin A1c units sent to 35-51 societies of laboratory medicine (mainly clinical chemistry) and returned by 28.12.2014

| Year | 2009 | 2011 | 2013 | 2014 |
| --- | --- | --- | --- | --- |
| Queries sent | 35 | 37 | 47 | 51 |
| Only % unit | 16 | 14 | 15 | 17 |
| Parallel units, % and mmol/mol | 4 | 9 | 12 | 11 |
| mmol/mol unit only in use | (1) | 6 | 11 | 13 |
| No reply | 15 | 8 | 9 | 10 |

1Germany was the first country which adopted the exclusive use of mmol/mol unit for HbA1c from 1.1.2010. The questionnaires have been dispatched by e-mail, telefax and mail mainly to Europe and to a small extent outside of Europe from 2009 to 2014.

Table 4  The relationships between the hemoglobin A1c units as % values to the mmol/mol values are presented at different levels of hemoglobin A1c

| HbA1c | HbA1c | HbA1c |
| --- | --- | --- |
|  | (%) | (mmol/mol) |
| Reference limits | 4.0-6.0 | 20-42 |
| Diagnosis limit | 6.5 | 48 |
| Treatment limits, adults | 7.0 | 53 |
| children < 6 yr | 7.5 | 69 |
| children 6-12 yr | 8.0 | 64 |
| children 13-19 yr | 8.5 | 58 |
| Poor diabetic balance | 9.0 | 75 |
| Very poor diabetic balance | 12.0 | 108 |

Equarions from the NGSP/DCCT mean values to the IFCC units: % = 0.095 ×  mmol/mol + 2.15; mmol/mol = 10.93 × % - 23.5. The table also shows the most important equations to convert the results described in % units of NGSP/DCCT into the values of IFCC expressed as mmol/mol units[23]. HbA1c: Hemoglobin A1c; NGSP/DCCT: National Glycohemoglobin Standardization Program/the Diabetes Control and Complicatioins Trial Research Group; IFCC: International Federation of Clinical Chemistry.