Copyright Information of the Article Published Online

| TITLE | Contraindications for video capsule endoscopy |
| --- | --- |
| AUTHOR(s) | Dirk Bandorski, Niehls Kurniawan, Peter Baltes, Reinhard Hoeltgen, Matthias Hecker, Dominik Stunder, Martin Keuchel |
| CITATION | Bandorski D, Kurniawan N, Baltes P, Hoeltgen R, Hecker M, Stunder D, Keuchel M. Contraindications for video capsule endoscopy. World J Gastroenterol 2016; 22(45): 9898-9908 |
| URL | http://www.wjgnet.com/1007-9327/full/v22/i45/9898.htm |
| DOI | http://dx.doi.org/10.3748/wjg.v22.i45.9898 |
| OPEN ACCESS | This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/ |
| CORE TIP | Video capsule endoscopy has emerged as a first line diagnostic tool for small bowel visualization. The few existing contraindications are discussed in this review and put into perspective. Special situations are to be considered for patients with gastrointestinal stenosis, swallowing and motility disorders, or implanted electromagnetic cardiac devices, pregnant women, young children, and magnetic resonance imaging for patients with a retained capsule. Appropriate precautions are discussed in this paper. |
| KEY WORDS | Video capsule endoscopy; Contraindications; Stenosis; Pacemaker; Aspiration; Pregnancy; Magnetic resonance imaging |
| COPYRIGHT | © The Author(s) 2016. Published by Baishideng Publishing Group Inc. All rights reserved. |
| NAME OF JOURNAL | World Journal of Gastroenterology |
| ISSN | 1007-9327 (print) and 2219-2840 (online) |
| PUBLISHER | Baishideng Publishing Group Inc, 8226 Regency Drive, Pleasanton, CA 94588, USA |
| WEBSITE | http://www.wjgnet.com |

REVIEW

Contraindications for video capsule endoscopy

Dirk Bandorski, Niehls Kurniawan, Peter Baltes, Reinhard Hoeltgen, Matthias Hecker, Dominik Stunder, Martin Keuchel

Dirk Bandorski, Matthias Hecker, University of Giessen and Marburg Lung Center, Member of the German Center for Lung Research (DZL), 35392 Giessen, Germany

Niehls Kurniawan, Peter Baltes, Martin Keuchel, Clinic for Internal Medicine, Bethesda Krankenhaus Bergedorf, Glindersweg, 21029 Hamburg, Germany

Reinhard Hoeltgen, Klinikum Westmünsterland, St. Agnes-Hospital Bocholt Rhede, Medical Clinic 1 Cardiology/Electrophysiology, Barloer Weg, 46397 Bocholt, Germany

Dominik Stunder, Research Center for Bioelectromagnetic Interaction (femu), Institute of Occupational and Social Medicine, RWTH Aachen University Hospital, Pauwelsstrasse, 52074 Aachen, Germany

Author contributions: Each author wrote a part of the paper; Bandorski D and Stunder D analyzed technical data of capsule endoscopy and implantable cardiac devices.

Correspondence to: Dirk Bandorski, MD, FESC, FACC, University of Giessen and Marburg Lung Center, Member of the German Center for Lung Research, Klinikstrasse 33, 35392 Giessen, Germany. dirk.bandorski@hkw.med.uni-giessen.de

Telephone: +49-641-98542351   Fax: +49-641-98542359

Received: August 27, 2016   Revised: October 9, 2016   Accepted: November 14, 2016

Published online: December 7, 2016

Abstract

Video capsule endoscopy (VCE) has been applied in the last 15 years in an increasing field of applications. Although many contraindications have been put into perspective, some precautions still have to be considered. Known stenosis of the gastrointestinal tract is a clear contraindication for VCE unless surgery is already scheduled or at least has been considered as an optional treatment modality. In patients with a higher incidence of stenosis, as in an established diagnosis of Crohn’s disease, clinical signs of obstruction, prior radiation or surgical small bowel resection, a preceding test with the self-dissolving patency capsule can override this contraindication. Endoscopic placement of the capsule should be considered in patients with swallowing disorders to avoid aspiration. Esophageal or gastric motility disorders may require endoscopic capsule transport or application of prokinetics if the real-time viewer proofs delayed transit. In pregnant women, VCE should be restricted to urgent cases where diagnosis cannot be postponed after delivery, as data on safety are missing. There is theoretical and clinical evidence that patients with implanted cardiac devices such as a pacemaker, cardioverters or left heart assist devices, can safely undergo VCE in spite of still existing contraindication by manufacturers. Children from the age of 2 years have safely undergone VCE. Although video capsules are not proven safe with magnetic resonance imaging (MRI), first single cases of patients incidentally undergoing MRI with an incorporated capsule have been reported, showing susceptibility artifacts but no signs of clinical harm.

Key words: Video capsule endoscopy; contraindications; stenosis; pacemaker; aspiration; pregnancy; Magnetic resonance imaging

© The Author(s) 2016. Published by Baishideng Publishing Group Inc. All rights reserved.

Bandorski D, Kurniawan N, Baltes P, Hoeltgen R, Hecker M, Stunder D, Keuchel M. Contraindications for video capsule endoscopy. World J Gastroenterol 2016; 22(45): 9898-9908  Available from: URL: http://www.wjgnet.com/1007-9327/full/v22/i45/9898.htm  DOI: http://dx.doi.org/10.3748/wjg.v22.i45.9898

Core tip: Video capsule endoscopy has emerged as a first line diagnostic tool for small bowel visualization. The few existing contraindications are discussed in this review and put into perspective. Special situations are to be considered for patients with gastrointestinal stenosis, swallowing and motility disorders, or implanted electromagnetic cardiac devices, pregnant women, young children, and magnetic resonance imaging for patients with a retained capsule. Appropriate precautions are discussed in this paper.

INTRODUCTION

Video capsule endoscopy (VCE) was introduced in 2001 as a well-tolerated, non-invasive, radiation free, disruptive method to visualize the gastrointestinal (GI) tract, in particular the small bowel. The wireless video capsule consists of one or more cameras with a corresponding lens and light source, batteries, a video chip, and an electronic circuit to either store or transmit the captured images. Depending on the manufacturer, the capsule measures 24-32 mm in length and 11-13 mm in diameter. The capsule is swallowed by the patient and then progresses through the gastrointestinal tract by peristalsis until it is excreted naturally. Only the colon capsule endoscopy needs an additional booster-solution during the procedure. The most commonly used VCE systems transmit the captured images in real-time to an external sensor array and recorder. The transmission technique is based on radiofrequency (Pillcam, Medtronic plc, Dublin, Ireland; EndoCapsule, Olympus Medical Systems Corp., Tokyo, Japan; OMOM capsule, Jinshan Science and Technology Co. Ltd., Chongqing, China) or electrical current via Human Body Communication (MiRoCam, IntroMedic Co. Ltd., Seoul, South Korea). Images captured by the CapsoCam capsule (CapsoVision Inc., Saratoga, CA, United States) are stored on-board in an integrated flash-drive, thus obviating the need for an external recorder, but requiring retrieval of the capsule to download the data[1].

Based on these properties of VCE systems and the modality of the procedure, contraindications were established by the manufacturers. Up to today, millions of VCE studies have been performed worldwide. For example, Covidien/Medtronic announced that more than 1.5 million PillCam capsules were used by the end of 2014. With this vast clinical experience, many of the initially pronounced contraindications can now be put into perspective.

This review summarizes the contraindications to VCE provided by the manufacturers and critically analyzes the theoretical reasons, the existing clinical evidence in the literature and technical data, as well as statements and guidelines of national or international societies.

CONTRAINDICATIONS FOR VCE BASED ON MANUFACTURERS’ RECOMMENDATIONS

Listed below (Table 1) are the contraindications and relative contraindications for VCE as stated by the manufacturers. Detailed contraindications are summarized based on the underlying pathophysiology (i.e., radiation enteritis, large small bowel tumor, extensive abdominal surgery, extensive small bowel or colon diverticulosis, GI perforation and fistulas are summarized under GI obstruction/obstacles).

VCE IN PATIENTS AT RISK FOR GASTROINTESTINAL STENOSIS

Known or suspected obstruction of the gastrointestinal tract bears the risk of capsule retention and consecutive complications. Intestinal obstacles like extensive diverticulosis or fistulas can have a similar effect. Capsule retention is defined by consensus as having a capsule endoscope remaining in the GI tract for a minimum of 14 d or if a directed medical, endoscopic or surgical intervention has to be implemented to retrieve the capsule[2]. The consequences of capsule retention can be a total or subtotal obstruction[3], gastrointestinal perforation[4,5], or capsule disintegration[6]. As these rare complications may occur late in previously asymptomatic patients, retrieval of a retained capsule should be considered. A case report has documented asymptomatic retention for up to 12 years. This 43-year-old patient underwent procto-colectomy for familial adenomatous polyposis (FAP) in 2000 and capsule endoscopy in a pilot study in 2004. The patient was lost at follow-up. In 2016 an abdominal computed tomography (CT) detected the capsule proximal of an anastomotic stricture. After failed endoscopic retrieval, the capsule was recovered surgically[7]. This very rare necessity of surgery for retrieval is the reason that some manufacturers include the inability to undergo surgery as a contraindication for VCE.

In a systematic review of 22840 VCE procedures, the overall retention rate was as low as 1.4% (CI: 1.2-1.6). Categorized by indication, retention rates in obscure gastro-intestinal bleeding (OGIB) were 1.2% (CI: 0.9-1.6), in Crohn’s disease (definite or suspected) 2.6% (CI: 1.6-3.9) and in the neoplastic lesions subgroup 2.1% (CI: 0.7-4.3). Out of 104 reported capsule retentions, 88 were asymptomatic (85%) and 16 had signs of partial or total intestinal obstruction. Of the retained capsules 58.7% were removed surgically, 12.5% endoscopically, or passed either spontaneously or after medical treatment in 15.8%, others were not reported in detail. In 136 cases a cause for retention was reported: Crohn’s disease 35.3%, neoplastic lesions 22.1%, NSAID-induced enteropathy 18.4%, postsurgical stenosis 7.4%, ulceration 3.7%, intestinal adhesion 2.9%, tuberculosis or radiation enteritis each 2.2%, ischemia-induced stenosis, Meckel’s diverticulum or pouch each 1.5%, peptic ulcer scar with stricture or cryptogenic multifocal ulcerous stenosing enteritis each 0.7%[8]. In addition, case reports documented capsule retention in a Zenker’s diverticulum[9], a duodenal diverticulum[10], in an ileo-rectal fistula[11], in a epiphrenic diverticulum[12], and the appendix orifice[13].

An analysis of 5428 VCE procedures in Spain came to a similar conclusion. The overall retention rate was 1.9%, and 1.5% in the OGIB subgroup and 3.3% in the inflammatory bowel disease subgroup. Retention rate raised to 5.7%-30% if at least two of these clinical symptoms were present prior to the VCE study: abdominal pain, distension and nausea/vomiting[14].

In patients with suspected GI obstruction a patency test capsule can be administered prior to the actual VCE study. If the capsule is excreted intact within 30 h, GI patency is presumed. After 30 h the lactose body of the patency capsule dissolves, leaving only a slim cellophane coating and a small tag[15]. Passage of an intact patency capsule predicts uneventful VCE[16].

Established Crohn’s disease is an indication for capsule endoscopy with an increased rate of capsule retention. A Swedish analysis found an odds ratio of 9.39 (95%CI: 3.32-26.54, p < 0.001) for capsule retention in patients with known Crohn’s disease compared to bleeding indication[17]. Even if current studies could not confirm retention rates of 13% as reported in the era before the advent of the patency capsule[18], a retention rate of 2%-3% seems to be realistic[19]. Clinical assessment, MR-enteroclysis and the use of a patency capsule can help to identify high-risk patients.

In a retrospective study 134 patients with known Crohn’s disease underwent VCE. Patients with obstructive symptoms, a history of bowel obstruction and NSAID/aspirin medication were previously excluded and 1/3 had prior small bowel follow through. Although no patency capsule test was performed on this selected group of patients, no cases of a capsule retention were observed[20]. This is in accordance with a recent retrospective multicenter study including 406 patients with known Crohn’s disease. A patency capsule test in every patient with Crohn’s disease did not show a reduction in the capsule retention rate compared to a selective use of the patency capsule in high risk patients with clinical signs of obstruction, or prior abdominal surgery[19]. In a prospective study, 57 patients with known Crohn’s disease and mild symptoms or in remission, who underwent MR-enteroclysis evaluated by two radiologists, had a good sensitivity (92.3% and 100%, respectively) and a negative predictive value (96.3% and 100%, respectively) for retention of the patency capsule as a predictor for functional stenosis test[21].

In 2009 the joint consensus of the Organisation Mondiale d’Endoscopie Digestive and the European Crohn’s and Colitis Organization recommended using imaging techniques before VCE in suspected Crohn’s disease[22]. However, in 2015 based on broader evidence, the European Society for Gastrointestinal Endoscopy (ESGE) recommended not using cross sectional imaging or patency capsule before VCE in patients with suspected Crohn’s disease in the absence of obstructive symptoms. In contrast, in established Crohn’s disease, imaging techniques and patency capsule are recommended to precede VCE[23].

Patients with a small bowel (SB) tumor seem to have a slightly higher risk of retention. A suspected tumor as an indication for VCE was associated with an odds ratio of 3.9 (95%CI: 1.2-12.8, p = 0.026)[17]. However, clinical symptoms of such tumors are typically bleeding or iron deficiency anemia. As tumors only present in a small subgroup of patients presenting with bleeding/anemia, retention even in this subgroup is rare, mostly asymptomatic, and diagnostic rather than a complication, ESGE recommends against routine precautious tests before VCE in bleeding patients. However, if a tumor is suspected by imaging techniques, device assisted enteroscopy with the option of obtaining histology is preferred over VCE[23].

In sum, suspected or known GI stenosis is a contraindication unless intestinal patency is proven, best by the passage of an intact patency capsule. The risk for capsule retention should be assumed in patients with known Crohn’s disease, clinical or radiologic signs of obstruction, a history of abomino-pelvic radiation, and after small bowel resection. Patients undergoing VCE for mid-GI bleeding without the above risks do not require preceding radiology or a patency capsule.

VCE IN PATIENTS WITH MOTILITY DISORDERS

VCE is not indicated for the diagnosis of GI motility disorders. For this purpose, a specifically designed, non-imaging wireless motility capsule (SmartPill, Medtronic plc, Dublin, Ireland) has been developed. Data from sensors measuring pH, pressure, and temperature are transmitted wirelessly for up to 5 d allowing diagnosis of gastroparesis, and prolonged transit times in the small bowel, colon or combined disorders[24-27].

Nevertheless, standard video capsule was applied in 18 patients with chronic intestinal dysmotility in the search for associated mucosal lesions. Three capsules were retained in the stomach for > 2 h, one of them during the entire recording time. However, no permanent retention, symptoms, or need for interventional treatment occurred[28]. Another study included 36 patients with severe symptomatic intestinal motor disorders for analysis of VCE image patterns compared with controls. No adverse events were mentioned in this report[29].

Although indication of VCE for diagnosis of GI motility disorders has yet to be considered as experimental, known or yet undiagnosed motility disorders may jeopardize routine VCE performed for other indications. Prolonged esophageal or gastric passage may lead to incomplete visualization of the small bowel, i.e., the cecum is not reached during working capacity of the batteries. Moderate prolongation seems to be compensated by longer battery life span in newer capsule generation[30].

VCE systems using an external recorder have the ability to display transmitted images in real-time during the procedure[31-33]. Significantly prolonged gastric transit time can be identified by this real-time viewer and a prokinetic agent can be administered[34]. A single center study reported a higher completion rate and diagnostic yield when a real time viewer was used and the capsule was placed endoscopically into the duodenum in the case of prolonged gastric transit time (> 60 min)[35]. The unselected primary endoscopic placement of the capsule into the duodenum to circumvent possible gastroparesis had no effect on complete small bowel visualization in a single center analysis of 687 hospitalized or out-patients compared to swallowing the capsule[36]. In a prospective single-center study including 100 VCE studies, a pathologic Gastroparesis Cardinal Symptoms Index questionnaire could not predict a prolonged gastric transit time nor did a delayed gastric passage have any clinical significance[37].

GI motility disorders are no contraindication for VCE. The routine use of a real time viewer directly after swallowing the capsule and after an hour enables detection of aspiration (see below) and esophageal or gastric retention and consecutive intervention.

VCE IN PATIENTS WITH IMPLANTABLE CARDIAC DEVICES

The radio transmitters of the first capsule endoscopes work with a carrier frequency of 434.1 MHz in PillCam and 433.8 MHz in EndoCapsule, similar to the C-Net mobile cellular system (450 MHz). The frequency in the newly available OMOM Capsule is 2.4 GHz. Two studies revealed electromagnetic interference (EMI) between cardiac pacemakers (PM) and the C-Net mobile cellular system in 22.4%-30.7% of the tested pacemakers[38,39]. However, the radiated power of C-Net mobile phones with 2 W is several factors higher than that of VCE with max. 100 nW. EMI with implantable cardiac devices at 2.4 GHz was also investigated in two studies[40,41] showing no risk of interference. Nevertheless, users of VCE estimated EMI between capsules and cardiac devices possibly being life-threatening for patients. Since the introduction of VCE, several in vitro and in vivo studies analyzed EMI between VCE (PillCam and EndoCapsule) and PMs (in vitro:[42-44]; in vivo:[44-54]), implantable cardioverter defibrillators (ICD) (in vitro:[55,56]; in vivo:[45-48,54,56-59]) and left ventricular assist devices (LVAD) (in vitro: none, in vivo:[52,60-68]).

In order to simulate electrical interactions under physiological conditions in patients, the authors of in vitro studies positioned PMs[43,44] or ICDs in a saline solution with a resistivity corresponding to that of muscle tissue. No interference with any of the PMs was observed. In Dubner’s study in one ICD (Belos DR, Biotronik), interference occurred reproducible when placing a test cap (technical data corresponding to first generation PillCam SB1 video capsule) over the ring and the shock coil electrode, but not over the pulse generator itself. This could still be verified even at 30 cm distance from the ICD system[56]. However, the reason for EMI remained unclear, and in vivo validation was missing. This observation is in contrast to our results. We tested five Belos ICDs and found no interference by the capsules at all, even though the devices were investigated in the most sensitive setting[54]. Furthermore, there are several in vivo studies investigating interference between VCE and PMs and ICDs. Interrogation of the devices (in all or some patients) either before and/or after VCE was performed in some studies (PM:[44,46-48,52-54]; ICD:[48,52,54,57-59]) whereas (all or some) patients in other studies were monitored with ECG monitor, telemetry or clinically (PM:[44-54] ICD:[44-49,52,54,57-59]). No interference with any of the PMs or ICDs in in vivo studies was observed. Relevant interference of wireless telemetry has been observed. In some cases, VCE videos had been corrupted[46,47,51]. If cardiac monitoring is necessary during VCE, wired systems should be used.

With regard to different capsule types, PillCam SB1, SB2, PillCam Colon1, and Olympus EndoCapsule have been studied. For the new PillCam SB3 and PillCam Colon2 with additional remote signals from the DR3 recorder to the capsule in order to adapt frame rates[69], studies are still warranted.

Only one study investigated EMI between the MiroCam endoscope that uses human body communication to transmit data and PMs (n = 3) and ICDs (n = 3)[70]. VCE was safely performed in patients with PMs and ICDs, and images from capsule endoscopy were not affected by cardiac devices. Studies relating to EMI between OMOM-Capsule and cardiac devices are lacking. For CapsoCam with on board storage of images without transmission, interference with cardiac devices is not possible.

EMI between VCE and LVAD was investigated in 10 in-vivo studies[52,60-68]. No interference was observed in any of the studies.

The United States Food and Drug Administration (FDA) and the manufacturers of transmitting capsules (Medtronic GI solutions, Olympus, IntroMedic, and Jinshan) recommend not using VCE in patients with cardiac devices. For CapsoCam without transmission technology there is no such formal contraindication.

Guidelines of the ESGE state that VCE is not contraindicated in patients with PM or ICD[71], whereas the American Society of Gastrointestinal Endoscopy guidelines consider cardiac devices as a relative contraindication for VCE[72]. The German Society of Gastroenterology, Digestive and Metabolic diseases recommends not withholding VCE in patients with a proper indication regardless of implanted cardiac devices[73].

In accordance with the recommendations of the Biotronik and Medtronic Cardio vascular group, VCE can be used in patients with cardiac devices[74,75], whereas statements from other manufacturers are not available. Technical data (maximum effective radiated power or output current and transmitter frequency) of VCE (Medtronic, Olympus, Jinshan, IntroMedic) and of the remote transmitting PillCam recorder DR3 were made available to two of the authors (Bandorski D, Stunder D). Based on this data, the maximum electromagnetic radiation in close proximity (5 mm) was calculated for VCE of Medtronic, Olympus, Jinshan as well as for Medtronic recorder DR3. Likewise, for VCE of IntroMedic the maximum obtainable interference voltage at the input of cardiac devices due to the human body communication was evaluated. The determined values are below the safety objectives set by the international product standard for cardiac devices (ISO 14117)[76] by a factor of 8 to 85.

In conclusion, VCE is safe in patients with PMs/ICDs based on technical data and in vitro/in vivo studies. The automatic frame rate control by transmitting a reverse signal from the recorder (DR3) to the capsule also remains without interference. Technical data of manual remote switching between different image acquisition rates in OMOM capsules are lacking. Wireless telemetry can impair recording of VCE images. Regarding patients with LVAD VCE seems to be safe according to in vivo results.

VCE IN PATIENTS WITH SWALLOWING DISORDERS

Capsule aspiration is a rare complication of VCE with a presumed incidence of 1 in 600-700[77,78]. Oral ingestion of the capsule is therefore contraindicated in patients with known swallowing disorder. Yet it is difficult to predict the patient’s ability to swallow the capsule safely. Aspiration was reported even if a patency capsule had been administered successfully prior to the procedure[79] or a barium swallow was uneventful[78]. In a series of 15 well-documented cases of capsule aspiration, only three patients had a history of dysphagia. The leading symptom during the aspiration was coughing (12/15)[80], which can stop even if the capsule is still within the trachea[81]. The aspiration resolved spontaneously by coughing (9/15) or via endoscopic retrieval (6/15)[80].

In one case, asymptomatic retention of a capsule for 6 d within in a bronchus and consecutive spontaneous passage through the GI tract was reported[82]. However, one patient with capsule aspiration experienced fatal extensive intracerebral hemorrhage, either provoked by initial coughing or during consecutive endoscopy for retrieval[83].

In case of an increased risk of aspiration, the capsule should be placed endoscopically directly into the duodenum[73]. This can be achieved via an overtube[84] or a special endoscopic delivery device (AdvanCE, US Endoscopy, Mentor, Oh, United States)[85]. Endoscopic placement with a Roth net is another alternative, but is more frequently associated with mucosal trauma in children than application with the dedicated delivery device[86].

In conclusion, swallowing disorders with the inability to safely swallow the capsule are a contraindication for standard procedure. However, if endoscopic placement is applied, VCE can be safely performed. The clinical challenge is the identification of patients at risk. Older patients, a history of cerebral stroke, bleeding or trauma, require a thorough history, and test for swallowing function. Children may have a test with swallowing a marshmallow.

VCE IN PREGNANCY

During pregnancy the growing uterus compresses the GI tract. Additionally, gastrointestinal transit is prolonged in the second and third trimester[87], which theoretically may jeopardize VCE procedure. There are only two published cases of VCE studies about pregnant women. Both reported no adverse events including no retention. The first case was a 30-year-old woman with extensive GI bleeding. A conventional upper endoscopy was uneventful. Lower endoscopy showed fresh blood coming out of the ileocecal valve. VCE revealed an ulcerated jejunal neuroendocrine tumor. Emergency surgery was successful and mother and child were alive and well[88]. The second case was a 20-year-old woman with a history of cavernous transformation of the portal vein with secondary thrombosis after omphalitis at the age of two. Esophageal varices were treated with sclerotherapy and banding at age 13 and 15. Due to the high risk of upper GI bleeding during pregnancy, the esophagus was examined through the PillCam ESO capsule. No esophageal or gastric varices were detected. The VCE study was uneventful with mother and child alive and well[89]. The theoretical short-term risk of retention due to altered GI motility in advanced pregnancy was not observed in either of these two cases.

However, there is no data on whether the electromagnetic field of the capsule-recorder-system could harm the unborn child. For comparison, mobile phones seem to have no negative effect[90]. In contrast, pregnancies of mothers reporting microwave use 6 mo prior to the pregnancy or during the first trimester were more likely to result in miscarriage (OR = 1.28, 95%CI: 1.02-1.59). The odds ratio was raised with an increasing level of exposure with an odds ratio of 1.59 for the highest exposure group (20 or more exposures/month)[91]. Although microwaves have a higher frequency - from 300 to 3000 MHz - than radio waves, the radio waves used by endoscopic capsules (e.g., 434 MHz for PillCam and EndoCapsule) are within the lower range of microwaves. Another comparator are effects caused by mobile phones with a much higher power than video capsules but not reaching proximity to the unborn as an intra-abdominal source of radio waves. This risk is not relevant for CapsoCam without electro-magnetic emission.

In conclusion, elective capsule endoscopy should be postponed after delivery due to missing data. Nevertheless, VCE may be considered in indications related to maternal symptoms not allowing delay of diagnosis as in relevant small bowel bleeding. Accordingly, the FDA assesses pregnancy only as a relative contraindication to VCE[92].

VCE IN CHILDREN

There has been an increased use of VCE in the pediatric population due to the possibility of avoiding ionizing radiation, deep sedation and general anesthesia[93]. The main issue of VCE in children seems to be the ability to voluntarily swallow the capsule and the fear of the capsule not being able to pass the narrow GI tract[94].

Since it was introduced, the minimum age of VCE has been lowered by the manufacturers and the FDA. In 2009 the FDA approved VCE for children of 2 years or older. The youngest age of a child undergoing a VCE study was 8 mo[95], and the lowest weight was 7.9 kg[96]. Voluntary ingestion seems feasible at an age older than 6-8 years[94], and has already been reported in a child of 4 years[86]. However, the manufacturer of PillCam recommends not letting children under the age of 8 years swallow the capsule. If endoscopic delivery is necessary, the AdvanCE delivery device was superior to the Roth-net, which caused significant mucosal trauma in 50% in a multicenter trial[86].

There have been no reports of a capsule aspiration, perforation or complete small bowel obstruction in the studies and meta-analyses of more than 1000 VCE studies with children[86,93,95-98]. In the largest meta-analysis, the retention rate was 2.3%. The risk for retention was higher in known inflammatory bowel disease (IBD 5.2%), a small bowel follow through suggestive of Crohn’s disease (CD 35.7%), and the combination of a body-mass-index below the 5th percentile and known IBD (43%). Retention rates by indication were 1.2% for OGIB, 2.6% for CD, and 2.1% for neoplastic lesions[93]. In patients with an increased risk of small bowel obstruction, a patency capsule test may reduce the risk of retention[97,98]. Guidelines of the Spanish Societies for Pediatric Gastroenterology, Hepatology, and Nutrition (SEGHNP) and for Digestive Diseases (SEPD) recommend that in suspected or established Crohn’s disease, magnetic resonance enterography or patency capsule should precede VCE in cases of obstruction symptoms[94].

MAGNETIC RESONANCE IMAGING IN PATIENTS WITH INCORPORATED CAPSULE

As no testing on magnetic resonance (MR) compatibility of VCEs has been conducted, the FDA requested a warning that a patient should not undergo magnetic resonance imaging (MRI) until excretion of the capsule has been verified[99]. The feared theoretical complication of performing an MRI scan while a capsule is still within the GI tract is migration of the capsule and the potential for bowel injury or perforation due to heat or high forces[100]. There are only few reported cases of MRI scans in patients with retained video capsules. In one case, an emergency MRI of the lumbar spine was ordered due to acute lumbar radiculopathy. The localizing sequence showed a focal susceptibility and the MRI was terminated, the capsule was excreted two days later[100]. In another case, an MRI was performed in a patient with a recurring Crohn’s disease. The MRI revealed a capsule that had been retained for two years due to a stenosis. The capsule was retrieved endoscopically with prior dilatation of the stenosis[101]. The third case was also a patient with symptoms of recurring Crohn’s disease. An MRI was performed shortly after VCE with the capsule still lying in the colon[102]. None of the three cases reported adverse events. Due to the interference of the MRI scan, VCE had no diagnostic value. Unpublished personal experience with three other patients incidentally undergoing abdominal MRI with an incorporated VCE confirms these initial reports.

COLON CAPSULE ENDOSCOPY IN PATIENTS WITH CONTRAINDICATION FOR SODIUM PHOSPHATE

The standard colon preparation prior to a colon capsule endoscopy consists of a PEG solution. In addition, sodium phosphate is used as the standard booster solution during the procedure to ensure that the capsule passes through the entire colon within the lifespan of the capsule’s battery. The ESGE guidelines for colon capsule endoscopy recommend the use of sodium phosphate as a booster for all patients with no contraindication[103]. However, sodium phosphate can cause severe complications like phosphate nephropathy, acute renal failure, hypertension, or mineral imbalance.

In the search for an alternate procedure, a pilot study showed feasibility of a low volume cleansing procedure for colon capsule endoscopy using PEG with ascorbic acid for bowel cleansing and as a boost after swallowing the capsule. CCE could be completed in 37/49 patients (76%)[104]. Another pilot trial from Japan, where sodium phosphate is contraindicated in hypertensive patients older than 63 years, proposed a diluted Gastrografin solution as an alternative booster based on a capsule excretion rate during recording of 97% (28/29 patients)[105].

CONCLUSION

Non-invasive VCE is safe, and formal contraindications can be put into perspective when observing some precautions. Based on uneventful clinical application in children, the minimum age has been lowered to 2 years. There is positive in vitro and in vivo evidence that cardiac pacemakers and defibrillators are no contraindication to VCE. Due to missing data, VCE in pregnancy should only be performed in very limited indications in cases where a delay of diagnosis until after delivery may put the mother or the unborn at risk. MRI with retained video capsule should be avoided, although the first reports describe only artifacts prohibiting proper image analysis but no harm to the patient. Suspected, known, or likely GI stenosis is a contraindication to VCE unless patency has been proven, or surgery is scheduled and preceding VCE might provide additional relevant information.

REFERENCES

1	Kurniawan N, Keuchel M. Technology. In: Keuchel M, Hagenmüller F, Tajiri H, editors. Video Capsule Endoscopy: A Reference Guide and Atlas. Berlin, Heidelberg: Springer Berlin Heidelberg, 2014: 15-20

2	Cave D, Legnani P, de Franchis R, Lewis BS. ICCE consensus for capsule retention. Endoscopy 2005; 37: 1065-1067 [PMID: 16189792 DOI: 10.1055/s-2005-870264]

3	Lin OS, Brandabur JJ, Schembre DB, Soon MS, Kozarek RA. Acute symptomatic small bowel obstruction due to capsule impaction. Gastrointest Endosc 2007; 65: 725-728 [PMID: 17383473 DOI: 10.1016/j.gie.2006.11.033]

4	Repici A, Barbon V, De Angelis C, Luigiano C, De Caro G, Hervoso C, Danese S, Preatoni P, Pagano N, Comunale S, Pennazio M, Rizzetto M. Acute small-bowel perforation secondary to capsule endoscopy. Gastrointest Endosc 2008; 67: 180-183 [PMID: 17981271 DOI: 10.1016/j.gie.2007.05.044]

5	Gonzalez Carro P, Picazo Yuste J, Fernández Díez S, Pérez Roldán F, Roncero García-Escribano O. Intestinal perforation due to retained wireless capsule endoscope. Endoscopy 2005; 37: 684 [PMID: 16010621 DOI: 10.1055/s-2005-861424]

6	Tachecí I, Ryska A, Rejchrt S, Kopácová M, Horava V, Bures J. Spontaneous disintegration of a retained video capsule in a patient with cryptogenic multifocal ulcerous stenosing enteritis: a rare complication. Endoscopy 2008; 40 Suppl 2: E104-E105 [PMID: 19085708 DOI: 10.1055/s-2007-966871]

7	Araujo IK, Pages M, Romero C, Castells A, González-Suárez B. Twelve-year asymptomatic retention of a colon capsule endoscope. Gastrointest Endosc 2016; Epub ahead of print [PMID: 27156654 DOI: 10.1016/j.gie.2016.04.045]

8	Liao Z, Gao R, Xu C, Li ZS. Indications and detection, completion, and retention rates of small-bowel capsule endoscopy: a systematic review. Gastrointest Endosc 2010; 71: 280-286 [PMID: 20152309 DOI: 10.1016/j.gie.2009.09.031]

9	Ziachehabi A, Maieron A, Hoheisel U, Bachl A, Hagenauer R, Schöfl R. Capsule retention in a Zenker’s diverticulum. Endoscopy 2011; 43 Suppl 2 UCTN: E387 [PMID: 22275012 DOI: 10.1055/s-0030-1256931]

10	Kim S, Bae SS, Chu HJ, Park JH, Kyung GC, An HD, Kim K, Gang EG. Capsule Endoscopy with Retention of the Capsule in a Duodenal Diverticulum: A Case Report. Korean J Gastroenterol 2016; 67: 207-211 [PMID: 27112247 DOI: 10.4166/kjg.2016.67.4.207]

11	Sulz MC, Anderson SH. Wireless capsule retained in an ileorectal fistula in a patient with undiagnosed Crohn’s disease. Endoscopy 2008; 40 Suppl 2: E5 [PMID: 18283619 DOI: 10.1055/s-2007-967057]

12	Rondonotti E, Martinez FJ, Barkin J, Gay G, Cheng MW. Complications: Prevention and Management. In: Keuchel M, Hagenmüller F, Tajiri H, editors. Video Capsule Endoscopy: A Reference Guide and Atlas. Berlin, Heidelberg: Springer Berlin Heidelberg, 2014: 413-422

13	Van Gossum A, François E, Hittelet A, Schmit A, Devière J. A prospective, comparative study between push enteroscopy and wireless video capsule in patients with obscure digestive bleeding. Gastroenterology 2003; 125: 276 [PMID: 12870498]

14	Fernández-Urién I, Carretero C, González B, Pons V, Caunedo Á, Valle J, Redondo-Cerezo E, López-Higueras A, Valdés M, Menchen P, Fernández P, Muñoz-Navas M, Jiménez J, Herrerías JM. Incidence, clinical outcomes, and therapeutic approaches of capsule endoscopy-related adverse events in a large study population. Rev Esp Enferm Dig 2015; 107: 745-752 [PMID: 26671587]

15	Spada C, Spera G, Riccioni M, Biancone L, Petruzziello L, Tringali A, Familiari P, Marchese M, Onder G, Mutignani M, Perri V, Petruzziello C, Pallone F, Costamagna G. A novel diagnostic tool for detecting functional patency of the small bowel: the Given patency capsule. Endoscopy 2005; 37: 793-800 [PMID: 16116528 DOI: 10.1055/s-2005-870246]

16	Herrerias JM, Leighton JA, Costamagna G, Infantolino A, Eliakim R, Fischer D, Rubin DT, Manten HD, Scapa E, Morgan DR, Bergwerk AJ, Koslowsky B, Adler SN. Agile patency system eliminates risk of capsule retention in patients with known intestinal strictures who undergo capsule endoscopy. Gastrointest Endosc 2008; 67: 902-909 [PMID: 18355824 DOI: 10.1016/j.gie.2007.10.063]

17	Höög CM, Bark LÅ, Arkani J, Gorsetman J, Broström O, Sjöqvist U. Capsule retentions and incomplete capsule endoscopy examinations: an analysis of 2300 examinations. Gastroenterol Res Pract 2012; 2012: 518718 [PMID: 21969823 DOI: 10.1155/2012/518718]

18	Cheifetz AS, Kornbluth AA, Legnani P, Schmelkin I, Brown A, Lichtiger S, Lewis BS. The risk of retention of the capsule endoscope in patients with known or suspected Crohn’s disease. Am J Gastroenterol 2006; 101: 2218-2222 [PMID: 16848804 DOI: 10.1111/j.1572-0241.2006.00761.x]

19	Nemeth A, Kopylov U, Koulaouzidis A, Wurm Johansson G, Thorlacius H, Amre D, Eliakim R, Seidman EG, Toth E. Use of patency capsule in patients with established Crohn’s disease. Endoscopy 2016; 48: 373-379 [PMID: 26561918 DOI: 10.1055/s-0034-1393560]

20	Mehdizadeh S, Chen GC, Barkodar L, Enayati PJ, Pirouz S, Yadegari M, Ippoliti A, Vasiliauskas EA, Lo SK, Papadakis KA. Capsule endoscopy in patients with Crohn’s disease: diagnostic yield and safety. Gastrointest Endosc 2010; 71: 121-127 [PMID: 19863957 DOI: 10.1016/j.gie.2009.06.034]

21	Rozendorn N, Klang E, Lahat A, Yablecovitch D, Kopylov U, Eliakim A, Ben-Horin S, Amitai MM. Prediction of patency capsule retention in known Crohn’s disease patients by using magnetic resonance imaging. Gastrointest Endosc 2016; 83: 182-187 [PMID: 26142554 DOI: 10.1016/j.gie.2015.05.048]

22	Bourreille A, Ignjatovic A, Aabakken L, Loftus EV, Eliakim R, Pennazio M, Bouhnik Y, Seidman E, Keuchel M, Albert JG, Ardizzone S, Bar-Meir S, Bisschops R, Despott EJ, Fortun PF, Heuschkel R, Kammermeier J, Leighton JA, Mantzaris GJ, Moussata D, Lo S, Paulsen V, Panés J, Radford-Smith G, Reinisch W, Rondonotti E, Sanders DS, Swoger JM, Yamamoto H, Travis S, Colombel JF, Van Gossum A. Role of small-bowel endoscopy in the management of patients with inflammatory bowel disease: an international OMED-ECCO consensus.Endoscopy 2009; 41: 618-637 [PMID: 19588292 DOI: 10.1055/s-0029-1214790]

23	Pennazio M, Spada C, Eliakim R, Keuchel M, May A, Mulder CJ, Rondonotti E, Adler SN, Albert J, Baltes P, Barbaro F, Cellier C, Charton JP, Delvaux M, Despott EJ, Domagk D, Klein A, McAlindon M, Rosa B, Rowse G, Sanders DS, Saurin JC, Sidhu R, Dumonceau JM, Hassan C, Gralnek IM. Small-bowel capsule endoscopy and device-assisted enteroscopy for diagnosis and treatment of small-bowel disorders: European Society of Gastrointestinal Endoscopy (ESGE) Clinical Guideline. Endoscopy 2015; 47: 352-376 [PMID: 25826168 DOI: 10.1055/s-0034-1391855]

24	Coleski R, Wilding GE, Semler JR, Hasler WL. Blunting of Colon Contractions in Diabetics with Gastroparesis Quantified by Wireless Motility Capsule Methods. PLoS One 2015; 10: e0141183 [PMID: 26510137 DOI: 10.1371/journal.pone.0141183]

25	Saad RJ, Hasler WL. A technical review and clinical assessment of the wireless motility capsule. Gastroenterol Hepatol (N Y) 2011; 7: 795-804 [PMID: 22347818]

26	Hasler WL. Gastroparesis. Curr Opin Gastroenterol 2012; 28: 621-628 [PMID: 23041675 DOI: 10.1097/MOG.0b013e328358d619]

27	Keuchel M, Kurniawan N, Baltes P, Bandorski D, Koulaouzidis A. Quantitative measurements in capsule endoscopy. Comput Biol Med 2015; 65: 333-347 [PMID: 26299419 DOI: 10.1016/j.compbiomed.2015.07.016]

28	Höög CM, Lindberg G, Sjoqvist U. Findings in patients with chronic intestinal dysmotility investigated by capsule endoscopy. BMC Gastroenterol 2007; 7: 29 [PMID: 17640373 DOI: 10.1186/1471-230X-7-29]

29	Malagelada C, Drozdzal M, Seguí S, Mendez S, Vitrià J, Radeva P, Santos J, Accarino A, Malagelada JR, Azpiroz F. Classification of functional bowel disorders by objective physiological criteria based on endoluminal image analysis. Am J Physiol Gastrointest Liver Physiol 2015; 309: G413-G419 [PMID: 26251472 DOI: 10.1152/ajpgi.00193.2015]

30	Rahman M, Akerman S, DeVito B, Miller L, Akerman M, Sultan K. Comparison of the diagnostic yield and outcomes between standard 8 h capsule endoscopy and the new 12 h capsule endoscopy for investigating small bowel pathology. World J Gastroenterol 2015; 21: 5542-5547 [PMID: 25987777 DOI: 10.3748/wjg.v21.i18.5542]

31	Spada C, Riccioni ME, Costamagna G. Rapid Access Real-Time device and Rapid Access software: new tools in the armamentarium of capsule endoscopy. Expert Rev Med Devices 2007; 4: 431-435 [PMID: 17605677 DOI: 10.1586/17434440.4.4.431]

32	Ogata H, Kumai K, Imaeda H, Aiura K, Hisamatsu T, Okamoto S, Iwao Y, Sugino Y, Kitajima M, Hibi T. Clinical impact of a newly developed capsule endoscope: usefulness of a real-time image viewer for gastric transit abnormality. J Gastroenterol 2008; 43: 186-192 [PMID: 18373160 DOI: 10.1007/s00535-007-2140-y]

33	Zhang JS, Ye LP, Zhang JL, Wang CY, Chen JY. Intramuscular injection of metoclopramide decreases the gastric transit time and does not increase the complete examination rate of capsule endoscopy: a prospective randomized controlled trial. Hepatogastroenterology 2011; 58: 1618-1621 [PMID: 21940331 DOI: 10.5754/hge11081]

34	Lai LH, Wong GL, Lau JY, Sung JJ, Leung WK. Initial experience of real-time capsule endoscopy in monitoring progress of the videocapsule through the upper GI tract. Gastrointest Endosc 2007; 66: 1211-1214 [PMID: 17945224 DOI: 10.1016/j.gie.2007.04.011]

35	Gao YJ, Ge ZZ, Chen HY, Li XB, Dai J, Ye CA, Xiao SD. Endoscopic capsule placement improves the completion rate of small-bowel capsule endoscopy and increases diagnostic yield. Gastrointest Endosc 2010; 72: 103-108 [PMID: 20304397 DOI: 10.1016/j.gie.2009.12.003]

36	Stanich PP, Guido J, Kleinman B, Betkerur K, Porter KM, Meyer MM. Video capsule endoscopy completion and total transit times are similar with oral or endoscopic delivery. Endosc Int Open 2016; 4: E228-E232 [PMID: 26878055 DOI: 10.1055/s-0041-110770]

37	Carter D, Eliakim R, Har-Noy O, Goldstein S, Derazne E, Bardan E. PillCam small bowel capsule endoscopy gastric passage interval association with patient’s complaints and pathological findings: a prospective study. Eur J Gastroenterol Hepatol 2014; 26: 47-51 [PMID: 24145864 DOI: 10.1097/01.meg.0000435548.11908.73]

38	Hofgärtner F, Müller T, Sigel H. [Could C- and D-network mobile phones endanger patients with pacemakers?]. Dtsch Med Wochenschr 1996; 121: 646-652 [PMID: 8635399 DOI: 10.1055/s-2008-1043051]

39	Irnich W, Batz L, Müller R, Tobisch R. Electromagnetic interference of pacemakers by mobile phones. Pacing Clin Electrophysiol 1996; 19: 1431-1446 [PMID: 8904533]

40	Mattei E, Censi F, Triventi M, Calcagnini G. Electromagnetic immunity of implantable pacemakers exposed to wi-fi devices. Health Phys 2014; 107: 318-325 [PMID: 25162422 DOI: 10.1097/hp.0000000000000113]

41	Tri JL, Trusty JM, Hayes DL. Potential for Personal Digital Assistant interference with implantable cardiac devices. Mayo Clin Proc 2004; 79: 1527-1530 [PMID: 15595337 DOI: 10.4065/79.12.1527]

42	Bandorski D, Irnich W, Brück M, Beyer N, Jakobs R. Kapselendoskopie und Herzschrittmacher - Ein Modellversuch zur Untersuchung möglicher Interferenzen. Endo heute 2007; 20: P8 [DOI: 10.1055/s-2007-974180]

43	Bandorski D, Irnich W, Brück M, Beyer N, Kramer W, Jakobs R. Capsule endoscopy and cardiac pacemakers: investigation for possible interference. Endoscopy 2008; 40: 36-39 [PMID: 18067067 DOI: 10.1055/s-2007-995353]

44	Payeras G, Piqueras J, Moreno VJ, Cabrera A, Menéndez D, Jiménez R. Effects of capsule endoscopy on cardiac pacemakers. Endoscopy 2005; 37: 1181-1185 [PMID: 16329014 DOI: 10.1055/s-2005-870558]

45	Bandorski D, Diehl KL, Jaspersen D. [Capsule endoscopy in patients with cardiac pacemakers: current situation in Germany]. Z Gastroenterol 2005; 43: 715-718 [PMID: 16088768 DOI: 10.1055/s-2005-858469]

46	Bandorski D, Jakobs R, Brück M, Hoeltgen R, Wieczorek M, Keuchel M. Capsule Endoscopy in Patients with Cardiac Pacemakers and Implantable Cardioverter Defibrillators: (Re)evaluation of the Current State in Germany, Austria, and Switzerland 2010. Gastroenterol Res Pract 2012; 2012: 717408 [PMID: 22253620 DOI: 10.1155/2012/717408]

47	Bandorski D, Lotterer E, Hartmann D, Jakobs R, Brück M, Hoeltgen R, Wieczorek M, Brock A, de Rossi T, Keuchel M. Capsule endoscopy in patients with cardiac pacemakers and implantable cardioverter-defibrillators - a retrospective multicenter investigation. J Gastrointestin Liver Dis 2011; 20: 33-37 [PMID: 21451795]

48	Cuschieri JR, Osman MN, Wong RC, Chak A, Isenberg GA. Small bowel capsule endoscopy in patients with cardiac pacemakers and implantable cardioverter defibrillators: Outcome analysis using telemetry review. World J Gastrointest Endosc 2012; 4: 87-93 [PMID: 22442746 DOI: 10.4253/wjge.v4.i3.87]

49	Dirks MH, Costea F, Seidman EG. Successful videocapsule endoscopy in patients with an abdominal cardiac pacemaker. Endoscopy 2008; 40: 73-75 [PMID: 18161651 DOI: 10.1055/s-2007-966785]

50	Dubner S, Dubner Y, Gallino S, Spallone L, Zagalsky D, Rubio H, Zimmerman J, Goldin E. Electromagnetic interference with implantable cardiac pacemakers by video capsule. Gastrointest Endosc 2005; 61: 250-254 [PMID: 15729234]

51	Guyomar Y, Vandeville L, Heuls S, Coviaux F, Graux P, Cornaert P, Filoche B. Interference between pacemaker and video capsule endoscopy. Pacing Clin Electrophysiol 2004; 27: 1329-1330 [PMID: 15461730 DOI: 10.1111/j.1540-8159.2004.00631.x]

52	Harris LA, Hansel SL, Rajan E, Srivathsan K, Rea R, Crowell MD, Fleischer DE, Pasha SF, Gurudu SR, Heigh RI, Shiff AD, Post JK, Leighton JA. Capsule Endoscopy in Patients with Implantable Electromedical Devices is Safe. Gastroenterol Res Pract 2013; 2013: 959234 [PMID: 23710168 DOI: 10.1155/2013/959234]

53	Leighton JA, Sharma VK, Srivathsan K, Heigh RI, McWane TL, Post JK, Robinson SR, Bazzell JL, Fleischer DE. Safety of capsule endoscopy in patients with pacemakers. Gastrointest Endosc 2004; 59: 567-569 [PMID: 15044901]

54	Stanich PP, Kleinman B, Betkerur K, Mehta Oza N, Porter K, Meyer MM. Video capsule endoscopy is successful and effective in outpatients with implantable cardiac devices. Dig Endosc 2014; 26: 726-730 [PMID: 24673381 DOI: 10.1111/den.12288]

55	Bandorski D, Irnich W, Brück M, Kramer W, Jakobs R. Do endoscopy capsules interfere with implantable cardioverter-defibrillators? Endoscopy 2009; 41: 457-461 [PMID: 19353490 DOI: 10.1055/s-0029-1214610]

56	Dubner S, Dubner Y, Rubio H, Goldin E. Electromagnetic interference from wireless video-capsule endoscopy on implantable cardioverter-defibrillators. Pacing Clin Electrophysiol 2007; 30: 472-475 [PMID: 17437569 DOI: 10.1111/j.1540-8159.2007.00695.x]

57	Leighton JA, Srivathsan K, Carey EJ, Sharma VK, Heigh RI, Post JK, Erickson PJ, Robinson SR, Bazzell JL, Fleischer DE. Safety of wireless capsule endoscopy in patients with implantable cardiac defibrillators. Am J Gastroenterol 2005; 100: 1728-1731 [PMID: 16086708 DOI: 10.1111/j.1572-0241.2005.41391.x]

58	Moneghini D, Lipari A, Missale G, Minelli L, Cengia G, Bontempi L, Curnis A, Cestari R. Lack of interference between small bowel capsule endoscopy and implantable cardiac defibrillators: an ‘in vivo’ electrophysiological study. United European Gastroenterol J 2016; 4: 216-220 [PMID: 27087949 DOI: 10.1177/2050640615608570]

59	Pelargonio G, Dello Russo A, Pace M, Casella M, Lecca G, Riccioni ME, Bellocci F. Use of video capsule endoscopy in a patient with an implantable cardiac defibrillator. Europace 2006; 8: 1062-1063 [PMID: 17098781 DOI: 10.1093/europace/eul116]

60	Amornsawadwattana S, Nassif M, Raymer D, LaRue S, Chen CH. Video capsule endoscopy in left ventricular assist device recipients with obscure gastrointestinal bleeding. World J Gastroenterol 2016; 22: 4559-4566 [PMID: 27182165 DOI: 10.3748/wjg.v22.i18.4559]

61	Daas AY, Small MB, Pinkas H, Brady PG. Safety of conventional and wireless capsule endoscopy in patients supported with nonpulsatile axial flow Heart-Mate II left ventricular assist device. Gastrointest Endosc 2008; 68: 379-382 [PMID: 18582876 DOI: 10.1016/j.gie.2008.03.1077]

62	Fenkel JM, Grasso MA, Goldberg EM, Feller ED. Capsule endoscopy is safe in patients with pulsatile Novacor PC left ventricular assist device. Gastrointest Endosc 2007; 65: 559-560; author reply 560 [PMID: 17321277 DOI: 10.1016/j.gie.2006.11.029]

63	Garatti A, Bruschi G, Girelli C, Vitali E. Small intestine capsule endoscopy in magnetic suspended axial left ventricular assist device patient. Interact Cardiovasc Thorac Surg 2006; 5: 1-4 [PMID: 17670498 DOI: 10.1510/icvts.2005.116871]

64	Girelli CM, Tartara P, Vitali E. Lack of reciprocal interference between capsule endoscope and left ventricular assist device. Endoscopy 2006; 38: 94-95; discussion 95 [PMID: 16429365 DOI: 10.1055/s-2005-870458]

65	Hanson BJ, Koene RJ, Roy SS, Eckman PM, John R, Chaudhary NA, Vega-Peralta J. Safety and Outcomes of Capsule Endoscopy in Patients with Left Ventricular Assist Device: a Single-Center Retrospective Case Series. J Cardiovasc Transl Res 2016; 9: 402-404 [PMID: 27250722 DOI: 10.1007/s12265-016-9701-5]

66	Letsou GV, Shah N, Gregoric ID, Myers TJ, Delgado R, Frazier OH. Gastrointestinal bleeding from arteriovenous malformations in patients supported by the Jarvik 2000 axial-flow left ventricular assist device. J Heart Lung Transplant 2005; 24: 105-109 [PMID: 15653390 DOI: 10.1016/j.healun.2003.10.018]

67	Seow CH, Zimmerman MJ. Capsule endoscopy in the detection of small-intestinal bleeding in patients supported by a nonpulsatile axial-flow Jarvik 2000 left ventricular assist device. Gastrointest Endosc 2006; 63: 1087 [PMID: 16733141 DOI: 10.1016/j.gie.2006.01.018]

68	Truss WD, Weber F, Pamboukian SV, Tripathi A, Peter S. Early Implementation of Video Capsule Enteroscopy in Patients with Left Ventricular Assist Devices and Obscure Gastrointestinal Bleeding. ASAIO J 2016; 62: 40-45 [PMID: 26501918 DOI: 10.1097/MAT.0000000000000303]

69	Eliakim R, Yassin K, Niv Y, Metzger Y, Lachter J, Gal E, Sapoznikov B, Konikoff F, Leichtmann G, Fireman Z, Kopelman Y, Adler SN. Prospective multicenter performance evaluation of the second-generation colon capsule compared with colonoscopy. Endoscopy 2009; 41: 1026-1031 [PMID: 19967618 DOI: 10.1055/s-0029-1215360]

70	Chung JW, Hwang HJ, Chung MJ, Park JY, Pak HN, Song SY. Safety of capsule endoscopy using human body communication in patients with cardiac devices. Dig Dis Sci 2012; 57: 1719-1723 [PMID: 22311369 DOI: 10.1007/s10620-012-2067-x]

71	Ladas SD, Triantafyllou K, Spada C, Riccioni ME, Rey JF, Niv Y, Delvaux M, de Franchis R, Costamagna G. European Society of Gastrointestinal Endoscopy (ESGE): recommendations (2009) on clinical use of video capsule endoscopy to investigate small-bowel, esophageal and colonic diseases. Endoscopy 2010; 42: 220-227 [PMID: 20195992 DOI: 10.1055/s-0029-1243968]

72	Wang A, Banerjee S, Barth BA, Bhat YM, Chauhan S, Gottlieb KT, Konda V, Maple JT, Murad F, Pfau PR, Pleskow DK, Siddiqui UD, Tokar JL, Rodriguez SA. Wireless capsule endoscopy. Gastrointest Endosc 2013; 78: 805-815 [PMID: 24119509 DOI: 10.1016/j.gie.2013.06.026]

73	Denzer U, Beilenhoff U, Eickhoff A, Faiss S, Hüttl P, In der Smitten S, Jakobs R, Jenssen C, Keuchel M, Langer F, Lerch MM, Lynen Jansen P, May A, Menningen R, Moog G, Rösch T, Rosien U, Vowinkel T, Wehrmann T, Weickert U. [S2k guideline: quality requirements for gastrointestinal endoscopy, AWMF registry no. 021-022]. Z Gastroenterol 2015; 53: E1-227 [PMID: 26783975 DOI: 10.1055/s-0041-109598]

74	Biotronik. Elektromagnetische Verträglichkeit von Herzschrittmachern, ICDs und CRT-Implantaten von BIOTRONIK. 2015 Available from: URL: https://www.biotronik.com/files/75891733AA975C3AC1257F16003351D8/$FILE/Stoerbeeinflussungen_Implantate_BIOTRONIK.pdf

75	Medtronic. Medizinische Verfahren und EMI-Vorsichtsmaßnahmen für implantierbare Kardioverter-Defibrillatoren und Defibrillatoren zur kardialen Resynchronisationstherapie, Handbuch für medizinisches Fachpersonal 2014. Available from: URL: http://www.medtronicheart.com/wcm/groups/mdtcom_sg/@emanuals/@era/@crdm/documents/documents/contrib_193764.pdf

76	International Organization for Standardization. Active implantable medical devices: electromagnetic compatibility: EMC test protocols for implantable cardiac pacemakers, implantable cardioverter defibrillators and cardiac resynchronization devices. ISO14117-2012. 2012 Available from: URL: http://www.iso.org/iso/catalogue_detail.htm?csnumber=54472

77	Lucendo AJ, González-Castillo S, Fernández-Fuente M, De Rezende LC. Tracheal aspiration of a capsule endoscope: a new case report and literature compilation of an increasingly reported complication. Dig Dis Sci 2011; 56: 2758-2762 [PMID: 21409372 DOI: 10.1007/s10620-011-1666-2]

78	Tabib S, Fuller C, Daniels J, Lo SK. Asymptomatic aspiration of a capsule endoscope. Gastrointest Endosc 2004; 60: 845-848 [PMID: 15557975]

79	Mannami T, Ikeda G, Seno S, Sonobe H, Fujiwara N, Komoda M, Edahiro S, Ohtawa Y, Fujimoto Y, Sato N, Kambara T, Waku T. Capsule Endoscope Aspiration after Repeated Attempts for Ingesting a Patency Capsule. Case Rep Gastroenterol 2015; 9: 347-352 [PMID: 26600772 DOI: 10.1159/000441382]

80	Amarna M, Vanlandingham A, Brahmbhatt P, Roy TM, Byrd RP. Late presentation of capsule endoscope aspiration with successful extraction by flexible bronchoscopy utilizing a snare wire loop. Endoscopy 2015; 47 Suppl 1 UCTN: E6-E7 [PMID: 25603524 DOI: 10.1055/s-0034-1377543]

81	Juanmartiñena Fernández JF, Fernández-Urien I, Vila Costas JJ. Asymptomatic bronchial aspiration of capsule endoscope: a significant complication. Rev Esp Enferm Dig 2016; 108: 605 [PMID: 27128495 DOI: 10.17235/reed.2016.4363/2016]

82	Pezzoli A, Fusetti N, Pizzo E. Capsule endoscopy diagnosis of intestinal Taenia. Gastrointest Endosc 2016; 83: 261-262 [PMID: 26320694 DOI: 10.1016/j.gie.2015.07.021]

83	Parker C, Davison C, Panter S. Tracheal aspiration of a capsule endoscope: not always a benign event. Dig Dis Sci 2012; 57: 1727-1728 [PMID: 22526588 DOI: 10.1007/s10620-012-2173-9]

84	Ito S, Hotta K, Imai K, Ono H. Overtube-assisted placement of a capsule endoscope in a patient with a swallowing disorder. Endoscopy 2016; 48 Suppl 1 UCTN: E47-E48 [PMID: 26829201 DOI: 10.1055/s-0042-100808]

85	Holden JP, Dureja P, Pfau PR, Schwartz DC, Reichelderfer M, Judd RH, Danko I, Iyer LV, Gopal DV. Endoscopic placement of the small-bowel video capsule by using a capsule endoscope delivery device. Gastrointest Endosc 2007; 65: 842-847 [PMID: 17466203 DOI: 10.1016/j.gie.2007.01.033]

86	Fritscher-Ravens A, Scherbakov P, Bufler P, Torroni F, Ruuska T, Nuutinen H, Thomson M, Tabbers M, Milla P. The feasibility of wireless capsule endoscopy in detecting small intestinal pathology in children under the age of 8 years: a multicentre European study. Gut 2009; 58: 1467-1472 [PMID: 19625281 DOI: 10.1136/gut.2009.177774]

87	Lawson M, Kern F, Everson GT. Gastrointestinal transit time in human pregnancy: prolongation in the second and third trimesters followed by postpartum normalization. Gastroenterology 1985; 89: 996-999 [PMID: 4043680]

88	Hogan RB, Ahmad N, Hogan RB, Hensley SD, Phillips P, Doolittle P, Reimund E. Video capsule endoscopy detection of jejunal carcinoid in life-threatening hemorrhage, first trimester pregnancy. Gastrointest Endosc 2007; 66: 205-207 [PMID: 17521645 DOI: 10.1016/j.gie.2006.11.021]

89	Wax JR, Pinette MG, Cartin A, Winn SS, Blackstone J. Cavernous transformation of the portal vein complicating pregnancy. Obstet Gynecol 2006; 108: 782-784 [PMID: 17018501 DOI: 10.1097/01.AOG.0000204872.46203.bf]

90	Baste V, Oftedal G, Møllerløkken OJ, Mild KH, Moen BE. Prospective study of pregnancy outcomes after parental cell phone exposure: the Norwegian Mother and Child Cohort Study. Epidemiology 2015; 26: 613-621 [PMID: 25906367 DOI: 10.1097/EDE.0000000000000293]

91	Ouellet-Hellstrom R, Stewart WF. Miscarriages among female physical therapists who report using radio- and microwave-frequency electromagnetic radiation. Am J Epidemiol 1993; 138: 775-786 [PMID: 8237966]

92	Savas N. Gastrointestinal endoscopy in pregnancy. World J Gastroenterol 2014; 20: 15241-15252 [PMID: 25386072 DOI: 10.3748/wjg.v20.i41.15241]

93	Oliva S, Cohen SA, Di Nardo G, Gualdi G, Cucchiara S, Casciani E. Capsule endoscopy in pediatrics: a 10-years journey. World J Gastroenterol 2014; 20: 16603-16608 [PMID: 25469028 DOI: 10.3748/wjg.v20.i44.16603]

94	Argüelles-Arias F, Donat E, Fernández-Urien I, Alberca F, Argüelles-Martín F, Martínez MJ, Molina M, Varea V, Herrerías-Gutiérrez JM, Ribes-Koninckx C. Guideline for wireless capsule endoscopy in children and adolescents: A consensus document by the SEGHNP (Spanish Society for Pediatric Gastroenterology, Hepatology, and Nutrition) and the SEPD (Spanish Society for Digestive Diseases). Rev Esp Enferm Dig 2015; 107: 714-731 [PMID: 26671584]

95	Nuutinen H, Kolho KL, Salminen P, Rintala R, Koskenpato J, Koivusalo A, Sipponen T, Färkkilä M. Capsule endoscopy in pediatric patients: technique and results in our first 100 consecutive children. Scand J Gastroenterol 2011; 46: 1138-1143 [PMID: 21615227 DOI: 10.3109/00365521.2011.584900]

96	Oikawa-Kawamoto M, Sogo T, Yamaguchi T, Tsunoda T, Kondo T, Komatsu H, Inui A, Fujisawa T. Safety and utility of capsule endoscopy for infants and young children. World J Gastroenterol 2013; 19: 8342-8348 [PMID: 24363526 DOI: 10.3748/wjg.v19.i45.8342]

97	Cohen SA, Ephrath H, Lewis JD, Klevens A, Bergwerk A, Liu S, Patel D, Reed-Knight B, Stallworth A, Wakhisi T, Gold BD. Pediatric capsule endoscopy: review of the small bowel and patency capsules. J Pediatr Gastroenterol Nutr 2012; 54: 409-413 [PMID: 21760541 DOI: 10.1097/MPG.0b013e31822c81fd]

98	Cohen SA, Gralnek IM, Ephrath H, Stallworth A, Wakhisi T. The use of a patency capsule in pediatric Crohn’s disease: a prospective evaluation. Dig Dis Sci 2011; 56: 860-865 [PMID: 20652742 DOI: 10.1007/s10620-010-1330-2]

99	US Food and Drug Administration. De novo classification request for PillCam Colon 2 capsule endoscsopy system. 2012 Available from: URL: http://www.accessdata.fda.gov/cdrh_docs/reviews/k123666.pdf

100	Anderson BW, Liang JJ, Dejesus RS. Capsule endoscopy device retention and magnetic resonance imaging. Proc (Bayl Univ Med Cent) 2013; 26: 270-271 [PMID: 23814387]

101	Zuber-Jerger I, Gelbmann CM, Endlicher E, Ott C, Obermeier F. Complicated wireless capsule enteroscopy in a patient with Crohn’s disease. Eur J Gastroenterol Hepatol 2009; 21: 952-954 [PMID: 19404201 DOI: 10.1097/MEG.0b013e3283200088]

102	Berry PA, Srirajaskanthan R, Anderson SH. An urgent call to the magnetic resonance scanner: potential dangers of capsule endoscopy. Clin Gastroenterol Hepatol 2010; 8: A26 [PMID: 19747984 DOI: 10.1016/j.cgh.2009.08.037]

103	Spada C, Hassan C, Galmiche JP, Neuhaus H, Dumonceau JM, Adler S, Epstein O, Gay G, Pennazio M, Rex DK, Benamouzig R, de Franchis R, Delvaux M, Devière J, Eliakim R, Fraser C, Hagenmüller F, Herrerias JM, Keuchel M, Macrae F, Munoz-Navas M, Ponchon T, Quintero E, Riccioni ME, Rondonotti E, Marmo R, Sung JJ, Tajiri H, Toth E, Triantafyllou K, Van Gossum A, Costamagna G. Colon capsule endoscopy: European Society of Gastrointestinal Endoscopy (ESGE) Guideline. Endoscopy 2012; 44: 527-536 [PMID: 22389230 DOI: 10.1055/s-0031-1291717]

104	Hartmann D, Keuchel M, Philipper M, Gralnek IM, Jakobs R, Hagenmüller F, Neuhaus H, Riemann JF. A pilot study evaluating a new low-volume colon cleansing procedure for capsule colonoscopy. Endoscopy 2012; 44: 482-486 [PMID: 22275051 DOI: 10.1055/s-0031-1291611]

105	Togashi K, Fujita T, Utano K, Waga E, Katsuki S, Isohata N, Endo S, Lefor AK. Gastrografin as an alternative booster to sodium phosphate in colon capsule endoscopy: safety and efficacy pilot study. Endosc Int Open 2015; 3: E659-E661 [PMID: 26716132 DOI: 10.1055/s-0034-1393075]

Footnotes

Manuscript source: Invited manuscript

Specialty type: Gastroenterology and hepatology

Country of origin: Germany

Peer-review report classification

Grade A (Excellent): A

Grade B (Very good): B, B

Grade C (Good): C

Grade D (Fair): 0

Grade E (Poor): 0

Conflict-of-interest statement: Baltes P and Keuchel M have received lecture fees from Covidien, Medtronic, Olympus; the other authors have nothing to declare.

Open-Access: This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

Peer-review started: August 31, 2016

First decision: September 12, 2016

Article in press: November 16, 2016

P- Reviewer: Annese V, Figueiredo PN, Kato J, Sauer CG    S- Editor: Gong ZM    L- Editor: A    E- Editor: Wang CH

# Text boxes

Table 1  Contraindications by manufacturer

| Product | Medtronic | Medtronic | Olympus | IntroMedic | Capsovision | Jinshan |
| --- | --- | --- | --- | --- | --- | --- |
| Condition | Patency | PillCam | EndoCapsule | Mirocam | CapsoCam | Science |
|  | capsule |  |  |  |  | OMOM |
|  |  |  |  |  |  | capsule |
| Known or suspected GI obstruction/obstacles, Fistulae, relevant (small bowel) diverticulosis |  | C | C | C | C | RC |
| Motility disorder incl. indigestion or slow gastric emptying |  |  | C | C | C |  |
| Cardiac pacemakers or other implanted electromedical devices |  | C | C | C |  | RC |
| Swallowing disorder (dysphagia) | C | C | C | C | C | RC |
| Pregnancy |  | RC | C | C | C | RC |
| Children under the age of (yr) | 2 | 2 (SB3) |  | 2 |  |  |
|  |  | 18 (Colon, Eso) |  |  |  |  |
| Strong electromagnetic fields i.e., MRI |  |  | C | C | C |  |
| Inability to endure capsule retrieval surgery |  |  | C |  |  | C |
| Inability to communicate sufficiently |  |  |  | C |  |  |
| Concomitant heart disease or epilepsy (due to electromagnetic radiation) |  |  |  | C |  |  |

PillCam Rapid 8 User Manual (DOC-2051-02) http://www.medtronic.com/content/dam/covidien/library/us/en/product/diagnostic-testing/rapid-v83-user-manual.pdf; Olympus EC 10 System User Manual (DE-8602257); IntroMedic Mirocam User Manual v3.9 (MM1100-U-1511); Capsovision CapsoCam SV1 Manual (Doc. No. 1151, Rev. G, ECO 11-0098; OMOM User’s Manual Version 1 (ZSSM-OM00-002). C: Contraindication; GI: Gastrointestinal; MRI: Magnetic resonance imaging; RC: Relative contraindication.