Terahertz Endoscopic Imaging for Colorectal Cancer Detection

Name of Journal: World Journal of Gastrointestinal Endoscopy

Manuscript Type: Review

Current Status and Future Perspectives

Pallavi Doradla, Cecil Joseph, Robert Giles

Pallavi Doradla, Wellman Center for Photomedicine, Harvard Medical School, Boston, Massachusetts 02114, USA

Pallavi Doradla, Cecil Joseph, Robert H. Giles, Biomedical Terahertz Technology center, Department of Physics, University of Massachusetts Lowell, Massachusetts 01854, USA

Author contributions:  Pallavi D performed the research work, analyzed the collected data, conducted the literature review, and wrote the manuscript. Cecil J contributed to the search and review of the literature and revised the draft. Robert G supervised the review.

Conflict-of-interest statement: Authors declare no conflicts of interests for this article.

Open-Access: This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

Correspondence to: Pallavi Doradla, PhD, Wellman center for photomedicine, Harvard Medical School, Boston, MA, United States. Email: pdoradla@mgh.harvard.edu, Telephone: 617-726-3417, Fax: 617-643-9208

Abstract

Terahertz (THz) imaging is emerging as a robust platform for myriad applications in the field of security, health, and material science. The terahertz regime, which comprises wavelengths spanning from microns to millimeters, is non-ionizing and has very low photon energy: making it inherently safe for biological imaging.  Colorectal cancer is the third most commonly diagnosed cancer in the world. The current standard of care for colorectal cancer screening is conventional colonoscopy, which relies exclusively on the physician’s experience. Researchers have been working on the development of a flexible terahertz endoscope, as an additional tool to aid in colorectal cancer screening. This involves building a single-channel terahertz endoscope, and investigating the terahertz response from cancerous colon, and demonstrating observable contrast levels when imaging in reflection modality. The current level of contrast provided by the prototype terahertz endoscopic system represents a significant step towards clinical endoscopic application of terahertz technology for in-vivo colon cancer screening. The aim of this paper is to provide a short review of the recent advances in terahertz endoscopic technology and cancer imaging. In particular, the potential of single-channel terahertz endoscopic imaging for colorectal cancer screening will be highlighted.

Key words: endoscopy; terahertz imaging; colonoscopy; colon; cancer detection; flexible waveguides; metal-coated; polarization-sensitive; polarization; cross-pol

The Author(s) 2016. Published by Baishideng Publishing Group Inc. All rights reserved.

Core tip: Terahertz (THz) imaging is emerging as a robust platform for a myriad of applications in the field of medicine. The terahertz regime with wavelengths spanning from microns to millimeters has the potential to obtain high-resolution images of an organ, effectively combining macroscopic and microscopic information. Terahertz reflection imaging provides an intrinsic contrast between normal and diseased tissues, in real-time. This review describes the design, development, and practical implication of flexible terahertz endoscopic system, while simultaneously obtaining an overview of the existing technology. In addition to the state-of-art terahertz endoscopy, the feasibility study of a single-channel terahertz endoscopic system for colorectal cancer screening will be highlighted.

SECTION 1: INTRODUCTION

- Colorectal cancer

Cancers figure among the leading causes of death worldwide with more than 14 million new cases diagnosed each year and causing 8.2 million deaths. Colorectal cancer (CRC) is the third most commonly diagnosed cancer in the world causing 0.7 million deaths per year (World Health Organization Data & Statistics 2012). Early diagnosis is an effective method of reducing cancer risk. The staging and subsequent treatment of CRC depends on current imaging technologies, such as colonoscopy[1,2], computed tomography (CT)[3,4], positron emission tomography (PET)[5,6], magnetic resonance imaging (MRI)[7,8], and optical coherence tomography (OCT)[9,10]. The current standard of care for CRC screening is the conventional colonoscopy, which relies on the visual inspection by a physician. During colonoscopy, the decision to remove abnormal growths is based on the physician's experience and the obtained abnormal tissue will be sent for pathological examination for diagnosis.

Besides conventional colonoscopy, CT, MRI and PET are current diagnostic imaging modalities for the detection of local and distant relapse of CRC. CT is a noninvasive technique that provides quick tomographic images of the entire colon and is better than MRI in detecting lesions smaller than 10 mm. However, it cannot detect tumors smaller than 5 mm diameter[11]. In addition, it uses a series of cross sectional X-rays that are ionizing[12]. Furthermore, X-ray imaging cannot be applied to the patients with renal failure[13]. On the other hand, MRI uses liquid enema for contrast and is more expensive than CT that uses air to achieve contrast[11]. PET provides high sensitivity and specificity, ranging from 80% to 90%, and can differentiate tumors from scar tissue created by surgery. However, it presents poor resolution unless the tumor is metabolically active and gives low sensitivity for lymph node staging in rectal cancer[14]. OCT provides high micrometer resolution and has great potential in detecting tumors. But it is limited by the high scattering of optical wavelengths in the tissue[15].

Conventional techniques such as X-ray and MRI provide images of living tissue at macroscopic levels but with low resolution and specificity. To obtain structural and functional information of an organ, microscopic imaging will be performed on biopsied samples. Currently there is no technique that can noninvasively provide in vivo high-resolution images of an organ or tissue in real-time. Using terahertz radiation with wavelengths ranging from microns to millimeter can fix this niche. In general, if a polyp or tumor is found during a standard colonoscopy, a biopsy or polypectomy may be performed during the test to determine the cancer’s presence[16]. Since 85% of all cancers are difficult to detect in early stage, clinicians recur to biopsy excisions for histopathology examination. If an imaging modality provides the ability of delineating the cancerous or abnormal region of fresh tissue in real-time, without staining the tissue, it’s not only time effective but also improves the screening capability of endoscopy. Since, terahertz rays are nonionizing and offer intrinsic contrast between normal and abnormal regions of the tissue[17], a THz endoscope integrated into a conventional endoscopic system can be used as a potential tool in the real-time examination and detection of cancerous regions of biological tissue.

Tissue abnormality and cell disorder:

Figure 1 displays the histology slides[18-20] of hyperplastic, normal, and various stages of colon cancer tissues. Usually, a tumorous tissue contains larger size nuclei with irregular shapes. The structures are disorganized and crowded. Figure 1(N) shows the enface section of normal mucosa with an inset showing normal mucus-secreting colon cell. In the enface direction, Figure 1(HP) shows both normal and hyperplastic mucosa structures for comparison. Crypts are the columnar structures in the mucosa layer of the colon tissue, made up of goblet cells, with approximately 100 μm diameter. The hyperplastic crypts tend to be order of magnitude larger and elliptical in shape. The hyperplastic crypts shown in the Figure 1(HP) are 4 times larger than normal crypts and oblong in shape. Figure 1(P) shows an enface cut of a sporadic juvenile polyp (benign). The smooth eroded surface with numerous mucus retention cysts is typical of these polyps. Figure 1(S1) shows a benign neoplasm of mucus-secreting colon cell. The Mu denoted in the figure is the mucin contained inside goblet cells. The benign tumor is characterized by crowded nuclei and shortage of mucin production in goblet cells. Figure 1(S2) exhibits neoplasm of mucus-secreting colon cell. The characteristics shown are larger nuclei, nuclei that are no longer arranged at the bottom of goblet cells, and

(N)

(P)

(S1)

(S2)

(S3)

100 µm

100 - 400 µm

(HP)

almost no mucin production.

Figure 1: Enface histology[19,20] sections of (HP) Hyperplastic mucosa, (N) Normal (inset: mucus secreting colon cell), (P) Sporadic juvenile benign polyp, (S1) low grade stage I, (S2) intermediate stage II, and (S3) high grade stage III colon.

The typical size of the nucleus is around 1 micron, whereas in neoplastic cells the nuclei tend to be larger and around 3 to 5μm. Figure 1(S3) shows a malignant neoplasm of mucus-secreting colon cell, which is characterized by the disorganization of cellular components such that it no longer resembles the normal colon tissue. The aggressive tumor cells are randomly arranged and contain large nuclei that vary in size and shape, dominating most of the cell volume. Therefore, the normal tissue is very homogenous at terahertz wavelengths while cancerous and dysplastic tissue has structures approaching the size of the wavelength. In addition, the dense structure of abnormal region can lead to higher refractive indices and can result in greater reflectance values. As a result, both of these mechanisms can engender an intrinsic contrast between normal and cancerous regions.

Screening Techniques:

Table 1: Merits and demerits of current techniques used in colorectal cancer screening

| Test | Advantages | Disadvantages |
| --- | --- | --- |
| Flexible sigmoidoscopy | - Quick and safe method - Biopsy or polypectomy can be done. - Usually doesn’t require full bowel preparation - Sedation is not required - Done every 5 years | - Bowel cleansing is required - Can miss small polyps - Views only the lower third of the colon - Can’t remove all polyps - If an abnormality is found, colonoscopy will be required. |
| Standard colonoscopy | - Very sensitive - Can view entire colon - Can do biopsy and remove polyps - Can diagnose other diseases - Done every 10 years | - Full bowel preparation needed - Can miss small polyps - More expensive - Minor sedation is required - Small risk of bleeding, bowel tears, or infection |
| Virtual colonoscopy | - Quick and noninvasive - Can view entire colon - No sedation is needed - Done every 5 years | - Need full bowel preparation - Cannot detect polyps <5mm - Possibility of false positive test results - Cannot remove polyps - If an abnormality is found, colonoscopy will be required |
| Fecal occult blood test (FOBT) | - Non-invasive - No bowel preparation is required - No sedation is required - Inexpensive - Sampling done at home | - May miss polyps and cancers that doesn’t cause bleeding - Some false positive results - Pre-test dietary limitations - Should be done every year - If an abnormality is found, colonoscopy will be required |

Colorectal cancer is one of the most common cancers across the world. The disease is slow to develop, but the early diagnosis and removal of abnormal growths is an effective method of reducing cancer risk. Expert groups recommend that people at average risk for colorectal cancer should start regular screening at the age of 50. Screening is the process of looking for cancer in people who have no symptoms of the disease. Several tests can be used to screen colorectal cancers[21]. These tests are divided into two types; tests that can find polyps as well as cancer, and tests that detect mainly cancer. The first kind looks at the structure of the colon to find abnormal regions. This can be done either with a scope inserted into the rectum or with a special x-ray imaging test. This test can prevent colorectal cancer, since the polyps found in the benign stage will be removed during the test. The second type diagnostics involve testing stools for the presence of cancer. These tests are less invasive and can be easily performed but are less likely to detect polyps. Although most expert groups generally recommend high sensitivity fecal occult blood test, sigmoidoscopy, and colonoscopy for cancer screening; several other tests such as virtual colonoscopy and barium enema are also used. Table 1 describes the merits and demerits of the techniques used for colorectal cancer screening.

- Terahertz Endoscopy

The Terahertz (THz) frequency regime of the electromagnetic spectrum extends from 1011 to 5×1012 Hz and lies between the microwave and infrared regions. Terahertz imaging has shown a great potential for in-vivo and ex-vivo identification of tissue abnormalities, hydration, and sub layer probing[22].  Recently, researchers have affirmed the use of THz frequencies in cancer detection[23,24] and dental imaging, as they can penetrate through several millimeters of tissue and enable the detection of differences in water content and tissue density[25]. The potential of THz imaging for colorectal cancer screening applications was encouraged by the positive results obtained from dental tissue[26], skin burns[27], breast[28], liver[29], gastric[30] and especially skin cancer studies[31].

Endoscopy is a minimally invasive diagnostic medical procedure to examine the interior surfaces of an organ or cavity of the body without the need for surgery. Endoscopes have traditionally been considered either “rigid” or “flexible,” depending on the geometry required by the physician in accessing different areas of the body[32]. Rigid endoscopes were the high-quality devices with stacks of lenses arranged to relay images from the tip of the scopes to eyepieces or large video cameras outside the body. These rigid endoscopes are surgical devices and are inserted through temporary access ports created by the physician. Unlike the rigid endoscope, flexible endoscope is more versatile and can be directly inserted from natural body cavities. Usually these flexible endoscopes are lower-quality devices, typically fiber-optic or with a miniature video camera at the tip directly coupled with an equally small camera lens[33]. Hence the cancer screening and decision to remove the abnormal region solely depends visual inspection and experience of a physician. In contrast, a terahertz endoscope integrated into a conventional endoscopic system can be used as a potential tool for the in vivo examination and detection of cancerous and/or precancerous colonic regions in real time. The terahertz endoscope is a medical device consisting of a long, thin, flexible THz waveguide instead of an optical fiber and uses a THz laser as a light source and a THz detector in place of a video camera for examining the interior surface of an organ or cavity and detects abnormal regions.

Terahertz endoscopic imaging has the potential to offer a safe, noninvasive medical imaging modality for delineating colorectal cancers. To test this hypothesis, the experimental measurements have to be performed in four steps: obtain ideal system-imaging frequency, evaluate base contrast, test flexible terahertz waveguides for use as an endoscope, and demonstrate terahertz waveguide based imaging of colorectal tissue. To confirm a system-imaging frequency, the frequency dependent absorption coefficient and refractive index of colorectal tissues must be acquired using a traditional time domain pulsed terahertz system with a frequency bandwidth of 0.1 to 5 THz. To evaluate the base contrast, terahertz reflectance images of human colonic tissues need to be obtained on ex vivo specimens and compared with the tissue histology. To test the flexible terahertz waveguides for use as an endoscope, flexible waveguides should be characterized at the desired imaging frequency prior to the determination of waveguide operational parameters. Finally to demonstrate terahertz endoscopic imaging, the requirement is to integrate flexible waveguide with the transceiver system, implement waveguide based reflection modality imaging, and obtain sensitivity and specificity of the device from colorectal specimens.

Previously proposed terahertz endoscopes fall into two categories. The first category uses an uncoated polymer tube to transmit THz radiation and works in transmission modality. This study based on anti-resonant hollow core waveguides used a Teflon tube to transmit terahertz radiation. Since the radiation is not confined inside the bent tube, it limits the guiding capability and possesses high bending losses[34]. Also, for endoscopic applications, which require lower bend radii, the guided field easily escapes into the air and interacts with other parts of the body thereby contaminating the resultant image. In addition, a recent near-field terahertz study relied on a polymer tube to propagate the terahertz beam with an attached bull’s-eye structure to obtain field enhancement[35]. This imaging setup works in transmission modality. However, the high absorption of terahertz radiation by tissue necessitates reflection modality imaging for in vivo applications. The second category uses a mode locked femtosecond laser. It relies on optical fibers for pulse propagation. It contains the THz source at the end of the optical fiber and is inserted into the patient[36]. Consequently, electrical connections to drive the terahertz source must be inserted into the patient. Also it requires two channels, including a first channel for guiding radiation to the sample and a second channel for guiding the reflected light to a detector. Here, the optical fiber is connected with a photoconductive antenna and is necessary to operate under high voltage which is undesirable for in-vivo applications.

A recent study by Doradla. et. al. demonstrated the design and development of a prototype terahertz endoscopic system based on flexible metal-coated terahertz waveguides[37]. The imaging system uses a single flexible waveguide channel to transmit and collect the back reflected intrinsic terahertz signal from the sample and is capable of operation in both transmission and reflection modalities. Using a metal-coated terahertz waveguide provides 99% inner surface reflectivity at all terahertz wavelengths and confines the terahertz radiation inside the waveguide. It preserves the linearly polarized launched mode and exhibits low bending loss even at larger bending angles. The hyper hemi spherical lens attached to the waveguide output end provides a diffraction limited ≈λ/2 beam waist that is free from lens aberrations. The two dimensional terahertz images obtained using polarization specific detection showed intrinsic contrast between cancerous and normal regions of both fresh and formalin fixed colorectal tissue, and show the potential of terahertz endoscopic imaging in cancer screening.

SECTION 2: OVERVIEW OF THE EXPERIMENTAL WORK

A. Terahertz Spectroscopy of Colorectal Tissue

Intrinsic contrast observed in terahertz images of human colonic tissues is indicative of a change in the complex refractive index between cancer and normal tissue at terahertz frequencies. Thus the first step to developing an imaging system is to measure the terahertz spectroscopic response of colorectal tissues and determine frequency regimes for the contrast. Reid et. al.[38] and Wahaia et. al.[39] performed THz time domain spectroscopy on cancerous and normal colon tissues, in this section we summarize their results.

Reid et. al. used a conventional THz-TDS system to image cancer, dysplastic and healthy colon tissues from 30 patients[38]. Their study was carried out in reflection mode and histopathological sections of the imaged tissues were used as the gold standard to classify tissue regions as cancer/ dysplastic/ normal and the corresponding optical properties were determined and averaged over different specimens. Their results for the frequency dependent absorption coefficient and refractive index are shown in Figure 2. As expected, the absorption increases with increasing frequency (this mimics how liquid water responds in this frequency region) and the refractive index (real part) is fairly steady in the region of interest. What is of interest for imaging applications, however, is the variation of these parameters between healthy and diseased tissue.

Figure 2: Terahertz spectroscopic results for the absorption coefficient (a) and refractive index (b) of fresh excisions of normal (blue) and cancerous (red) of colon tissues[38]. (Printed with permission)

Wahaia, et. al. investigated the terahertz reflectance and transmittance of specimens of paraffin embedded colon cancer and normal tissues[39]. The samples investigated were cut to 2 mm thicknesses. Their results are shown in Figure 3.

Figure 3: Absorption (a) and reflectance (b) measurements of paraffin embedded dehydrated fixed specimens of normal and cancerous colon tissue[39]. (Printed with permission)

As seen in Figure 3, there is a difference in the absorption coefficient of normal and cancerous colon in dehydrated formalin fixed samples. This result is extremely interesting as it indicates that water is not the sole contributor to terahertz contrast. Other studies[40], have shown that while water does contribute to the observed contrast in fresh tissues, it is not the sole mechanism. Other factors such as tissue morphology leading to scattering might also measurably affect the tissue response. We discuss possible contrast mechanisms later in this review, however, at this point the complete mechanism for the intrinsic contrast seen in terahertz images of colon tissue is not completely understood.

When determining the exposure frequency for terahertz imaging of colon cancers, it is the difference in the complex refractive index that ultimately determines contrast. Transmission images rely primarily on differences in the absorption coefficient while sample reflectance is dominated by changes to the real part of the refractive index. As seen in Reid et. al.’s data (Figure 3) while the absorption increases with increasing frequency, the difference in absorption decreases- thus frequencies above approximately 0.7 THz are not suited for transmittance based images. For reflection based imaging, however, the frequency region of interest spans 0.2 – 1 THz, as the difference in refractive index is fairly constant.

Different imaging approaches are discussed in later sections; however, there are other factors that also influence modality and frequency selection that we discuss below. In vivo applications require reflection based imaging- terahertz radiation is strongly absorbed by tissue, thus transmission through thick sections is not feasible. Lower frequencies correspond to longer wavelengths and thus lower resolution in the far field. However, higher frequencies experience stronger absorption, thus exhibit less penetration into tissue. Most single frequency systems used for imaging colon tissue thus far work at around 0.6 THz (500 µm) as a compromise between these two factors. A notable exception is the work done by Chen et. al. in transmission imaging at 0.3 THz (~1mm)[41]; they compensate for the lower far-field resolution by using a specially designed aperture which allows for significantly higher resolution in the near field.

B. Terahertz Imaging of Human Colorectal Tissues

Continuous-Wave (CW), or single frequency imaging of ex vivo samples of normal and cancerous colon tissues has been demonstrated. Wahaia et. al. have imaged paraffin embedded, formalin fixed specimens of normal and cancerous colon in both transmission and reflection modalities[39]. They used a solid-state multiplier/ amplifier chain as their source (VDI Systems) and a microbolometer array for detection. The frequency of the imaging system was 0.6 THz (500 µm). As can be seen in Figure 4, the transmittance images of dehydrated fixed tissue still show contrast between normal and cancer, while the reflection images do not show any appreciable difference. This is expected based on the spectroscopy results for these tissues, which are discussed in a prior section. The change in absorption persists in the imaging while the lack of difference in the real part of the refractive index shows up as no observable contrast in the reflectance images.

Figure 4: Photographs, absorption (transmittance) and reflection images of formalin fixed dehydrated colon tissue[39]. (Printed with permission)

Reid et. al. have imaged excised cancerous, dysplastic and healthy colonic tissues obtained from 30 patients in reflection modality[38]. They used a stand-alone portable terahertz imaging system TIP imaga1000. The frequency of the imaging system was 0.03 - 1 THz. An image is shown in figure 5, from which it can be seen that the cancerous and dysplastic tissue regions show contrast relative to the normal tissue, indicating differences in the reflected waveforms.

Figure 5: (a) An example THz image of tissue containing healthy regions, dysplasia and cancerous tissue, (b) The histology results (drawn onto a photographic image of the tissue samples) and (c) The histology results are overlaid on the THz image. In this example, regions a and b are normal tissue, c is dysplastic tissue and d is cancerous tissue[38]. (Printed with permission)

Doradla et. al. measured the terahertz reflectance of thick, fresh excisions of cancerous and normal colon tissue using a CW terahertz imaging system operating at 0.584 THz (513 µm)[42]. They used a CO2 pumped far-infrared gas laser source and a liquid helium cooled silicon bolometer as a detector. Figure 6, below shows a schematic of the imaging system. The beam is incident normal to the sample and the reflectance is measured using a beam-splitter. Wire-grid polarizers allow for the selection of specific polarizations and both images comprised of the co-polarized and cross-polarized remittance of the samples can be measured. The samples used in this study were thick fresh excisions of normal and cancerous colon tissues. The samples were measured the same day as the surgical procedure and were backed with saline soaked gauze during the measurement process to make sure that they did not dry out. Figure 7 shows some of the images collected alongside digital photographs of the sample.

Figure 6: Schematic of continuous-wave terahertz reflection imaging system[42].

As shown in Figure 7, the cross-polarized reflectance of normal colon is lower than the cross-polarized reflectance of cancerous colon. This result was found for all 4 sample sets (each sample set consisted of one normal and one cancerous tissue) measured in this study. The authors also computed the percent reflectivity difference between normal and cancerous colon tissue for both the cross-polarized and co-polarized data for each sample set while accounting for background.

Figure 7: Digital photograph (a) and corresponding terahertz reflectance images of Normal (N) and Cancerous (C) colon tissue[42].

As can be seen in Table 2, the cross-polarization channel exhibits a significantly larger reflectivity difference than the co-polarized channel. Moreover, the difference observed in cross-polarization is consistent across different sample sets (i.e. it is consistent across samples from different patients), thus it presents a potential quantitative screening tool for cancer detection.

Table 2: The relative reflectance difference between normal and cancerous colonic tissue[42].

| Sample # | Co-pol (×10-1%) | Cross-pol (%) |
| --- | --- | --- |
| Set 1 | 1.53 | 7.74 |
| Set 2 | 3.03 | 7.74 |
| Set 3 | 1.56 | 7.75 |
| Set 4 | 2.44 | 7.30 |

C. Waveguide Based Imaging

For biomedical endoscopic applications it is necessary to propagate terahertz radiation through a flexible tube that has low loss even at large bending angles. For imaging applications mode-quality is crucial. Therefore, flexible terahertz waveguide with low propagation and bending losses along with reasonable mode preservation characteristics is an essential tool for in vivo terahertz imaging of colon cancer. Terahertz waveguides have been fabricated from a variety of materials[43-45] and with different cross sectional designs[46-48]. Most of the reported terahertz waveguides are either rigid or not quite flexible at larger bore diameters[49,50]. On the other hand, flexible terahertz waveguides suffer from higher propagation losses[51,52]. Matsuura et. al. demonstrated flexible low-loss cylindrical waveguides, however, the fabrication technique employed is not feasible for bore diameters less than 3 mm[53]. Doradla et. al. reported the transmission characteristics of low loss (less than 1 dB/m), hollow, flexible, cylindrical terahertz waveguides fabricated with an inner metal, and metal/dielectric coatings that are small enough in diameter for endoscopic applications[54, 55].

Figure 8: Experimental setup for the transmission loss measurement in metal and metal dielectric (inset: (a) 4mm Ag (top), 3mm Ag, 2mm Au, and 2mm Ag, and (b) 3mm Ag/PS (top), 2mm Au/PS, and 4mm Ag/PS) coated terahertz waveguides[54].

The three operational parameters of waveguides that are crucial in determining the transmission losses and modal characteristics are waveguide inner diameter, coating material, and material thickness. Selection of coating materials such as silver and polystyrene were described in the preliminary investigation[56]. Also, the optimal coating thicknesses and requisite fabrication processes were detailed in Ref 55. In order to choose a suitable candidate waveguide for terahertz endoscopic imaging, the characterization was done in 3 steps: measuring propagation loss of the waveguide as a function of its inner diameter and coating material, obtaining bending loss as a function of bend radius and bending angle, and acquiring modal characteristics as a function of bending angle, bend radius, waveguide inner diameter, and coating material. The transmission losses and modal characteristics for flexible waveguides can be obtained using the optical layout shown in Figure 8. Harrington et.al. showed that when the waveguide bore size is about 17 times wavelength (λ) the guide is multimode[49], but when it is 12 times the wavelength or less then the waveguide becomes essentially single mode[54]. The characterization of flexible waveguides at the selected frequency[56], suitable for terahertz endoscopic system, is described in Ref 56.

Two groups so far have demonstrated waveguide based terahertz imaging. Chen et. al. used a polymethyl-methacrylate waveguide to demonstrate terahertz transmission imaging of human colon tissue[41] and Doradla et. al. demonstrated reflection modality imaging using a single channel terahertz waveguide[37].

Figure 9: Schematic of waveguide based terahertz near-field transmission imaging system[41]. (Printed with permission)

The work by Chen et. al. utilized a Continuous-wave source (Gunn diode) centered at 0.3 THz. Figure 9 shows a schematic of their system. As mentioned in Section 2, while 0.3 THz offers contrast in absorption, the resolution is limited in the far-field by the relatively long wavelength (1 mm). Chen et. al. overcome the resolution restriction by using a bull’s eye structure with a sub-wavelength aperture to get a near field resolution of 0.2 mm.

Figure 10: Terahertz transmittance images and stained histology sections showing cancerous and normal colon tissue[41]. (Printed with permission)

Figure 10 shows terahertz transmission images at 0.3 THz of specimens of fresh normal and cancerous colon that have been sectioned into 30 µm thick slices. As expected, the cancerous specimens exhibit significantly higher absorption than normal tissue. Chen et. al. investigated the response of 30 specimens and were able to demonstrate 100% sensitivity and specificity.

The work by Doradla et. al.[42,57,58] utilized a CO2 optically pumped far-infrared gas laser operating at 0.58 THz frequency. The terahertz endoscopic system contains three parts; a THz transceiver system to generate and detect the signal, system optics to guide the beam as per coupling requirements, and a flexible low-loss waveguide hollow metal-coated waveguide for propagating the terahertz beam. To obtain maximum coupling efficiency and attain maximum transmission through the waveguide, off axis parabolic mirror (OAP1 of Figure 11) has been adjusted to maintain the ratio of beam size and waveguide diameter as 0.77. This prototype endoscopic system utilized a single-channel to transmit and collect the back-reflected intrinsic terahertz signal from the sample and works in both transmission and reflection modalities to fulfil the in-vivo imaging criteria and overcome the higher THz absorption associated with the tissue. It uses a highly reflective flexible waveguide lens assembly to propagate terahertz beam.

Figure 11: Schematic of single-channel prototype terahertz (THz) endoscopic imaging setup (inset: Terahertz (a) transmission imaging of a small 10mm leaf, and (b) reflection imaging of a 25-cent coin).

Metal-coated waveguides provide 99% inner surface reflectivity at all terahertz wavelengths and confine the radiation inside the tube. Metal-coated waveguides preserve the linearly polarized launched mode and exhibits low bending loss even at larger bending angles. The hyper hemispherical lens attached to the waveguide output end provides an aberration free diffraction limited ~/2 beam waist. The technique in accordance with the present work acquires both co- and cross-polarized terahertz images using polarization sensitive detection. The data analysis indicates utilizing the cross-polarized component not only helps in obtaining Fresnel reflection free volume sampling but also in achieving a reflectance parameter that doesn’t vary with patient/individual[42,58]. Also, the current system doesn’t require contrast agents for the measurement. Since the device uses a single-channel, it can be easily integrated with the conventional optical endoscope. Furthermore, the dimensions of the terahertz probe can be shrunk for the practical usage, as the current endoscopic design relies on flexible terahertz metal-coated waveguides that are small enough in diameter (2mm to 100μm).

Figure 12: Digital photograph, cross-polarized THz reflection images of normal N versus cancerous C human colonic formalin fixed (a and b) and fresh (c and d) tissue sets[58].

This study[37,58] demonstrated the first prototype continuous-wave terahertz endoscopic system for cancer detection. The two dimensional terahertz images obtained using polarization specific detection showed intrinsic contrast between cancerous and normal regions of both fresh and formalin fixed colorectal tissue (Figure 12), and demonstrated the potential of terahertz imaging in cancer screening. Consequently, the data analysis indicated that the imaging system is capable of registering reflectance difference between normal and cancerous colon without the need for contrast agent. The experimental demonstration confirms the potential and feasibility of the device in accessing terahertz reflectivity data from previously inaccessible organs[59], thereby significantly increasing the overall impact of THz imaging for biomedical detection/screening applications[60].

SECTION 3: SUMMARY

Colorectal cancer is the third most commonly diagnosed cancer in the world. Early detection and treatment can significantly reduce the number of deaths caused by colon cancer. Current standard of care for CRC screening is an optical endoscopy which relies on visual inspection by the physician followed by histological analysis of biopsied specimens. Thus, an in vivo imaging modality capable of measuring quantitative differences between cancerous and diseased colon can significantly improve the screening of CRCs. Terahertz imaging, which is non-ionizing and highly sensitive to tissue water content can potentially fill this niche.

This review article has outlined the steps required for clinical application of terahertz imaging of CRC and provided an update on the current status of the technology. The first step was to measure the refractive indices and absorption coefficients of normal and diseased colon tissue in the terahertz region. This has been accomplished by several groups using both fresh and formalin fixed dehydrated colon tissue. The results indicate a measurable difference in specific frequency ranges for both fresh and fixed tissue. The second step was to image normal and cancerous colon at desired frequencies to confirm the contrast can be imaged. This has also been accomplished for both fresh and fixed tissue in both transmission and reflection imaging modalities using terahertz imaging systems. In order to proceed to clinical systems, the next step was to develop thin, flexible waveguides capable of endoscopic applications. This step has also been accomplished by several groups postulating a variety of endoscopic setups. The fourth step was to integrate the waveguide with a terahertz imaging system and test if waveguide enabled image acquisition was feasible. This was also demonstrated by different groups in both transmission and reflection modalities on fresh colon tissues. There is considerable evidence that terahertz imaging can potentially screen for colon cancers. A lot of the technological barriers have been overcome and the next step for the field is the development and testing of an in vivo terahertz endoscopy system capable of providing sensitivity and specificity numbers for the technique in identifying multiple stages of colon cancer.

References

- Lieberman, D. A., Weiss, D. G., Bond, J. H., Ahnen, D. J., Garewal, H., Harford, W. V., and Chejfec, G. Use of colonoscopy to screen asymptomatic adults for colorectal cancer. New England Journal of Medicine 2000; 343: 162-168 [PMID: 10900274, DOI: 10.1056/NEJM200007203430301].

- Muller AD, Sonnenberg A. Prevention of colorectal cancer by flexible endoscopy and polypectomy: a case-control study of 32 702 veterans. Annals of Internal Medicine 1995; 123(12): 904-10 [PMID: 7486484, DOI: 10.7326/0003-4819-123-12-199512150].

- Veerappan, G. R., Ally, M. R., Choi, J. H. R., Pak, J. S., Maydonovitch, C., and Wong, R. K. Extracolonic Findings on CT Colonography Increases Yield of Colorectal Cancer Screening. American Journal of Roentgenology 2010; 195: 677-686 [PMID: 20729446, DOI: 10.2214/AJR.09.3779].

- Mathur, P., Smith, J. J., Ramsey, C., Owen, M., Thorpe, A., Karim, S., and Dawson, P. M. Comparison of CT and MRI in the pre operative staging of rectal adenocarcinoma and prediction of circumferential resection margin involvement by MRI. Colorectal Disease 2003; 5: 396-401 [PMID: 12925069].

- Kantorová, I., Lipská, L., Bêlohlávek Schneiderová, M. Routine 18F-FDG PET preoperative staging of colorectal cancer: comparison with conventional staging and its impact on treatment decision making. Journal of Nuclear Medicine 2003; 44: 1784-1788 [PMID: 14602860].

- Whiteford, M. H., Whiteford, H. M., Yee, L. F., Ogunbiyi, O. A., Dehdashti, F., Siegel, B. A., and Read, T. E. Usefulness of FDG-PET scan in the assessment of suspected metastatic or recurrent adenocarcinoma of the colon and rectum. Diseases of the colon and rectum 2000; 43: 759-767 [PMID: 10859074, DOI: 10.1007/BF02238010].

- Hadfield, M. B., Nicholson, A. A., MacDonald, A. W., Farouk, R., Lee, P. W. R., Duthie, G. S., and Monson, J. R. T. Preoperative staging of rectal carcinoma by magnetic resonance imaging with a pelvic phased array coil. British journal of surgery 1997; 84: 529-531 [PMID: 9112909, DOI: 10.1046/j.1365-2168.1997.02591.x].

- Ajaj, W., Ruehm, S. G., Gerken, G., and Goyen, M. Strengths and weaknesses of dark lumen MR colonography: Clinical relevance of polyps smaller than 5 mm in diameter at the moment of their detection. Journal of Magnetic Resonance Imaging 2006; 24: 1088-1094 [PMID: 17031836].

- Hariri, L. P., Tumlinson, A. R., Besselsen, D. G., Utzinger, U., Gerner, E. W., and Barton, J. K. Endoscopic optical coherence tomography and laser induced fluorescence spectroscopy in a murine colon cancer model. Lasers in surgery and medicine 2006; 38: 305-313 [DOI: 10.1002/lsm.20305].

- Fercher AF, Drexler W, Hitzenberger CK, Lasser T. Optical coherence tomography-principles and applications. Reports on progress in physics 2003; 66(2): 239 [DOI: 10.1088/0034-4885/66/2/204].

- Fletcher, J, and Luboldt, W. CT and MR colonography: current status, research directions and comparison. European radiology 2000; 10: 786-801 [PMID: 10823635].

- Van Reeth, E., Tham, I. W., Tan, C. H., and Poh, C. L. Super resolution in magnetic resonance imaging: A review. Concepts in Magnetic Resonance Part A 2012; 40: 306-325 [DOI: 10.1002/cmra.21249].

- Sun, L., Wu, H., and Guan, Y. S. Colonography by CT, MRI and PET/CT combined with conventional colonoscopy in colorectal cancer screening and staging. World journal of gastroenterology WJG 2008; 14: 853-863 [PMID: 18240342].

- Schaefer, O., and Langer, M. Detection of recurrent rectal cancer with CT, MRI and PET/CT. European radiology 2007; 17: 2044-2054 [PMID: 17404742].

- Schmitt, J. M. Optical coherence tomography (OCT): a review. IEEE Journal of Selected Topics in Quantum Electronics 1999; 5: 1205-1215 [DOI: 10.1109/2944.796348].

- Rex DK, Bond JH, Winawer S, Levin TR, Burt RW, Johnson DA, Kirk LM, Litlin S, Lieberman DA, Waye JD, Church J. Quality in the technical performance of colonoscopy and the continuous quality improvement process for colonoscopy: recommendations of the US Multi-Society Task Force on Colorectal Cancer. The American journal of gastroenterology. 2002; 97(6): 1296 [PMID: 12094842].

- Woodward, R. M., Cole, B. E., Wallace, V. P., Pye, R. J., Arnone, D. D., Linfield, E. H., and Pepper, M. Terahertz pulse imaging in reflection geometry of human skin cancer and skin tissue. Physics in Medicine and Biology 2002; 47: 3853- 3863 [PMID: 12452577].

- O’Connell JB, Maggard MA, Ko CY. Colon cancer survival rates with the new American Joint Committee on Cancer sixth edition staging. Journal of the National Cancer Institute. 2004; 96(19): 1420-5 [PMID: 15467030].

- Burkitt H. G. Wheather’s Basic Histopathology. Churchill Livingston Inc., New York, NY, 1996.

- Andreoni, B., Crosta, C., Sonzogni, A., Pirola, M. E., Pavan, A., Bisanti, L, and Chiappa, A. C. Comparison between endoscopic and surgical treatment of screen-detected versus non-screen-detected colorectal cancers. Ecancer 2009; 3: 142 [PMID: 22276007].

- Burt RW. Colon cancer screening. Gastroenterology 2000; 119(3): 837-53 [PMID: 10982778].

- Mittleman DM, Gupta M, Neelamani R, Baraniuk RG, Rudd JV, Koch M. Recent advances in terahertz imaging. Applied Physics B. 1999; 68(6): 1085-94 [DOI: 10.1007/s003400050750].

- Fitzgerald, A. J., Wallace, V. P., Jimenez-Linan, M., Bobrow, L., Pye, R. J., Purushotham, A. D., and Arnone, D. D. Terahertz pulsed imaging of human breast tumors 1. Radiology 2006; 239: 533-540 [PMID: 16543586].

- Enatsu, T., Kitahara, H., Takano, K., and Nagashima, T. Terahertz spectroscopic imaging of paraffin-embedded liver cancer samples. Joint 32nd Int. Conf. Infrared and Millimeter Waves and the 15th Int. Conf. Terahertz Electronics. IRMMW-THz 2007; 1: 557-558 [DOI: 10.1109/ICIMW.2007.4516627].

- Crawley, D., Longbottom, C., Wallace, V. P., Pepper, M., Cole, B., and Arnone, D. Three-dimensional terahertz pulse imaging of dental tissue. Journal of Biomedical Optics 2003; 8: 303-307 [PMID: 12683858].

- Longbottom, C., Crawley, D. A., Cole, B. E., Arnone, D. D., Wallace, V. P., and Pepper, M. Potential uses of terahertz pulse imaging in dentistry: caries and erosion detection. International Society for Optics and Photonics, Proceedings of SPIE 2002; 4610: 109-112 [oi: 10.1117/12.469310].

- Tewari, P., Kealey, C. P., Bennett, D. B., Bajwa, N., Barnett, K. S., Singh, R. S., and Taylor, Z. D. In vivo terahertz imaging of rat skin burns. Journal of biomedical optics 2012; 17: 0405031-0405033 [PMID: 22559669].

- Ashworth, P. C., Pickwell-MacPherson, E., Provenzano, E., Pinder, S, Purushotham, A, Pepper, M., and Wallace, V. P. THz pulsed spectroscopy of freshly excised human breast cancer. Optics Express 2009; 17: 12444-12454 [PMID: 19654646].

- Enatsu, T., Kitahara, H., Takano, K., and Nagashima, T. Terahertz spectroscopic imaging of paraffin-embedded liver cancer samples. Joint 32nd Int. Conf. Infrared and Millimeter Waves and the 15th Int. Conf. Terahertz Electronics. IRMMW-THz 2007; 1: 557-558 [DOI: 10.1109/ICIMW.2007.4516627].

- Hou D, Li X, Cai J, Ma Y, Kang X, Huang P, Zhang G. Terahertz spectroscopic investigation of human gastric normal and tumor tissues. Physics in medicine and biology 2014; 59(18): 5423 [PMID: 25164759].

- Woodward RM, Cole BE, Wallace VP, Pye RJ, Arnone DD, Linfield EH, Pepper M. Terahertz pulse imaging in reflection geometry of human skin cancer and skin tissue. Physics in medicine and biology 2002; 47(21): 3853 [PMID: 12452577].

- Hake LW, inventor; Hake Lawrence W, assignee. Endoscope construction with means for controlling rigidity and curvature of flexible endoscope tube. United States patent US 4,890,602. 1990 Jan 2.

- Dennis C. L. Digital Endoscope Design, SPIE Press Book, Volume: SL10, 2016. ISBN: 9781510601611.

- Chen, H. Z., Lu, J. Y., and You, B. Terahertz endoscope based on anti- resonant reflecting hollow core waveguides. Conference on Lasers and Electro- Optics, CLEO 2011; 86146: 08855-1331 [doi:10.1364/CLEO_AT.2011.JThB113].

- Chiu, C. M., Chen, H. W., Huang, Y. R., Hwang, Y. J., Lee, W. J., Huang, H. Y., and Sun, C. K. All-terahertz fiber-scanning near-field microscopy. Optics Letters 2009; 34: 1084-1086 [PMID: 19340227].

- Ji, Y. B., Lee, E. S., Kim, S. H., Son, J. H., and Jeon, T. I. A miniaturized fiber-coupled terahertz endoscope system. Optics Express 2009; 17: 17082-17087 [PMID: 19770926].

- Doradla, P., Alavi, K., Joseph, C., and Giles, R. Single-channel prototype terahertz endoscopic system. Journal of biomedical optics 2014; 19: 080501-080501 [PMID: 25111144].

- Reid, Caroline B., Anthony Fitzgerald, George Reese, Robert Goldin, Paris Tekkis, P. S. O’Kelly, Emma Pickwell-MacPherson, Adam P. Gibson, and Vincent P. Wallace. "Terahertz pulsed imaging of freshly excised human colonic tissues." Physics in medicine and biology 2011; 56(14): 4333 [PMID: 21709342].

- Wahaia, Faustino, Irmantas Kasalynas, Rimvydas Venckevicius, Dalius Seliuta, Gintaras Valusis, Andrzej Urbanowicz, Gediminas Molis, Fatima Carneiro, Catia D. Carvalho Silva, and Pedro L. Granja. "Terahertz absorption and reflection imaging of carcinoma-affected colon tissues embedded in paraffin." Journal of Molecular Structure 2016; 1107: 214-219 [doi:10.1016/j.molstruc.2015.11.048].

- Sy, Stanley, Shengyang Huang, Yi-Xiang J. Wang, Jun Yu, Anil T. Ahuja, Yuan-ting Zhang, and Emma Pickwell-MacPherson. "Terahertz spectroscopy of liver cirrhosis: investigating the origin of contrast." Physics in medicine and biology 2010; 55(24): 7587 [PMID: 21098916].

- Chen, Hua, Shihua Ma, Xiumei Wu, Wenxing Yang, and Tian Zhao. "Diagnose human colonic tissues by terahertz near-field imaging." Journal of biomedical optics 2015; 20(3): 036017-036017 [PMID: 25815882].

- Doradla, P., Alavi, K., Joseph, C., and Giles, R. "Detection of colon cancer by continuous-wave terahertz polarization imaging technique." Journal of Biomedical Optics 2013; 18, (9): 090504-090504 [PMID: 25111144].

- Lu, J. T., Hsueh, Y. C., Huang, Y. R., Hwang, Y. J., and Sun, C. K. Bending loss of terahertz pipe waveguides. Optics Express 2010; 18: 26332-26338 [PMID: 21164983].

- George, R., and Harrington, J. A. Infrared transmissive, hollow plastic waveguides with inner Ag AgI coatings. Applied Optics 2005; 44: 6449-6455 [PMID: 16252656].

- Nielsen, K., Rasmussen, H. K., Adam, A. J., Planken, P. C., Bang, O., and Jepsen, P. U. Bendable, low-loss Topas fibers for the terahertz frequency range. Optics Express 2009; 17: 8592-8601 [PMID: 19434192].

- Yeh, C., Shimabukuro, F., and Siegel, P. H. Low-loss terahertz ribbon waveguides. Applied Optics 2005; 44: 5937-5946 [PMID: 16231801].

- Mitrofanov, O., James, R., Fernández, F. A., Mavrogordatos, T. K., and Harrington, J. A. Reducing transmission losses in hollow THz waveguides. IEEE Transactions Terahertz Science and Technology 2011; 1: 124-132 [DOI: 10.1109/TTHZ.2011.2159547].

- Lu, J. T., Lai, C. H., Tseng, T. F., Chen, H., Tsai, Y. F., Chen, I. J., and Sun, C. K. Terahertz polarization-sensitive rectangular pipe waveguides. Optics Express 2011; 19: 21532-21539 [PMID: 22109002].

- Harrington, J., George, R., Pedersen, P., and Mueller, E. Hollow polycarbonate waveguides with inner Cu coatings for delivery of terahertz radiation. Optics Express 2004; 12: 5263-5268 [PMID: 19484085].

- Bowden, B., Harrington, J. A., and Mitrofanov, O. Silver/polystyrene- coated hollow glass waveguides for the transmission of terahertz radiation. Optics Letters 2007; 32: 2945-2947 [doi: 10.1364/OL.32.002945].

- Vitiello, M. S., Xu, J, Beltram, F., Tredicucci, A., Mitrofanov, O., Harrington, J, and Ritchie, D. A. Guiding a terahertz quantum cascade laser into a flexible silver-coated waveguide. Journal of Applied Physics 2011; 110: 0631121-0631125 [doi: 10.1063/1.3639300].

- Dupuis, A., Stoeffler, K., Ung, B., Dubois, C., and Skorobogatiy, M. Transmission measurements of hollow-core THz Bragg fibers. JOSA B 2011; 28: 896- 907 [doi: 10.1364/JOSAB.28.000896].

- Matsuura, Y., and Takeda, E. Hollow optical fibers loaded with an inner dielectric film for terahertz broadband spectroscopy. JOSA B 2008; 25: 1949-1954 [doi: 10.1364/JOSAB.25.001949].

- Doradla, P., Joseph, C, Kumar, J., and Giles, R. Characterization of bending loss in hollow flexible terahertz waveguides. Optics Express 2012; 20: 19176-19184 [PMID: 23038558].

- Doradla, P., Joseph, C. S., Kumar, J., and Giles, R. H. Propagation loss optimization in metal/dielectric coated hollow flexible terahertz waveguides. International Society for Optics and Photonics, Proceedings of SPIE 2012; 8261: 82610P1-82610P10 [doi:10.1117/12.906833].

- Doradla, P., and Giles, R. H. Dual-frequency characterization of bending loss in hollow flexible terahertz waveguides. International Society for Optics and Photonics, Proceedings of SPIE 2014; 8985: 8985181-10 [doi:10.1117/12.2038596].

- Doradla, P., Alavi, K., Joseph, C., and Giles, R. Continuous wave terahertz reflection imaging of human colorectal tissue. Proc. SPIE 8624, Terahertz, RF, Millimeter, and Submillimeter-Wave Technology and Applications VI, 2013, 86240O [doi:10.1117/12.2004385].

- Doradla, P., Joseph, C. S., Kumar, J., and Giles R. H.  Development of terahertz endoscopic system for cancer detection. Proc. SPIE 9747, Terahertz, RF, Millimeter, and Submillimeter-Wave Technology and Applications IX, 2016; 97470F [doi:10.1117/12.2209404].

- Doradla, P., Alavi, K., Joseph, C.S. and Giles, R.H. Flexible waveguide enabled single-channel terahertz endoscopic system. Proc. SPIE 9362, Terahertz, RF, Millimeter, and Submillimeter-Wave Technology and Applications IX, 2015; 93620D [doi: 10.1117/12.2079417].

- Doradla, P. Terahertz Endoscopic System for Cancer Detection, Lambert Academic Publishing, Volume: 1, 2015. ISBN: 3659681555.