Name of Journal: World Journal of Gastrointestinal Surgery

ESPS Manuscript NO: 18202

Manuscript Type: Editorial

Mesh implants: An overview of crucial mesh parameters

Zhu LM et al. Crucial mesh implant parameters

Lei-Ming Zhu, Philipp Schuster, Uwe Klinge

Lei-Ming Zhu, Department of Surgery, Tongren Hospital, School of Medicine, Shanghai Jiaotong University, Shanghai 200336, China

Philipp Schuster, FEG Textiltechnik mbH, 52074 Aachen, Germany

Uwe Klinge, Department for General, Visceral and Transplant Surgery, the University Hospital of the RWTH Aachen, 52074 Aachen, Germany

Author contributions: Zhu LM and Schuster P conceived the study and wrote the manuscript; Klinge U reviewed the manuscript and made valuable contributions to the intellectual content and structure of the paper.

Conflict-of-interest statement: Lei-Ming Zhu declares no conflict of interest. Philipp Schuster is an employee of FEG Textiltechnik mbH, Aachen, Germany. Uwe Klinge received fees for serving as a speaker, a consultant for FEG Textiltechnik mbH, and an expert in litigation. Uwe Klinge received research funding from the German government and commercial manufacturers.

Open-Access: This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

Correspondence to: Dr. Uwe Klinge, Professor, Med, Department for General, Visceral and Transplant Surgery, the University Hospital of the RWTH Aachen, Pauwelsstraβe 30, 52074, Aachen, Germany. uklinge@ukaachen.de
Telephone: +49-241-8089352
Fax: +49-241-8082442

Received: April 10, 2015

Peer-review started: April 11, 2015

First decision: June 24, 2015

Revised: August 24, 2015

Accepted: September 25, 2015

Article in press:

Published online:

Abstract

Hernia repair is one of the most frequently performed surgical interventions that use mesh implants. This article evaluates crucial mesh parameters to facilitate selection of the most appropriate mesh implant, considering raw materials, mesh composition, structure parameters and mechanical parameters. A literature review was performed using the PubMed database. The most important mesh parameters in the selection of a mesh implant are the raw material, structural parameters and mechanical parameters, which should match the physiological conditions. The structural parameters, especially the porosity, are the most important predictors of the biocompatibility performance of synthetic meshes. Meshes with large pores exhibit less inflammatory infiltrate, connective tissue and scar bridging, which allows increased soft tissue ingrowth. The raw material and combination of raw materials of the used mesh, including potential coatings and textile design, strongly impact the inflammatory reaction to the mesh. Synthetic meshes made from innovative polymers combined with surface coating have been demonstrated to exhibit advantageous behavior in specialized fields. Monofilament, large-pore synthetic meshes exhibit advantages. The value of mesh classification based on mesh weight seems to be overestimated. Mechanical properties of meshes, such as anisotropy/isotropy, elasticity and tensile strength, are crucial parameters for predicting mesh performance after implantation.

Key words: Hernia repair; Hernia mesh; Incontinence mesh implant; Synthetic mesh; Mesh properties; Textile structure; Structure parameters; Mechanical parameters; Mesh weight; Synthetic raw materials

© The Author(s) 2015. Published by Baishideng Publishing Group Inc. All rights reserved.

Core tip: Hernia repair is one of the most frequently performed surgical interventions that use mesh implants. This article evaluates crucial mesh parameters to facilitate selection of the most appropriate mesh implant based on raw material, mesh composition, and structural and mechanical parameters. The structural parameters of the mesh, especially the porosity, are the most important predictors of the biocompatibility performance of synthetic meshes. Monofilament large-pore meshes exhibit less inflammatory infiltrate, connective tissue and scar bridging, which allows increased soft tissue ingrowth. The value of mesh classification based on the mesh weight seems to be overestimated. Other properties, such as the isotropy, elasticity and tensile strength, are crucial parameters for predicting the performance of meshes after implantation.

Zhu LM, Schuster P, Klinge U. Mesh implants: An overview of crucial mesh parameters. World J Gastrointest Surg 2015; In press

INTRODUCTION

Synthetic mesh implants are frequently used in many surgical interventions, especially in hernia repair. Mesh implants are composed of polypropylene (PP), polyethylene terephthalate (PET), expanded polytetrafluoroethylene (ePTFE), polyvinylidenefluoride (PVDF), and absorbable materials, such as polylactide (PLA), polyglycolic acid (PGA), polycaprolactone (PCL) and polydioxanone (PDO). Potential mesh-related complications include chronic infections, chronic pain and mesh rupture[1–3]. The reasons for chronic pain and the impact of mesh fixation in this context are controversial[4,5]. Chronic infections are favored by concomitant inflammatory and fibrotic reactions to the foreign body, hindering the local clearance from bacterial which leads to a chronic inflammatory wound with marked scarring, loss of compliance, mesh contraction, migration, physiochemical changes, seroma, infection, and in some cases, eventual mesh removal to resolve the problem[6]. A basic understanding of the physicochemical properties of meshes is essential for rational selection of the most appropriate device. This article evaluates the following crucial mesh parameters to facilitate selection of the most appropriate mesh implant: raw material, mesh composition, and structural and mechanical parameters (Figure 1).

BIOCHEMICAL FUNDAMENTALS

Implantation of a mesh triggers a foreign-body reaction, which plays a crucial role in the incorporation of the mesh into the host tissue. Incorporation of mesh into tissues is a complicated biochemical healing process. Implantation initiates an acute inflammatory cellular response that is initiated by protein absorption at the surface and attracts local inflammatory cells, such as macrophages, that converge to foreign body giant cells and eventually create a chronic wound around the mesh fibers. New blood vessels and collagen form around the mesh[7]. A relatively high level of macrophage invasion is detectable 20 min after mesh implantation, and these levels increase slightly and then decrease within 24 mo[8]. More than 80% of the cells in the mesh infiltrate positively express CD68, CD8, CD45R0 and vimentin, which indicates a mixture of cells of various origins and confirms the existence of multiple transition forms that are involved in the inflammatory response[9]. Complement and mast cell activation may also be involved in the mediation of local tissue responses to synthetic hernia meshes[10,11]. Cell migration is followed by collagen deposition, with an increase in the type I to type III collagen ratio over time[12]. The majority of tissue ingrowth and strength may be completed 2 wk after mesh implantation, but the final remodeling process is a very significant challenge[13]. Mesh-induced foreign body responses must be balanced to result in normal wound healing. Swift and adequate tissue ingrowth into the mesh results in superior biocompatibility and likely improved clinical performance. Intense or prolonged inflammation, bad infiltration, and immature collagen deposition result in scar plate formation, which can be accompanied by increased stiffness of the abdominal wall, shrinkage or deformation of the biomaterial, recurrence, adhesion, fistula or erosion of nearby tissue[14].

TEXTILE FUNDAMENTALS

Textile structures consist of mono- or multifilament fibers. Figure 2 shows the schematic appearance of knitted, warp-knitted, nonwoven and woven structures. Table 1 provides definitions of these different textile structures.

Table 2 presents the general essential properties of these textile structures. These properties are adjustable in a wide range through the selection of production technology and through the specific settings of the production process parameters. Most textile mesh implants are warp-knitted because of the ability of these implants to provide large pores and elasticity under load. Warp-knitted meshes also do not lose material or structural strength at margins when trimmed to the size of the surgical need. Nonwoven meshes are used as mesh implants in exceptional cases.

RAW MATERIAL AND MESH COMPOSITION

Raw material

The polymer and fiber surface affect the inflammatory response within the granuloma. Most synthetic meshes use one of following raw materials: nonabsorbable materials, such as PP, PET, PVDF and ePTFE, or absorbable materials, such as PLA, PGA, PCL, PDO and PHB. These materials may also be used in combination with each other or a range of additional materials, such as titanium and hyaluronate. The foreign body reaction is fairly uniform regardless of the type of mesh implanted, but the different raw materials affect the extent of the reaction. PP meshes result in an intensified inflammatory reaction with deposition of more collagen fibers and significantly higher collagen Type I/III ratios within the resulting scar neotissue compared with ePTFE meshes[15]. PET meshes induce the greatest foreign body reaction and longest-lasting chronic inflammatory response, which may be enhanced by the construction of PET fibers as a multifilament. Marked fibrosis and encapsulation surround ePTFE films[16]. PTFE is a more reactogenic material than PP, and it primarily stimulates the local production of pro-inflammatory cytokines. Therefore, the local anti-inflammatory effect of PP is less pronounced in comparison, but the inflammation persists for a longer time[17]. PVDF meshes produce a significantly reduced foreign body granuloma size compared with PP. PP is less stable than PVDF in vivo. Clear cracks in the surface of PP filaments have been detected 4 wk after implantation (Figure 3)[18]. These findings suggest that the raw material strongly influences the inflammatory and fibrotic responses.

Mesh composition

The primary aspects of mesh compositions are the use of different raw materials with or without surface coating in various textile designs.

Coatings may influence the degree of the inflammatory response. Nonabsorbable and absorbable materials are used for coatings. Absorbable materials are preferred if the coating provides a drug-eluting function. However, the degradation products may also influence the inflammatory response. A comparison of PP meshes, PP + polyglactin (PP + PG) meshes and PP + titanium (PP + TI) meshes demonstrated a reduced inflammatory reaction in the PP mesh group and increased reaction in the PP + PG mesh group. The PP mesh induced large early elevations in vascular endothelial growth factor (VEGF), cyclooxygenase-2 (COX2) and collagen levels, whereas the PP + PG mesh caused only small elevations in the levels of these factors. PP + TI meshes induced inflammatory response levels in between those of the other 2 meshes[19]. Human fibroblasts colonized on the macroporous PP side of a composite mesh made of two PP layers, but no cell growth occurred on the film PP side[20]. The suppressive effect of the mesh on the transforming growth factor beta1 (TGF-beta1) was more pronounced for partially absorbable materials compared with pure PP meshes, which suggests that a change in raw material composition and type affects the early biological reaction of connective tissue cells to the mesh[21]. Woven and nonwoven meshes have received less attention. Raptis et al[22] demonstrated that woven PP meshes became fully peritonealized intraperitoneally but generated thicker and more plentiful adhesions than nonwoven PP. PP nonwoven prosthesis are comparable to conventional warp-knitted meshes[23].

The textile design markedly influences the inflammatory reaction to the mesh. Using the best polymer in a poor textile design may lead to pronounced inflammation and scar formation. In contrast, an adequate tissue reaction may be achieved with a suboptimal polymer if the essential parameters of the textile design (e.g., the filament structure and pore size) are considered. The particular type of mesh used in hernia repair may affect the wound healing response and clinical outcome[24].

STRUCTURE PARAMETERS

Pore characteristics

The characteristics of the mesh used – primarily the pore characteristics especially the collapse of pores under strain, amount of mesh material, prosthesis weight, and mechanical properties – crucially influence the dynamic incorporation. In 1997, Amid[25] identified mesh porosity as the decisive factor for risk of infection. Amid defined pores larger than 75 μm as macropores before large-pore meshes (3-5 mm) were developed. Klinge et al[26] evaluated a remarkable number of explanted meshes and found that the mesh porosity was the most important determinant of the tissue reaction and risk of scar entrapment. The pore size must be much larger than 75 µm to preserve tissue integration without filling the pores with scar tissue. A pore size > 1 mm is required for PP, and the pore size should be > 3 mm in cases of mechanical strain. Meshes with large pores exhibit less inflammatory infiltrate, connective tissue, fistula formation, calcification, and bridging (i.e., the pores are filled by scar tissue) than meshes with small pores[27,28]. Granulomas normally form around individual mesh fibers as part of the foreign body reaction, but the term “bridging” describes the process whereby individual granulomas become confluent with each other and encapsulate the entire mesh, which leads to a stiff scar plate and reduced flexibility[29]. A pore that is not completely filled by scar tissue is considered “effective” according to Mühl et al[30]. Therefore, large pore sizes preserve the “effective porosity” and thus avoid formation of scar bridges (Figure 4).

It is difficult to define a best pore size a priori because different raw materials result in different “effective porosities”. Bridging of granuloma and encapsulation of the entire mesh is more likely for PP meshes with small pores (< 800 (m)[31]. In contrast, PVDF meshes do not exhibit bridging even for pore sizes of < 650 (m[32,33]. Klinge et al[26] characterized large-pore meshes using a textile porosity > 60% or an effective porosity > 0%. Pore shape may also determine integration. Lake et al[34] found that hexagonal pores resulted in the strongest tissue ingrowth, followed by square pores and diamond pores.

Mesh weight

Synthetic meshes may be classified as heavyweight or lightweight. The mesh weight depends on the polymer weight (raw material) and the amount of material used.

Coda et al[35] proposed a classification system based on the mesh weight that includes simple, composite and combined meshes. Meshes with weight per unit area of greater than 140 g/m² are defined as heavyweight meshes, meshes with weight per unit area in the range 35-70 g/m2 are defined as lightweight meshes and meshes with weight per unit area in the range of 70-140 g/m2 are defined as standard-weight meshes. Lightweight meshes generally contain less material and induce a less-pronounced foreign body reaction and decreased inflammatory response, which results in better tissue incorporation, increased prosthesis compliance, and decreased patient discomfort and pain. In an animal study, restriction of the abdominal wall mobility was significantly reduced and the inflammatory reaction and connective tissue formation were markedly diminished with lightweight meshes compared with heavyweight meshes[36]. Randomized prospective trials compared lightweight and heavyweight meshes for ventral hernia repair and found that they had equal outcomes in terms of ventral hernia recurrence[37]. Patients with lightweight mesh hernia repair exhibited better outcomes in terms of pain and seroma and an earlier return to activity[38].

Most current lightweight meshes have larger pores than heavyweight small-pore constructions[39]. However, mesh classification only in terms of weight disregards fiber and pore characteristics. Weyhe et al[40] considered the textile mesh construction, which they characterized in terms of the pore size and filament structure, as a more important determinant of foreign body reactions after implantation than absolute material reduction. This result attenuates the importance of mesh weight for the prediction of biocompatibility[41]. The advantages of lightweight meshes may be primarily related to their tendency to utilize a large pore size and/or monofilament.

However, an excessive reduction of mesh weight may also decrease the tensile strength. Lightweight meshes are sufficiently strong to resist abdominal wall pressure, but these meshes lose some burst strength compared with heavyweight meshes[32,42]. Experiments using small animals suggest that heavyweight small-pore meshes may withstand greater forces of scar contraction than large-pore lightweight meshes and may exhibit less shrinkage. Zogbi et al[43] found that lightweight PP mesh exhibited greater median shrinkage than heavyweight PP mesh in rats 7, 28 and 90 d after implantation.

MECHANICAL PARAMETERS

Mechanical properties are important parameters to consider when determining the suitability of a particular mesh for a specific clinical situation. However, surgeons typically implant meshes to provide maximum overlap over the defect with little regard for the mechanical properties of the mesh. Each synthetic mesh is composed of a unique combination of the material properties of the polymer and the textile design. The textile properties depend on the manufacturing process and the manufacturing process parameters.

The choice of raw materials determines a material’s properties, which in many cases implies a combination of the properties of more than one raw material. These features ultimately determine the mechanical properties of the resulting mesh. An important consequence of the manufacturing process is the anisotropy of the tensile strength, elasticity, burst strength and stiffness. The American Society for Testing and Materials (ASTM) specification (D 4850 Terminology of textile structures) provides definitions of these properties (Table 3)[44].

The actual load on the abdominal wall is of major relevance for the selection of the suitability of meshes for use in ventral hernia repair. Different groups often perform simple tensile tests (N/cm), measurements of the inner abdominal pressure (Pa = N/mm²) or calculations of the abdominal wall tension (N/cm) using the Young–Laplace equation to characterize the native abdominal wall properties. The different measuring methods and different units should be considered when comparing these measurement results. Conversion of the inner abdominal wall pressure (Pa = N/mm²) (using the Young-Laplace equation) to the abdominal wall tension (N/cm) is only possible if the circumference of the patient is also provided. Use of the Young-Laplace equation requires a distinction between the sphere-like anatomy of the groin and the cylinder-like anatomy of the abdominal wall.

Hollinsky et al[45] measured the tensile load of the linea alba, the anterior and posterior rectus sheath, and scar tissue following median laparotomy in fresh cadavers and found that the tissue in the epigastric region ruptured at a mean horizontal load of 10 N/mm2 in the linea alba and 6.9 N/mm2 in scar tissue and at a mean vertical load of 4.5 N/mm2 in the linea alba and 3.3 N/mm2 in scar tissue. In earlier research, Williams et al[46] estimated the maximum force applied to the abdominal wall after hernia repair surgery as 22 N/cm in the cranial/caudal direction and 32 N/cm in the lateral direction. Cobb et al[47] investigated the intra-abdominal pressure using a transurethral bladder (Foley) catheter under different physical situations, including standing, sitting, bending at the waist, bending at the knees, performing abdominal crunches, jumping, climbing stairs, bench-pressing 25 pounds, arm curling 10 pounds, and performing a valsalva and coughing while sitting and standing, and identified a pressure of 22.7 kPa (171 mmHg) as the maximum pressure during coughing. Deeken et al[48] argued that stress in the transverse direction can reach levels of 47.8 N/cm in obese males with a large abdominal circumference. The true peak pressure in situations such as expectoration or sternutation in the abdominal wall was not fully addressed, but it is accepted that 22 N/cm in the cranial/caudal and 32 N/cm in the lateral direction are the maximum forces applied to the abdominal wall after hernia repair surgery[49]. A load of 16 N/cm is accepted as the maximum load in the groin because of the more sphere-like anatomy of the groin[50].

The natural elasticity of the abdominal wall at 32 N/cm is approximately 38%, with higher resilience in the horizontal direction than the longitudinal direction[45,46]. DuBay et al[51] indicate that the use of meshes in ventral hernia repair increases abdominal wall elasticity, which results in lower recurrence rates. Lightweight meshes exhibit an elasticity of approximately 20%-35% at 16 N/cm, but heavyweight meshes exhibit half of this elasticity (4%-15% at 16 N/cm), which may restrict abdominal distension[39]. An inappropriate mesh tensile strength, which results in an inappropriate ability of the mesh material to stretch, may potentially lead to poor functional results, with pain, hernia recurrence or prolapse. Elongation rates of greater than 30% indicate that these materials may stretch more than the native human abdominal wall. These meshes may not maintain functional repair, which could result in bulging or recurrence[48].

Tensile strengths of greater than 100 N/cm of conventional heavyweight meshes (e.g., Prolene) are disproportionate and not necessary for effective repair[39]. Most synthetic meshes, even the lightest meshes, reach a tensile strength of at least 32 N/cm and are sufficiently strong. The mean burst strength and stiffness of lightweight meshes 5 mo after implantation in a pig was significantly less than those of heavyweight and middleweight meshes, but the burst strength for all meshes tested was much greater than the strengths measured for the abdominal wall fascia alone[32]. Bellón et al[52] demonstrated that the tensile strengths of lightweight and heavyweight meshes were comparable 90 d after implantation. However, Petro et al[53] recently reported 7 cases of mechanical failure or fracturing of lightweight monofilament polyester meshes after open incisional hernia repair. Zuvela et al[3] and Lintin et al[54] reported central ruptures of low-weight PP meshes after initial sublay incisional hernia repair. These isolated case reports are insufficient to question the use of lightweight meshes in ventral hernia repair, but one should consider that the maximum initial tensile strength of synthetic meshes did not predict long-term strength after implantation[55]. Eliason et al[56] demonstrated that BardMesh, Dualmesh, and Prolene exhibited significantly reduced tensile strength, and BardMesh, Proceed, Prolene, ProLite, ProLite Ultra, and Ultrapro exhibited significantly increased permanent elongation after exposure to 1000 cycles of repetitive loading sequences that simulated changes in the intra-abdominal pressure. Mesh elongation also led to the loss of effective porosity in most meshes, which is an important aspect for scar formation and foreign body reaction[57]. Stiffness and breaking strength also vary widely among available meshes for hernia repair, and most meshes exhibit significant anisotropy in terms of their mechanical behavior. Pott et al[49] compared six meshes composed of different raw materials and different textile structures. All six mesh types exhibited differences in maximum tensile strength (11.1 ± 6.4 to 100.9 ± 9.4 N/cm), stiffness (0.3 ± 0.1 to 4.6 ± 0.5 N/mm), and elongation at break (150% ± 6% to 340% ± 20%) based on the load direction: the warp direction, or “longitudinal direction”, versus the weft direction, or “orthogonal direction”. Deeken et al[58] recently evaluated 13 mesh types that exhibited a wide range of mechanical properties. Some meshes were nearly isotropic, with nearly similar properties in the vertical and horizontal strain directions [C-QUR™, DUALMESH(®), PHYSIOMESH™, and PROCEED(®)], but other meshes were highly anisotropic (Ventralight™ ST, Bard™ Mesh, and Bard™ Soft Mesh). Some meshes exhibited a nearly linear behavior (Bard™ Mesh), but other meshes were non-linear, with a long toe region followed by a sharp rise in tension.

Meshes with different mechanical properties are treated as uniform and interchangeable, but it is important to understand the characteristics of the meshes to identify an appropriate mesh for each patient and place the mesh in an appropriate position to avoid mechanical mismatch, which may impair graft fixation, and enable optimized integration into the host tissue[59,60]. Therefore, surgeons may use meshes with isotropic properties regardless of the mesh orientation, but surgeons should pay attention to the orientation of meshes with anisotropic properties, which should be placed with their major elasticity in the appropriate direction to match the physiological stretch abilities (Table 4).

NEW DEVELOPMENTS

The evolution of meshes is not complete. New synthetic meshes are continuously developed, and new polymers and innovative coatings are continuously introduced. Ulrich et al[61] examined 3 new warp-knitted synthetic meshes composed of different polymers with different tensile properties, polyetheretherketone (PEEK), polyamide (PA) and a composite, gelatin-coated PA (PA + G), in a rat model. All new materials exhibited better tissue integration, new collagen deposition and sustained neovascularization compared with PP meshes. Therefore, these new materials provide a promising alternative for future mesh developments. Meshes manufactured from native spider dragline revealed rapid cell migration, complete degradation, formation of a stable scar with constant tensile strength values and the highest relative elongation among standard biological and synthetic meshes[62].

Biosynthetic meshes are a possible cost-effective alternative to synthetic and biological meshes. Biodegradable polymers, instead of animal or cadaver tissue, provide a temporary scaffold for deposition of proteins and cells that are necessary for tissue ingrowths, neovascularization, and host integration[63]. Powell et al[64] reported good results in the early phase for ‘‘synthetic remodeling meshes’’ made from PGA/trimethylene carbonate in a study of 70 patients who underwent hiatal hernia repair. However, Symeonidis et al[65] used the same ‘‘synthetic remodeling mesh’’ in a pilot study of inguinal hernia repair and reported discouraging results, with a 38% recurrence rate after a mean follow-up of 2 years, which questions the general suitability of this mesh. Another fully absorbable mesh composed of knitted poly-4-hydroxybutyrate monofilament fibers, named the Phasix mesh, exhibited a strength that was 80%, 65%, 58%, 37% and 18% greater than the native abdominal wall at 8, 16, 32, and 48 wk post-implantation, respectively. The significant reduction of the polymers’ molecular weight over time demonstrated successful transfer of load-bearing from the mesh to the repaired abdominal wall[66].

Configurations that include a metal component may also add new properties to standard synthetic meshes. Mesh shrinkage, migration, and configuration changes in the host tissue cause severe complications and discomfort after mesh implantation. There is no way to revise an implanted mesh postoperatively except for access to samples that have been explanted because of severe infection, chronic pain and recurrence. However, incorporation of small iron particles into the polymer provides an effective option for noninvasive revision using magnetic resonance imaging (MRI)[67]. Another promising metal to improve mesh performance is nitinol. Nitinol-containing memory frame mesh is a valuable tool to achieve complete deployment in transinguinal preperitoneal repair for inguinal hernias that offers an acceptable morbidity and a low recurrence rate[68].

Coatings are another effective method to modify the properties of synthetic meshes. A titanium-coated PP mesh was associated with less postoperative pain in the short term, lower analgesic consumption and shorter convalescence compared with the Parietex composite mesh[69]. Intraperitoneal implantation of PP meshes is not recommended because of the likeliness of inducing intense adhesion and intestinal fistula. A PP mesh coated with poly(L-lactic acid) exhibited an additional property of anti-adhesion in a rat model[70]. Extracellular matrix-coated PP meshes attenuated the pro-inflammatory response with reduced cell accumulation, fewer foreign body giant cells and decreased collagen density without changes in the mechanical properties of the mesh[71,72]. Chitosan-coated PP meshes elicited preferential attachment of myoblasts over fibroblast attachment in vitro, which was associated with the restoration of functional skeletal muscle with histomorphological characteristics that resembled native muscle in vivo[73]. Degradable drug delivery coatings with incorporated antibiotics provide a specific approach to reduce post-surgical infections[74]. These promising laboratory and animal trial results may be incorporated in clinical practice in the future.

The use of electro-spun nanofibers of various polymers as tissue scaffolds in hernia repair has been an active research topic in recent years. Electro-spun materials feature three-dimensional nanofibrous structure with high surface-to-volume ratios and high porosity with high pore-interconnectivity that are similar to the native extracellular matrix (ECM). Drugs and growth factors for the prevention of incisional hernia formation have also been incorporated into electro-spun nanofibers[75]. Recent research revealed that PET and PET/chitosan electro-spun meshes performed well during incisional hernia surgery. However, the formation of foreign body granuloma in response to electro-spun structures was greater than when conventional meshes were used[76]. Further studies are required to elucidate the mechanisms that underlie the interactions between cells/tissues and nanofibrous materials.

CONCLUSION

Large-pore, monofilament, lightweight synthetic meshes are the current standard of practice. However, the risk of infection and other complications associated with the use of meshes are inevitable. An ideal synthetic mesh should consist of a monofilamentous large-pore structure with anisotropic mechanical properties that are similar to the native properties of the healthy host tissue and composed of a highly biocompatible raw material with long-term stability. An optimal mesh for intraperitoneal use must resist visceral adhesions to limit the risk of bowel obstruction and intestinal fistula. The use of innovative raw materials or coatings of currently available raw materials are promising approaches to realize these ideals. The individual response of the patient influences the local response after mesh implantation. Therefore, a thorough understanding of the biological processes of tissue formation and remodeling in the context of wound-healing processes after hernia repair is needed.

REFERENCES

1 Scott NW, McCormack K, Graham P, Go PM, Ross SJ, Grant AM. Open mesh versus non-mesh for repair of femoral and inguinal hernia. Cochrane Database Syst Rev 2002; (4): CD002197 [PMID: 12519568]

2 den Hartog D, Dur AH, Tuinebreijer WE, Kreis RW. Open surgical procedures for incisional hernias. Cochrane Database Syst Rev 2008; (3): CD006438 [PMID: 18646155 DOI: 10.1002/14651858.CD006438.pub2]

3 Zuvela M, Galun D, Djurić-Stefanović A, Palibrk I, Petrović M, Milićević M. Central rupture and bulging of low-weight polypropylene mesh following recurrent incisional sublay hernioplasty. Hernia 2014; 18: 135-140 [PMID: 24309998 DOI: 10.1007/s10029-013-1197-1]

4 Sanders DL, Waydia S. A systematic review of randomised control trials assessing mesh fixation in open inguinal hernia repair. Hernia 2014; 18: 165-176 [PMID: 23649403 DOI: 10.1007/s10029-013-1093-8]

5 Gutlic N, Rogmark P, Nordin P, Petersson U, Montgomery A. Impact of Mesh Fixation on Chronic Pain in Total Extraperitoneal Inguinal Hernia Repair (TEP): A Nationwide Register-based Study. Ann Surg 2015; Epub ahead of print [PMID: 26135697 DOI: 10.1097/SLA.0000000000001306]

6 Deligiannidis N, Papavasiliou I, Sapalidis K, Kesisoglou I, Papavramidis S, Gamvros O. The use of three different mesh materials in the treatment of abdominal wall defects. Hernia 2002; 6: 51-55 [PMID: 12152638 DOI: 10.1007/s10029-002-0054-4]

7 Klink CD, Binnebösel M, Kaemmer D, Schachtrupp A, Fiebeler A, Anurov M, Schumpelick V, Klinge U. Comet-tail-like inflammatory infiltrate to polymer filaments develops in tension-free conditions. Eur Surg Res 2011; 46: 73-81 [PMID: 21196740 DOI: 10.1159/000322250]

8 Gerullis H, Georgas E, Borós M, Klosterhalfen B, Eimer C, Arndt C, Otto S, Barski D, Ysebaert D, Ramon A, Otto T. Inflammatory reaction as determinant of foreign body reaction is an early and susceptible event after mesh implantation. Biomed Res Int 2014; 2014: 510807 [PMID: 24783209 DOI: 10.1155/2014/510807]

9 Klinge U, Dietz U, Fet N, Klosterhalfen B. Characterisation of the cellular infiltrate in the foreign body granuloma of textile meshes with its impact on collagen deposition. Hernia 2014; 18: 571-578 [PMID: 24500375 DOI: 10.1007/s10029-014-1220-1]

10 Kourtzelis I, Rafail S, DeAngelis RA, Foukas PG, Ricklin D, Lambris JD. Inhibition of biomaterial-induced complement activation attenuates the inflammatory host response to implantation. FASEB J 2013; 27: 2768-2776 [PMID: 23558338 DOI: 10.1096/fj.12-225888]

11 Orenstein SB, Saberski ER, Klueh U, Kreutzer DL, Novitsky YW. Effects of mast cell modulation on early host response to implanted synthetic meshes. Hernia 2010; 14: 511-516 [PMID: 20526725 DOI: 10.1007/s10029-010-0680-1]

12 Vaz M, Krebs RK, Trindade EN, Trindade MR. Fibroplasia after polypropylene mesh implantation for abdominal wall hernia repair in rats. Acta Cir Bras 2009; 24: 19-25 [PMID: 19169537 DOI: 10.1590/s0102-86502009000100005]

13 Majercik S, Tsikitis V, Iannitti DA. Strength of tissue attachment to mesh after ventral hernia repair with synthetic composite mesh in a porcine model. Surg Endosc 2006; 20: 1671-1674 [PMID: 17001442 DOI: 10.1007/s00464-005-0660-1]

14 Wood AJ, Cozad MJ, Grant DA, Ostdiek AM, Bachman SL, Grant SA. Materials characterization and histological analysis of explanted polypropylene, PTFE, and PET hernia meshes from an individual patient. J Mater Sci Mater Med 2013; 24: 1113-1122 [PMID: 23371769 DOI: 10.1007/s10856-013-4872-y]

15 Baktir A, Dogru O, Girgin M, Aygen E, Kanat BH, Dabak DO, Kuloglu T. The effects of different prosthetic materials on the formation of collagen types in incisional hernia. Hernia 2013; 17: 249-253 [PMID: 22907153 DOI: 10.1007/s10029-012-0979-1]

16 Orenstein SB, Saberski ER, Kreutzer DL, Novitsky YW. Comparative analysis of histopathologic effects of synthetic meshes based on material, weight, and pore size in mice. J Surg Res 2012; 176: 423-429 [PMID: 22099590 DOI: 10.1016/j.jss.2011.09.031]

17 Grigoryuk AA, Turmova EP. Effects of polypropylene and polytetrafluoroethylene prostheses for abdominal plasty on local and systemic cytokine production. Bull Exp Biol Med 2014; 156: 530-534 [PMID: 24771444 DOI: 10.1007/s10517-014-2391-6]

18 Klink CD, Junge K, Binnebösel M, Alizai HP, Otto J, Neumann UP, Klinge U. Comparison of long-term biocompability of PVDF and PP meshes. J Invest Surg 2011; 24: 292-299 [PMID: 22047202 DOI: 10.3109/08941939.2011.589883]

19 Pereira-lucena CG, Artigiani Neto R, de Rezende DT, Lopes-Filho Gde J, Matos D, Linhares MM. Early and late postoperative inflammatory and collagen deposition responses in three different meshes: an experimental study in rats. Hernia 2014; 18: 563-570 [PMID: 24370605 DOI: 10.1007/s10029-013-1206-4]

20 Canuto RA, Saracino S, Oraldi M, Festa V, Festa F, Muzio G, Chiaravalloti A. Colonization by human fibroblasts of polypropylene prosthesis in a composite form for hernia repair. Hernia 2013; 17: 241-248 [PMID: 22996952 DOI: 10.1007/s10029-012-0996-0]

21 Weyhe D, Hoffmann P, Belyaev O, Mros K, Muller C, Uhl W, Schmitz F. The role of TGF-beta1 as a determinant of foreign body reaction to alloplastic materials in rat fibroblast cultures: comparison of different commercially available polypropylene meshes for hernia repair. Regul Pept 2007; 138: 10-14 [PMID: 16973225 DOI: 10.1016/j.regpep.2006.07.010]

22 Raptis DA, Vichova B, Breza J, Skipworth J, Barker S. A comparison of woven versus nonwoven polypropylene (PP) and expanded versus condensed polytetrafluoroethylene (PTFE) on their intraperitoneal incorporation and adhesion formation. J Surg Res 2011; 169: 1-6 [PMID: 20400113 DOI: 10.1016/j.jss.2009.12.014]

23 Reis Pdos S, Chagas VL, Silva JM, Silva PC, Jamel N, Schanaider A. Nonwoven polypropylene prosthesis in large abdominal wall defects in rats. Acta Cir Bras 2012; 27: 671-680 [PMID: 23033127 DOI: 10.1590/s0102-86502012001000002]

24 Asarias JR, Nguyen PT, Mings JR, Gehrich AP, Pierce LM. Influence of mesh materials on the expression of mediators involved in wound healing. J Invest Surg 2011; 24: 87-98 [PMID: 21345009 DOI: 10.3109/08941939.2010.548904]

25 Amid PK. Classification of biomaterials and their related complications in abdominal wall hernia surgery. Hernia 1997; 1: 15–21 [DOI: 10.1007/BF02426382]

26 Klinge U, Klosterhalfen B. Modified classification of surgical meshes for hernia repair based on the analyses of 1,000 explanted meshes. Hernia 2012; 16: 251-258 [PMID: 22562353 DOI: 10.1007/s10029-012-0913-6]

27 Klosterhalfen B, Klinge U. Retrieval study at 623 human mesh explants made of polypropylene--impact of mesh class and indication for mesh removal on tissue reaction. J Biomed Mater Res B Appl Biomater 2013; 101: 1393-1399 [PMID: 24591221 DOI: 10.1002/jbmb.32958]

28 Greca FH, Souza-Filho ZA, Giovanini A, Rubin MR, Kuenzer RF, Reese FB, Araujo LM. The influence of porosity on the integration histology of two polypropylene meshes for the treatment of abdominal wall defects in dogs. Hernia 2008; 12: 45-49 [PMID: 17823771 DOI: 10.1007/s10029-007-0276-6]

29 Jerabek J, Novotny T, Vesely K, Cagas J, Jedlicka V, Vlcek P, Capov I. Evaluation of three purely polypropylene meshes of different pore sizes in an onlay position in a New Zealand white rabbit model. Hernia 2014; 18: 855-864 [PMID: 25033941 DOI: 10.1007/s10029-014-1278-9]

30 Mühl T, Binnebösel M, Klinge U, Goedderz T. New objective measurement to characterize the porosity of textile implants. J Biomed Mater Res B Appl Biomater 2008; 84: 176-183 [PMID: 17497684 DOI: 10.1002/jbm.b.30859]

31 Brown CN, Finch JG. Which mesh for hernia repair? Ann R Coll Surg Engl 2010; 92: 272-278 [PMID: 20501011 DOI: 10.1308/003588410X12664192076296]

32 Cobb WS, Burns JM, Peindl RD, Carbonell AM, Matthews BD, Kercher KW, Heniford BT. Textile analysis of heavy weight, mid-weight, and light weight polypropylene mesh in a porcine ventral hernia model. J Surg Res 2006; 136: 1-7 [PMID: 16996087 DOI: 10.1016/j.jss.2006.05.022]

33 Conze J, Junge K, Weiss C, Anurov M, Oettinger A, Klinge U, Schumpelick V. New polymer for intra-abdominal meshes--PVDF copolymer. J Biomed Mater Res B Appl Biomater 2008; 87: 321-328 [PMID: 18435400 DOI: 10.1002/jbm.b.31106]

34 Lake SP, Ray S, Zihni AM, Thompson DM, Gluckstein J, Deeken CR. Pore size and pore shape--but not mesh density--alter the mechanical strength of tissue ingrowth and host tissue response to synthetic mesh materials in a porcine model of ventral hernia repair. J Mech Behav Biomed Mater 2015; 42: 186-197 [PMID: 25486631 DOI: 10.1016/j.jmbbm.2014.11.011]

35 Coda A, Lamberti R, Martorana S. Classification of prosthetics used in hernia repair based on weight and biomaterial. Hernia 2012; 16: 9-20 [PMID: 21837484 DOI: 10.1007/s10029-011-0868-z]

36 Novitsky YW, Cristiano JA, Harrell AG, Newcomb W, Norton JH, Kercher KW, Heniford BT. Immunohistochemical analysis of host reaction to heavyweight-, reduced-weight-, and expanded polytetrafluoroethylene (ePTFE)-based meshes after short- and long-term intraabdominal implantations. Surg Endosc 2008; 22: 1070-1076 [PMID: 18188649 DOI: 10.1007/s00464-007-9737-3]

37 Conze J, Kingsnorth AN, Flament JB, Simmermacher R, Arlt G, Langer C, Schippers E, Hartley M, Schumpelick V. Randomized clinical trial comparing lightweight composite mesh with polyester or polypropylene mesh for incisional hernia repair. Br J Surg 2005; 92: 1488-1493 [PMID: 16308855 DOI: 10.1002/bjs.5208]

38 Agarwal BB, Agarwal KA, Sahu T, Mahajan KC. Traditional polypropylene and lightweight meshes in totally extraperitoneal inguinal herniorrhaphy. Int J Surg 2010; 8: 44-47 [PMID: 19853672 DOI: 10.1016/j.ijsu.2009.08.014]

39 Klosterhalfen B, Junge K, Klinge U. The lightweight and large porous mesh concept for hernia repair. Expert Rev Med Devices 2005; 2: 103-117 [PMID: 16293033 DOI: 10.1586/17434440.2.1.103]

40 Weyhe D, Belyaev O, Müller C, Meurer K, Bauer KH, Papapostolou G, Uhl W. Improving outcomes in hernia repair by the use of light meshes--a comparison of different implant constructions based on a critical appraisal of the literature. World J Surg 2007; 31: 234-244 [PMID: 17180568 DOI: 10.1007/s00268-006-0123-4]

41 Weyhe D, Schmitz I, Belyaev O, Grabs R, Müller KM, Uhl W, Zumtobel V. Experimental comparison of monofile light and heavy polypropylene meshes: less weight does not mean less biological response. World J Surg 2006; 30: 1586-1591 [PMID: 16855805 DOI: 10.1007/s00268-005-0601-0]

42 Muftuoglu MA, Gungor O, Odabasi M, Ekinci O, Teyyareci A, Sekmen U, Saglam A. The comparison of heavyweight mesh and lightweight mesh in an incisional animal model. Hernia 2010; 14: 397-400 [PMID: 20229105 DOI: 10.1007/s10029-010-0647-2]

43 Zogbi L, Trindade EN, Trindade MR. Comparative study of shrinkage, inflammatory response and fibroplasia in heavyweight and lightweight meshes. Hernia 2013; 17: 765-772 [PMID: 23344668 DOI: 10.1007/s10029-013-1046-2]

44 ASTM D4850. Standard Terminology Relating to Fabrics and Fabric Test Methods. ASTM International, West Conshohocken, PA, 2013 [DOI: 10.1520/D4850]

45 Hollinsky C, Sandberg S. Measurement of the tensile strength of the ventral abdominal wall in comparison with scar tissue. Clin Biomech (Bristol, Avon) 2007; 22: 88-92 [PMID: 16904247 DOI: 10.1016/j.clinbiomech.2006.06.002]

46 Williams JF, Kirkpatrick J, Syme GA. Force measurement in the abdominal wall. Biomed Eng 1975; 10: 181-183 [PMID: 123792]

47 Cobb WS, Burns JM, Kercher KW, Matthews BD, James Norton H, Todd Heniford B. Normal intraabdominal pressure in healthy adults. J Surg Res 2005; 129: 231-235 [PMID: 16140336 DOI: 10.1016/j.jss.2005.06.015]

48 Deeken CR, Abdo MS, Frisella MM, Matthews BD. Physicomechanical evaluation of polypropylene, polyester, and polytetrafluoroethylene meshes for inguinal hernia repair. J Am Coll Surg 2011; 212: 68-79 [PMID: 21115372 DOI: 10.1016/j.jamcollsurg.2010.09.012]

49 Pott PP, Schwarz ML, Gundling R, Nowak K, Hohenberger P, Roessner ED. Mechanical properties of mesh materials used for hernia repair and soft tissue augmentation. PLoS One 2012; 7: e46978 [PMID: 23071685 DOI: 10.1371/journal.pone.0046978]

50 Klinge U, Klosterhalfen B, Conze J, Limberg W, Obolenski B, Ottinger AP, Schumpelick V. Modified mesh for hernia repair that is adapted to the physiology of the abdominal wall. Eur J Surg 1998; 164: 951-960 [PMID: 10029391 DOI: 10.1080/110241598750005138]

51 DuBay DA, Wang X, Adamson B, Kuzon WM, Dennis RG, Franz MG. Mesh incisional herniorrhaphy increases abdominal wall elastic properties: a mechanism for decreased hernia recurrences in comparison with suture repair. Surgery 2006; 140: 14-24 [PMID: 16857438 DOI: 10.1016/j.surg.2006.01.007]

52 Bellón JM, Rodríguez M, García-Honduvilla N, Gómez-Gil V, Pascual G, Buján J. Comparing the behavior of different polypropylene meshes (heavy and lightweight) in an experimental model of ventral hernia repair. J Biomed Mater Res B Appl Biomater 2009; 89: 448-455 [PMID: 18837449 DOI: 10.1002/jbm.b.31234]

53 Petro CC, Nahabet EH, Criss CN, Orenstein SB, von Recum HA, Novitsky YW, Rosen MJ. Central failures of lightweight monofilament polyester mesh causing hernia recurrence: a cautionary note. Hernia 2015; 19: 155-159 [PMID: 24659227 DOI: 10.1007/s10029-014-1237-5]

54 Lintin LA, Kingsnorth AN. Mechanical failure of a lightweight polypropylene mesh. Hernia 2014; 18: 131-133 [PMID: 22824989 DOI: 10.1007/s10029-012-0959-5]

55 Melman L, Jenkins ED, Hamilton NA, Bender LC, Brodt MD, Deeken CR, Greco SC, Frisella MM, Matthews BD. Histologic and biomechanical evaluation of a novel macroporous polytetrafluoroethylene knit mesh compared to lightweight and heavyweight polypropylene mesh in a porcine model of ventral incisional hernia repair. Hernia 2011; 15: 423-431 [PMID: 21279663 DOI: 10.1007/s10029-011-0787-z]

56 Eliason BJ, Frisella MM, Matthews BD, Deeken CR. Effect of repetitive loading on the mechanical properties of synthetic hernia repair materials. J Am Coll Surg 2011; 213: 430-435 [PMID: 21705242 DOI: 10.1016/j.jamcollsurg.2011.05.018]

57 Alizai PH, Schmid S, Otto J, Klink CD, Roeth A, Nolting J, Neumann UP, Klinge U. Biomechanical analyses of prosthetic mesh repair in a hiatal hernia model. J Biomed Mater Res B Appl Biomater 2014; 102: 1485-1495 [PMID: 24599834 DOI: 10.1002/jbm.b.33128]

58 Deeken CR, Thompson DM, Castile RM, Lake SP. Biaxial analysis of synthetic scaffolds for hernia repair demonstrates variability in mechanical anisotropy, non-linearity and hysteresis. J Mech Behav Biomed Mater 2014; 38: 6-16 [PMID: 24997427 DOI: 10.1016/j.jmbbm.2014.06.001]

59 Li X, Kruger JA, Jor JW, Wong V, Dietz HP, Nash MP, Nielsen PM. Characterizing the ex vivo mechanical properties of synthetic polypropylene surgical mesh. J Mech Behav Biomed Mater 2014; 37: 48-55 [PMID: 24942626 DOI: 10.1016/j.jmbbm.2014.05.005]

60 Anurov MV, Titkova SM, Oettinger AP. Biomechanical compatibility of surgical mesh and fascia being reinforced: dependence of experimental hernia defect repair results on anisotropic surgical mesh positioning. Hernia 2012; 16: 199-210 [PMID: 21909779 DOI: 10.1007/s10029-011-0877-y]

61 Ulrich D, Edwards SL, White JF, Supit T, Ramshaw JA, Lo C, Rosamilia A, Werkmeister JA, Gargett CE. A preclinical evaluation of alternative synthetic biomaterials for fascial defect repair using a rat abdominal hernia model. PLoS One 2012; 7: e50044 [PMID: 23185528 DOI: 10.1371/journal.pone.0050044]

62 Schäfer-Nolte F, Hennecke K, Reimers K, Schnabel R, Allmeling C, Vogt PM, Kuhbier JW, Mirastschijski U. Biomechanics and biocompatibility of woven spider silk meshes during remodeling in a rodent fascia replacement model. Ann Surg 2014; 259: 781-792 [PMID: 23873006 DOI: 10.1097/SLA.0b013e3182917677]

63 Kim M, Oommen B, Ross SW, Lincourt AE, Matthews BD, Heniford BT, Augenstein VA. The current status of biosynthetic mesh for ventral hernia repair. Surg Technol Int 2014; 25: 114-121 [PMID: 25396323]

64 Powell BS, Wandrey D, Voeller GR. A technique for placement of a bioabsorbable prosthesis with fibrin glue fixation for reinforcement of the crural closure during hiatal hernia repair. Hernia 2013; 17: 81-84 [PMID: 22581201 DOI: 10.1007/s10029-012-0915-4]

65 Symeonidis D, Efthimiou M, Koukoulis G, Athanasiou E, Mamaloudis I, Tzovaras G. Open inguinal hernia repair with the use of polyglycolic acid/trimethylene carbonate absorbable mesh: a critical update of the long-term results. Hernia 2013; 17: 85-87 [PMID: 23138860 DOI: 10.1007/s10029-012-1016-0]

66 Martin DP, Badhwar A, Shah DV, Rizk S, Eldridge SN, Gagne DH, Ganatra A, Darois RE, Williams SF, Tai HC, Scott JR. Characterization of poly-4-hydroxybutyrate mesh for hernia repair applications. J Surg Res 2013; 184: 766-773 [PMID: 23582230 DOI: 10.1016/j.jss.2013.03.044]

67 Ciritsis A, Hansen NL, Barabasch A, Kuehnert N, Otto J, Conze J, Klinge U, Kuhl CK, Kraemer NA. Time-dependent changes of magnetic resonance imaging-visible mesh implants in patients. Invest Radiol 2014; 49: 439-444 [PMID: 24651663 DOI: 10.1097/RLI.0000000000000051]

68 Berrevoet F, Vanlander A, Bontinck J, Troisi RI. Open preperitoneal mesh repair of inguinal hernias using a mesh with nitinol memory frame. Hernia 2013; 17: 365-371 [PMID: 23686406 DOI: 10.1007/s10029-013-1110-y]

69 Köckerling F, Schug-Pass C. What do we know about titanized polypropylene meshes? An evidence-based review of the literature. Hernia 2014; 18: 445-457 [PMID: 24253381 DOI: 10.1007/s10029-013-1187-3]

70 Zhang Z, Zhang T, Li J, Ji Z, Zhou H, Zhou X, Gu N. Preparation of poly(L-lactic acid)-modified polypropylene mesh and its antiadhesion in experimental abdominal wall defect repair. J Biomed Mater Res B Appl Biomater 2014; 102: 12-21 [PMID: 23661568 DOI: 10.1002/jbm.b.32947]

71 Wolf MT, Carruthers CA, Dearth CL, Crapo PM, Huber A, Burnsed OA, Londono R, Johnson SA, Daly KA, Stahl EC, Freund JM, Medberry CJ, Carey LE, Nieponice A, Amoroso NJ, Badylak SF. Polypropylene surgical mesh coated with extracellular matrix mitigates the host foreign body response. J Biomed Mater Res A 2014; 102: 234-246 [PMID: 23873846 DOI: 10.1002/jbm.a.34671]

72 Faulk DM, Londono R, Wolf MT, Ranallo CA, Carruthers CA, Wildemann JD, Dearth CL, Badylak SF. ECM hydrogel coating mitigates the chronic inflammatory response to polypropylene mesh. Biomaterials 2014; 35: 8585-8595 [PMID: 25043571 DOI: 10.1016/j.biomaterials.2014.06.057]

73 Udpa N, Iyer SR, Rajoria R, Breyer KE, Valentine H, Singh B, McDonough SP, Brown BN, Bonassar LJ, Gao Y. Effects of chitosan coatings on polypropylene mesh for implantation in a rat abdominal wall model. Tissue Eng Part A 2013; 19: 2713-2723 [PMID: 23859182 DOI: 10.1089/ten.TEA.2012.0739]

74 Letouzey V, Lavigne JP, Garric X, Coudane J, de Tayrac R, Callaghan DO. Is degradable antibiotic coating for synthetic meshes provide protection against experimental animal infection after fascia repair? J Biomed Mater Res B Appl Biomater 2012; 100: 471-479 [PMID: 22102417 DOI: 10.1002/jbm.b.31973]

75 Plencner M, East B, Tonar Z, Otáhal M, Prosecká E, Rampichová M, Krejčí T, Litvinec A, Buzgo M, Míčková A, Nečas A, Hoch J, Amler E. Abdominal closure reinforcement by using polypropylene mesh functionalized with poly-ε-caprolactone nanofibers and growth factors for prevention of incisional hernia formation. Int J Nanomedicine 2014; 9: 3263-3277 [PMID: 25031534 DOI: 10.2147/IJN.S63095]

76 Veleirinho B, Coelho DS, Dias PF, Maraschin M, Pinto R, Cargnin-Ferreira E, Peixoto A, Souza JA, Ribeiro-do-Valle RM, Lopes-da-Silva JA. Foreign body reaction associated with PET and PET/chitosan electrospun nanofibrous abdominal meshes. PLoS One 2014; 9: e95293 [PMID: 24740104 DOI: 10.1371/journal.pone.0095293]

P-Reviewer: Dmochowski R, Garg P, Vilallonga R S-Editor: Ji FF L-Editor: E-Editor:

Figure 1 Crucial mesh parameters for selection of an ideal mesh. The impact of mesh implants on clinical results is the current subject of much litigation in the field of stress urinary incontinence and pelvic prolapse, and some manufacturers were sued because of allegedly defective implants. However, many other factors besides mesh parameters must be considered in evaluations of the overall outcome of an intervention, including the patient’s constitution, the selection of a proper operation technique and the operation performance, which are essential for the success or failure of a therapy.

Figure 2 Textile structures from left to right: Knitted structure, warp-knitted structure, nonwoven structure, and woven structure.

Figure 3 Comparison of the in vivo stability of the surface of polypropylene and polyvinylidenefluoride four weeks after implantation[18]. PP: Polypropylene; PVDF: Polyvinylidenefluoride.

A                                 B

Figure 4 Comparison of the textile porosity (A) and the effective porosity (B).

Table 1 Definitions of the knit, warp-knit, nonwoven and woven textile structures

| Textile structure | Definition |
| --- | --- |
| Knitted fabric | Knitted fabric consists of a number of consecutive rows of loops, called stitches. Knitted structures are manufactures from single yarn systems. Thus, knitted structures can be ribbed off. Trimming of knitted structures often leads to a complete falling apart |
| Warp-knitted fabric | Warp-knitted fabric consists of a number of consecutive courses of loops, called stitches. Warp-knitted structures are manufactures from multi yarn systems whereby the number of separate strands of yarn equals the number of stitches in a row. In contrast to knitted structures warp-knitted structures can be trimmed and sewed |
| Nonwoven fabric | Nonwoven fabric consists of non orientated or to a certain degree orientated staple or endless fibers. After the nonwoven formation the structure needs to be bonded which either is realised by mechanical, thermal or chemical bonding |
| Woven fabric | Woven structures consist of two distinct sets of yarns or threads which are interlaced at right angles to form a fabric |

Table 2 Essential properties of the knit, warp-knit, nonwoven and woven textile structures

| Textile structure | Porosity (macropores) | Elasticity | Mechanical behaviour | Trim-ability |
| --- | --- | --- | --- | --- |
| Knitted fabric | ++ | ++ | Anisotropic | -- |
| Warp-knitted fabric | ++ | ++ | Isotropic, anisotropic | ++ |
| Nonwoven fabric | - | - | Isotropic | ++ |
| Woven fabric | - | -- | Isotropic | ++ |

Table 3 Definitions of mechanical mesh properties based on the definitions given by the American Society for Testing and Materials[44]

| Property | Definition |
| --- | --- |
| Tensile strength | Tensile strength is the maximum force that can be applied to a mesh without tearing or breaking of the mesh. The tensile strength is measured in Newton (N) and is usually given in relation to the clamping width as Newton per centimeter (N/cm) |
| Burst strength | The burst strength is the maximum uniformly distributed pressure applied at right angle to its surface that a material will withstand under standardized conditions. The burst strength is given in pressure per unit area (Pa/cm²) |
| Elasticity (elastic elongation) | Elasticity (elastic elongation) is the property of a material whereby it changes its shape and size under the action of opposing forces (%), but recovers its original configuration when the forces are removed. In contrast, to the elastic elongation the plastic elongation indicates the elongation ratio which does not recover after unloading the structure |
| Stiffness | Stiffness can be expressed as ratio of steadily increasing or decreasing force acting on a deformable elastic material to the resulting displacement or deformation. Stiffness is a crucial aspect that reflects the drapablity of a textile structure, means the ability of a textile structure to be adapted to a 3-dimensional geometry |

Table 4 Essential properties of hernia meshes used for groin and abdominal wall hernia repair

| Property | Recommendation |
| --- | --- |
| Tensile strength (abdominal wall) | 22 N/cm (cranial/caudal ) |
|  | 32 N/cm (lateral) |
| Tensile strength (groin) | 16 N/cm |
| Elongation | 20%-40% |
| Orientation | No specific orientation for meshes with isotropic properties |
|  | For meshes with anisotropic properties: orientation in the appropriate direction to match the physiological stretchability |
| Pore size | Depending on the used raw material and the foreign body reaction, respectively. To achieve a high effective porosity: for PP meshes a pore size ≥ 1000 µm should be used; for PVDF meshes a pore size ≥ 600 µm should be used |

PVDF: Polyvinylidenefluoride.