**Name of Journal: World Journal of Cardiology**

**Manuscript NO: 42942**

**Manuscript Type: Review**

**Percutaneous devices for left atrial appendage occlusion: A contemporary review**

Moussa Pacha H *et al*. Left atrial appendage occlusion

Homam Moussa Pacha, Yasser Al-khadra, Mohamad Soud, Fahed Darmoch, Abdulghani Moussa Pacha, M Chadi Alraies

**Homam Moussa Pacha,** Mohamad Soud, MedStar Washington Hospital Center, Washington, DC 20010, United States

**Yasser Al-khadra,** Cleveland Clinic, Medicine Institute, Cleveland, OH 44195, United States

**Fahed Darmoch,** Beth Israel Deaconess Medical center/Harvard medical school, Boston, MA 02215, United States

**Abdulghani Moussa Pacha,**University Hospital of Mainz, Mainz 55122, Germany

**M Chadi Alraies,** Wayne State University, Detroit Medical Center, Detroit Heart Hospital, Detroit, MI 48201, United States

**ORCID number:** Homam Moussa Pacha (0000-0003-4403-0941); Yasser Al-khadra (0000-0002-6618-1401); Mohamad Soud (0000-0003-2265-0875);Fahed Darmoch (0000-0003-2094-7368); Abdulghani Moussa Pacha (0000-0001-9159-2811); M Chadi Alraies (0000-0002-7874-4566).

**Author contributions:** All authors equally contributed to this paper with conception and design of the study, literature review and analysis, drafting and critical revision and editing, and final approval of the final version.

**Conflict-of-interest statement:** No potential conflicts of interest.

**Open-Access:** This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

**Manuscript source:** Unsolicited manuscript

**Corresponding author: M Chadi Alraies, MD,** Wayne State University, Detroit Medical Center, Detroit Heart Hospital, 311 Mack Ave, Detroit, MI 48201, United States. makraies@dmc.org

**Telephone:** +1-216-2550008

**Fax:** +1-313-7453500

**Received:** October 16, 2018

**Peer-review started:** October 16, 2018

**First decision:** October 26, 2018

**Revised:** December 12, 2018

**Accepted:** December 17, 2018

**Article in press:** December 17, 2018

**Published online:** February 26, 2019

**Abstract**

Patient with atrial fibrillation (AF) are at risk of developing stroke with the left atrial appendage (LAA) being the most common site for thrombus formation. If left untreated, AF is associated with 4 to 5 folds increase in the risk of ischemic stroke in all age groups. About 5% to 15% of AF patients have atrial thrombi on transesophageal echocardiography, and 91% of those thrombi are located in the LAA in patient with nonrheumatic AF. Although oral anticoagulants are the gold-standard treatment for stroke prevention in patients with non-valvular AF, some patients are at high risk of bleeding and deemed not candidates for anticoagulation. Therefore, LAA occlusion (LAAO) has emerged as alternative approach for stroke prevention in those patients. Surgical LAAO is associated with high rate of unsuccessful closure and recommended only in patients with AF and undergoing cardiac surgery. Percutaneous LAAO uses transvenous access with trans-septal puncture and was first tested using the PLAATO device. Watchman is the most common and only Food and Drug Administration (FDA) approved device for LAAO. LAAO using Watchman device is non-inferior to warfarin therapy in preventing ischemic stroke/systemic thromboembolism. However, it is associated with lower rates of hemorrhagic stroke, bleeding and death. Amplatzer is another successful LAAO device that has CE mark and is waiting for FDA approval. Optimal antithrombotic therapy post LAAO is still under debate and highly patient-specific. The aim of this paper is to systematically review the current literature to evaluate the efficacy and safety of different LAAO devices.

**Key words**: Left atrial appendage; Atrial fibrillation; Anticoagulation; Stroke; Mortality

**© The Author(s) 2019.** Published by Baishideng Publishing Group Inc. All rights reserved.

**Core tip:** Left atrial appendage occlusion (LAAO) is a reasonable alternative approach that is used in patients with atrial fibrillation who are not candidates for anticoagulation. A number of key trials have shown that Watchman device is non-inferior to warfarin therapy in preventing ischemic stroke/systemic thromboembolism. However, it is associated with lower rates of hemorrhagic stroke, bleeding and death. Multiple retrospective and prospective studies of Amplatzer device (ACP and Amulet) reported high success rates in device implantation and stroke prevention. Our objective is to consolidate the current literature to better delineate the safety, efficacy and indication of LAAO for stroke prevention.

**Citation:** Moussa Pacha H, Al-khadra Y, Soud M, Darmoch F, Moussa Pacha A, Alraies MC. Percutaneous devices for left atrial appendage occlusion: A contemporary review. *World J Cardiol* 2019; 11(2): 57-70

**URL:** https://www.wjgnet.com/1949-8462/full/v11/i2/0000.htm

**DOI:** https://dx.doi.org/10.4330/wjc.v11.i2.0000

**Introduction**

Atrial fibrillation (AF) affects 2.7 to 6.1 million in the United states and 33.5 million worldwide[1-3]. The projected prevalence of AF in the United states is expected to be 12.1 million by 2030[4]. AF-associated stroke is the most feared complication and the leading cause of disability in the United states[5]. If left untreated, AF is associated with 4 to 5 folds increase in the risk of ischemic stroke in all age groups[5,6]. Furthermore, AF is associated with increased risk of extracranial thromboembolic events to the aorta; and renal, mesenteric, and peripheral arteries[7]. The proportion of strokes attributed solely to AF increases with age and may reach up to 23.5%[6,8]. Oral anticoagulants (OACs) remain to be the gold standard treatment for stroke prevention, and their role in preventing AF-related strokes is well established[9,10]. Yet, OACs are contraindicated in a subset of patients who are at high risk of bleeding. As a result, left atrial appendage occlusion (LAAO) has emerged as an alternative approach in this group. In the current article, we present the most updated studies describing safety, efficacy and outcome of different LAAO devices.

**Literature search**

A systematic literature search was conducted using PubMed, EMBASE, and Cochrane Library to identify relevant articles from 1990 to 2018. The following search terms were used: “atrial fibrillation”, “stroke”, “left atrial appendage”, “occlusion” or “closure”, and “percutaneous” or “surgical.” A total of 78 studies were included for review. Of the included studies on LAAO, 3 studies contained surgical LAAO, two contained Atriclip device, two contained Tiger Paw system, 6 contained Lariat device, 4 contained PLAATO device, 19 contained Watchman device, and 12 contained Amplatzer (ACP/Amulet) device.

**Left atrial appendage and thrombus formation**

Left atrial appendage (LAA) is trabeculated long tubular structure that has narrow junction with the venous component of left atrium. it varies greatly in sizes and shapes and has bent or spiral axis in 70% of patients[11]. Anatomically, LAA is best divided into the ostium, neck, and lobar region[12]. In patients with chronic AF, remodeling of LAA leads to dilation, stretching and reduction in pectinate muscle volume[13].

Approximately, 5% to 15% of AF patients have atrial thrombi on Transesophageal echocardiography (TEE)[14-17], and 91% of those thrombi are located in LAA in patients with nonrheumatic AF[18].The reason for LAA predilection for thrombus formation in AF is still not well known. One theory suggests that the extent of LAA filling and emptying is influenced more by changes in the left ventricular (which is impaired in AF) than LAA function[19]. Ventricular filling creates intracavitary suction effect which influences the emptying and filling of left atrium and LAA.

**Imaging assessment of LAA**

Accurate assessment of anatomic LAA characteristics is crucial prior to LAAO due to substantial variations in LAA anatomy that impact device selection and efficacy. TEE is the most widely used imaging tool for periprocedural LAA assessment. It is used for the detection of thrombi in the LA and LAA as well other cardiac masses and thrombi prior to LAAO[12,20]. Features on TEE associated with increased risk of thrombus formation include: reduced LAA flow velocity, spontaneous left atrial contrast, and aortic atheroma[16]. TEE is very important imaging to support fluoroscopy during device implantation. 3D TEE has shown to be more accurate than 2D TEE in LAA assessment and thrombi detection[21,22]; and therefore, it is recommended for the guidance of LAAO[23]. It is used to guide trans-septal puncture, verify catheter and sheath position in the LAA, aid device delivery and positioning, confirm adequate LAA sealing, and detect complications[12]. Follow-up TEE is also recommended after LAAO to reassess the implanted device, confirm complete LAA closure, and rule out complications. Intracardiac echocardiography (ICE) is comparable imaging to TEE for guiding LAAO and performing the tasks typically provided by TEE during implantation. In one study LAA measurements by ICE during LAAO were significantly correlated to angiography and TEE (Pearson correlation coefficient r = 0.94, p < 0.0001 for both)[24].

Multidetector computed tomography is another imaging modality that is used for the assessment of thrombus formation, LAA anatomy and function, device assessment and detection of complications post procedure[12]. It provides 3D images of the heart by using numerous planes at different points in time during the cardiac cycle and has 100% sensitivity for excluding LAA thrombus[25]. However, its use is limited due to ionizing radiation, lower temporal resolution than TEE and inability to perform during device deployment. angiography has been used for in LAA thrombi detection[26]. However, it is expensive and invasive procedure, and rarely used nowadays due to presence of TEE and other less invasive imaging modalities.

**Guideline therapy for stroke prevention**

The 2014 American Heart Association/American College of Cardiology (AHA/ACC) guidelines for management of AF recommends the use of anticoagulation for prevention of thromboembolism when CHA2DS2-VASC score is ≥ 2 (class I (A))[16]. The 2016 European Society of Cardiology (ESC) guidelines differentiate between males and females regarding anticoagulation recommendations[27]. While anticoagulation is class I (A) indication for males with a score ≥ 2 and females with a score ≥ 3, it's considered class IIa (B) indication for males with a score of 1 and females with a score of 2. Both American and European guidelines recommend considering surgical excision of LAA in patients who have AF and undergoing cardiac surgery (class IIb (level of evidence is "C" in AHA/ACC and "B" in the ESC guidelines))[16,27]. While AHA/ACC guidelines have no recommendations for LAAO, the ESC guidelines have class IIb (B) recommendation for LAAO in patients with AF and contra-indications for long-term anticoagulation[27]. Similarly, National Heart Foundation of Australia and the Cardiac Society of Australia and New Zealand state that LAAO may be considered for stroke prevention in patients with non-valvular AF at moderate to high risk of stroke and with contraindications to OAC (GRADE quality of evidence: Low; GRADE strength of recommendation: Strong)[28].

**LAA surgical closures/excision**

Surgical exclusion of LAA is recommended for patients with AF and undergoing concomitant cardiac surgeries. Different surgical methods to isolate LAA include: suture ligation, excision and suture closure, and stapling exclusion with or without excision[29,30]. Surgical isolation of LAA is associated with high rate of unsuccessful closure. For instance, a previous study reported only 40% (55 out of 137) complete LAA closure noted on TEE following surgical closure[30]. Despite that, Friedman *et al*[31] reported a lower risk of readmission for thromboembolism (4.2% *vs* 6.2%, HR = 0.67; 95%CI: 0.56-0.81) and all-cause mortality (17.3% *vs* 23.9%, HR = 0.88; 95%CI: 0.79-0.97) among Medicare patients (age > 65) with AF undergoing concomitant cardiac surgery and surgical LAAO, compared with no surgical LAAO. This the largest study to date supporting the role of surgical LAAO during cardiac surgery as a mean of preventing thromboembolism in patients over the age of 65 with AF.

The Atriclip Device System (Atricure, Inc., West Chester, OH, United States) is a surgical LAA exclusion device composed of self-closing, sterile, implantable clip with a reusable deployment tool (figure 1). It is applied epicardially by either an open surgical or a minimally-invasive technique and placed at the base of the appendage. The clip is made of 2 parallel rigid titanium tubes with elastic nitinol springs covered with a knit-braided polyester sheath (table 1)[32,33]. The EXCLUDE study (Exclusion of LAA with AtriClip Exclusion Device in Patients Undergoing Concomitant Cardiac Surgery) is a nonrandomized multicenter trial that included 70 patients to evaluate the efficacy of Atriclip device[33]. They enrolled adult patients undergoing elective primary cardiac operations via median sternotomy (coronary artery bypass grafting, valve re- pair or replacement, surgical Maze procedures, or atrial septal defect repair) and have CHADS2 > 2. 67 out of 70 patients (95.7%) had successful intraoperative LAA exclusion, and 60 out of 61 patients (98.4%) had successful LAA exclusion seen on computed tomography angiography or TEE imaging after 3 mo[33]. Tiger Paw System (Terumo Cardiovascular Systems, Ann Arbor, Mi, United States) is another LAA exclusion device that is used as a concomitant procedure during open cardiac surgical procedures (figure 1). The device contains implantable fastener of titanium connectors that staples the LAA tissue and is embedded in two rims of silicone that adapts to the LAA morphology and seals the puncture sites (table 1)[34]. Despite its efficacy in achieving complete LAA closure on prior study[34], a class 1 recall from the market by FDA was made in 2015 due to device malfunction[35].

**Lariate device closure system**

Lariat device (SentreHEART, Inc., Redwood City, California) is LAA closure system that is approved by the United States Food and Drug Administration (FDA) for soft tissue closure, but not LAAO (figure 1). It is composed of 15-mm compliant occlusion balloon catheter (EndoCATH), 0.025-inch and 0.035-inch magnet-tipped guidewires (FindrWIRZ), and a 12-F suture delivery device (LARIAT) (table 1). During the procedure, magnet-tipped guidewires are advanced through epicardial and transvenous accesses and connected in the LAA. Then, a suture fashioned as a Lariate or lasso is advanced over the epicardial access guidewire and tightened to occlude LAA base[29,36]. The largest prospective study of Lariate device included patients who: were ≥ 18-year-old; had nonvalvular AF; had CHADS2 ≥ 1; were poor candidate for or failed warfarin therapy; and had a life expectancy of at least 1 year[36]. They reported 95% (81 of 85 patients) complete LAA closure documented on TEE one month after the procedure. 98% of those who underwent TEE (*n =* 65) had complete LAA closure after 1 year, including cases of incomplete closure at earlier time. Complications in the same study were limited to only two cases of severe pericarditis, two cases of strokes, and one case with pericardial effusion[36].Another study demonstrated similar efficacy of the Lariate device for stroke prevention[37]. Dar *et al*[38] demonstrated that LAAO using Lariate device might improve the mechanical function of the left atrium (LA) and reverse LA remodeling based on 2-dimensional speckle tracking echocardiography (a novel method for functional assessment of the LA). However, due to steep learning curve for device deployment (especially epicardial access), LAA leak and lack of direct efficacy comparison with oral anticoagulation, the device was not widely used in the United States[39-42].

**Percutaneous LAA closure**

The most commonly used percutaneous LAAO devices are shown in figure 1 and described in Table 1. Percutaneous LAAO uses transvenous access with trans-septal puncture and was first tested using the Percutaneous LAA Transcatheter Occlusion (PLAATO) device (Appriva Medical Inc., Sunnyvale, CA) in 2001.

***PLAATO device***

This device consists of self-expanding nitinol cage that is covered with polymeric membrane in order to close off blood flow into the LAA (Table 1)[43,44]. It was first tested on 15 patients with non-valvular AF and contraindication to warfarin therapy and are at high risk of thromboembolism based on CHADS2 score[43]. Successful occlusion of LAA was observed in all cases and no device related complications were reported. A larger prospective study enrolled patients using similar inclusion criteria to undergo LAAO using PLAATO device[44]. Similarly, they reported high successful device Implantation in 108 out of 111 patients (97.3%) with only 2 patients developed stroke on follow up (2.2% annual risk of stroke). Subsequently, the European PLAATO2 trial reported successful LAAO in 90% (126 out of 140) of patients with reduction of stroke rate from 6.6% (based on CHADS2 score) to 2.3% per year[45]. Besides, a single center prospective study on 73 cases who had PLAATO device reported death due to device embolization in one patient and implant instability requiring open heart surgery in another one[46]. Interestingly, there was no incidence of stroke for 24 mo of follow-up in the same study. Despite this success, the device was discontinued for unspecified reasons and replaced by Watchman device.

***Watchman device***

The Watchman device (Boston Scientific, Marlborough, MA), is the only FDA-approved percutaneous device for LAAO. The device is composed of self-expanding nitinol frame structure with fixation barbs and a permeable polyester fabric that covers the atrial facing surface of the device (Table 1)[47]. Multiple trials were done to evaluate the safety, efficacy and outcomes of watchmen device.

Pilot study was a non-randomized trial that included 75 patients and was done to assess the feasibility and safety of watchman device[47]. They enrolled adult patients who: had non-valvular AF for 2 years, were eligible for warfarin therapy, and had CHADS2 of at least 1. Although this was the first human trial to evaluate the efficacy and safety of Watchman device, the success rate of LAAO was very high and complications were relatively low. 88% of patients had successful device implantation and 93% of them had complete LAAO. Reported complications included; device embolization in 2 patients, device-related thrombus formation in 4 patients, and transient ischemic attack in 2 patients. There was no reported major strokes or procedure-related mortality.

PROTECT AF study(WATCHMAN LAA System for Embolic Protection in Patients with Atrial Fibrillation)was the first randomized trial to compare the efficacy and safety of LAAO using Watchman device with chronic warfarin therapy in patients with non-valvular AF and had CHADS2 of 1 or more[48]. Exclusion criteria included contraindications to warfarin, chronic warfarin use, LAA thrombus, a patent foramen ovale with atrial septal aneurysm and right-to-left shunt, mobile aortic atheroma, and symptomatic carotid artery disease. This trial enrolled 707 patients from 59 centers worldwide and assigned them randomly to LAAO with Watchman device (*n =* 463) or warfarin therapy (*n =* 244) with INR goal of 2 to 3. Watchman group was treated with warfarin for 45 d after device deployment to allow proper endothelialization. Warfarin was discontinued if TEE showed complete closure or significantly decreased flow around the device. Afterward, patients were given aspirin and clopidogrel for 6 mo followed by lifelong aspirin. At 1065 patient-years (PY) of follow-up (mean follow up 18 mo), Watchman device was non-inferior to warfarin for primary efficacy endpoint of stroke (either ischemic or hemorrhagic), cardiovascular death, or systemic thromboembolism. The Event rates of primary efficacy endpoint were 3% and 4.9% for Watchman and warfarin groups, respectively. Since then, two studies were published with two different follow up period[49,50]. At 2.3 ± 1.1 years (2621 PY), Watchman device continued to be non-inferior to warfarin therapy with 3% and 4.3% event rates of primary efficacy endpoint for Watchman and warfarin groups, respectively[49]. The second trial with 3.8 ± 1.7 years of follow up (2621 PY) showed event rate of 2.3% in the watchman group and 3.8% in the warfarin group (*p =* 0.0348), leading to 40% risk reduction in primary efficacy endpoint with Watchman device[50].

PREVAIL study(Evaluation of the Watchman LAA Closure Device in Patients With Atrial Fibrillation *vs* Long Term Warfarin Therapy) was another randomized trial that assessed the safety and efficacy of Watchman device in patients non-valvular AF[51]. Investigators included a higher risk patients than PROTECT AF (CHADS2 score of 1 plus any of the following higher-risk characteristics: female age ≥ 75 years, baseline ejection fraction ≥ 30% but < 35%, age 65 to 74 years and either diabetes or coronary disease, and age ≥ 65 years with congestive heart failure). Patients were assigned randomly to receive LAAO using Watchman (*n =* 269) or warfarin therapy (*n =* 138) in 2:l ratio. Warfarin and antiplatelet therapy post device implantion was in a similar fashion to PROTECT AF trial. Although non-inferiority criteria was not achieved for overall efficacy endpoint (stroke, systemic embolization or cardiovascualr death), the rate of second efficacy endpoint (stroke or systemic embolization) was 2.5% in the Watchman group and 2% in the warfarin group at 18 mo follow-up, achieving criteria for non-inferiority. Compared to PROTECT AF study, procedural success increased from 90.9% to 95.1% (*p =* 0.04), while all 7-d procedure-related complications (composite of cardiac perforation, pericardial effusion with tamponade, ischemic stroke, device embolization, and other vascular complications) decreased from 8.7% to 4.2% in PREVAIL (*p =* 0.004).

PROTECT AF and PREVAIL results were pooled for patient level meta-analysis and with combined follow-up of 5 years (4343 PY)[52]. The primary efficacy endpoint (stroke, systemic embolization or cardiovascualr death) was similar between LAAO and warfarin groups (2.8 *vs* 3.4 events/100 PY; *p =* 0.27). In subgroup analysis of the same meta-analysis, the rate of all stroke or systemic embolism was similar between both groups (1.7 *vs* 1.8 events/100 PY; *p =* 0.87). However, there was statistically significant decrease in the rates of hemorrhagic stroke (0.17% *vs* 0.87%, *p =* 0.002), disabling / fatal stroke (0.37% *vs* 0.94%, *p =* 0.027), cardiovascular/unexplained death (1.3% *vs* 2.2%, *p =* 0.027), all-cause death (3.6% *vs* 4.9%, *p =* 0.035), and post-procedure bleeding (1.7% *vs* 3.6%, *p =* 0.0003) in LAAO arm when compared with warfarin arm. This meta-analysis underscores the mortality reduction and stroke prevention, patrticularly hemorrhagic stroke, associated with LAAO using Wathcman device.

Continued Access to PROTECT AF (CAP)[53] and Continued Access to PREVAIL (CAP2)[54] Registries were designed to treat patients with similar baseline characteristics and according to same protocols after PROTECT AF and PREVAIL trials enrollment had been completed. Procedural performance and associated medications were identical in each registry. However, registries did not mandate 1-year neurological assessment. A Meta-analysis of 2406 patients from the PROTECT AF and PREVAIL trials and their respective registries (CAP and CAP2) with 5,931 PY of follow-up (mean of 2.69 years) reported: similar rate of all-cause stroke between both arms (1.75 *vs*. 1.87 events/100 PY, *p =* 0.94): higher rate of ischemic stroke in Watchman group (1.6 *vs*. 0.9 events/100 PY, *p =* 0.05); and lower rates of hemorrhagic stroke, cardiovascular death (1.1 *vs*. 2.3 events/100 PY, *p =* 0.006), and non-procedural bleeding (6.0% *vs* 11.3%, *p =* 0.02) in Watchman group[54]. Although the rate of all-cause stroke was similar between both arms, the reduction in hemorrhagic stroke with Watchman device was balanced by a relative increase in ischemic stroke rates. This may relate to possible technical failures of the device: failure to completely obliterate LAA flow, anatomical remodeling of the LAA ostium over time resulting in more leaks, or the development of thrombus on the device[54]. Compared with the pooled results of PROTECT AF and PREVAIL trials mentioned above, the difference in ischemic stroke rate was not observed between LAAO and warfarin groups at longer and combined follow-up of 5 years[52].

EWOLUTION study (Registry on Watchman Outcomes in Real-Life Utilization) is a multicenter, prospective, non-randomized cohort that aimed to collect peri-procedural and long-term outcome data for patients implanted with Watchman device for LAAO[55]. This world-wide registry enrolled 1025 patients at 47 centers from the United States, Europe, Middle east and Russia who are more than 18-year-old and require LAAO based on ESC guidelines[55-57]. The device was successfully implanted in 98.5% and complete LAAO was achieved in 99.3% noted on TEE[56,57]. the rates of procedure-related serious adverse events (defined as; perforation, tamponade, embolism, neurological events, thrombosis, and bleeding) were 2.8% at 7 d and 3.6% at 30 d with bleeding being the most common adverse event[57]. This is lower than the 7-d procedure-related serious adverse events observed in PROTECT AF (8.7%) and PREVAIL (4.2%) trials. At 1 year follow up; mortality was 9.8%, device-related thrombus was seen in 3.7% of patients, and 1.1% of patients suffered from ischemic stroke, leading to 84% risk reduction of stroke. There was no hemorrhagic stroke observed during follow-up[56].

The ASAP study (ASA Plavix Feasibility Study with Watchman LAA Closure Technology Trial to assess) was a European multicenter, prospective, non-randomized study of Watchman device in patients with non-valvular AF who had CHADS2 score ≥1 and were not eligible for OACs[58]. After the device implantation, participants were given thienopyridine antiplatelet agent (clopidogrel or ticlopidine) for 6 mo and aspirin indefinitely. Out of 150 patients, 142 (94.7%) had successful implantation and 13 (8.7%) developed device-related adverse event. During mean follow up of 14.4 ± 8.6 mo, 4 patients developed stokes (2.3% per year) and 3 of them were ischemic (1.7% per year). There was 77% risk reduction in stroke compared to expected stroke risk based on CHADS2 score(7.3% per year). Till this moment, there is no published randomized data on the safety and efficacy of LAAO in patients with contraindications to anticoagulation. The ASAP TOO study(The Assessment of the Watchman Device in Patients Unsuitable for Oral Anticoagulation) is ongoing multi-center prospective randomized trial plan is to enroll up to 888 patients with non-valvular AF who are not candidate for OAC and have CHA2DS2-VASC ≥2[59]. The study will randomize patients to Watchman *vs* control. Control patients will be prescribed single antiplatelet therapy, or no therapy based on physician discretion.

***Amplatzer cardiac plug and amulet***

Amplatzer cardiac plug (ACP) (AGA, St. Jude Medical, Minneapolis, MN, United States) is another LAAO device that consists of a lobe and disc made of nitinol mesh and polyester patch, connected by central waist. Amulet® is a second-generation device of the Amplatzer with several incremental design improvements. It is larger in size and has higher number of stabilizing wires, which allows successful closure of more LAA anatomies (Table 1). Comparative studies have shown similar results with ACP and Amulet AMPLATZER devices in terms of safety, implantation success and appropriate LAAO[60,61]. Multiple retrospective and prospective studies for ACP and Amulet reported successful device implantation in 95% to 100% patients, with major periprocedural adverse events (death, stroke/TIA, device embolization, MI/perforation/tamponade/effusion, and major bleeding) ranging from 3.2% to 8%[62-67]. An FDA approval trial is currently ongoing, with the aim of collecting randomized controlled data from the Amulet and Watchman devices from 1,600 patients worldwide.PRAGUE 17 is another ongoing prospective, multicenter, randomized trial That plan to enroll 396 patients with non-valvular AF and assign them to LAAO using Amulet or Watchman *vs* non-vitamin K oral anticoagulants (NOACs). The aim at 24 mo of follow-up is to determine whether LAAO is non-inferior to NOACs in terms of primary efficacy endpoint and peri-procedural complications[68].

**Comparison of multiple LAA occlusion devices**

A meta-analysis on 2779 patients who had percutaneous LAAO with multiple devices [PLAATO (18%), Watchman (57%), and ACP (24%)] showed successful implantation in in 2611 patients (94%). The adjusted pooled incidence of stroke was 1.2 per 100 PY (95%CI: 0.9-1.6/100) and the combined efficacy outcome (stroke, systemic embolism, or cardiovascular death) rate was 2.7 per 100 PY (95%CI: 1.9-3.4/100). For combined adverse events, the random effect pooled rate was 6.5% (95%CI: 4.9%-8.2%)[69]. One single-center retrospective study in Italy compared the use ACP *vs* Watchman in 156 patients (ACP in 99 and watchman in 66 patients) and demonstrated procedural success in 99.4%. During follow-up, only 1 patient suffered from transient ischemic attack and 2 from cardiac death. Furthermore, the data showed excellent safety and efficacy with similar clinical outcomes in both devices[70]. Another multicenter retrospective registry for LAAO using various devices showed an overall success of 92.5%. The combined adverse event rate was 3.5%, leading to annual relative risk reduction for ischemic stroke, thromboembolic events, and major bleeding of 90.1%, 87.2%, and 92.9%, respectively[71]. RELEXAO (Registry on Real-Life Experience With LAA Occlusion) registry is a French retrospective cohort of patients with AF who were treated with LAAO[72]. In the study cohort from RELEXAO, Fauchier *et al*[72] reported no differences in death, ischemic stroke, major bleeding, or device related thrombus between Watchman and Amplatzer devices. Those studies underscore the high success rate in placing various LAAO devices, and their safety and efficacy in preventing strokes and adverse events.

**Antithrombotic therapy after device implantation**

Optimal anticoagulation/antiplatelet protocol post LAAO is highly patient-specific and recommended for a limited period post LAAO to prevent device associated thrombus[72]. Different anticoagulation strategies have been described in multiple studies including: warfarin, NOACs, DAPT, single antiplatelet (SAPT), or no therapy at all (Table 2). The anticoagulation protocol described In PROTECT AF and PREVAIL trials consists of warfarin for 45 d followed by aspirin and clopidogrel for 6 mo, then aspirin indefinitely[48-51]. In EWOLUTION registry for example, anticoagulation regimens post LAAO were variable and included: warfarin in 16%, NOAC in 11%, DAPT in 60%, single antiplatelet (SAPT) in 7%, and no therapy in 6%[56]. A study on post LAAO anticoagulation in patients from EWOLUTION registry demonstrated that NOAC and DAPT were similar to warfarin in terms of device thrombus, stroke or bleeding risks[73]. Compared with EWOLUTION registry, antithrombotic regimen post LAAO In RELEXAO registry was different and included: OACs 28.8%, SAPT in 36.2%, DAPT in 23.2%, OACs plus DAPT in 4.3%, and no therapy in 7.5%. In ASAP study, patients were given DAPT for 6 mo followed by aspirin indefinitely as they were ineligible for OACs[58]. A Questionnaire sent by European Heart Rhythm Association Electrophysiology to the participating centers to assess the indications and anticoagulation regimen post LAAO, showed that DAPT for 6 wk to 6 mo followed by aspirin monotherapy as the most common regimen[74]. Interestingly, 41% of centers would prescribe no therapy and less than 10% followed PROTECT AF and PREVAIL protocol. The European Heart Rhythm Association/European Association of Percutaneous Cardiovascular Interventions (EHRA/EAPCI) expert consensus statement recom­mends treatment with DAPT for 1 to 6 mo followed by aspirin indefinitely in patients with high bleeding risk[75].

**Complications**

Complications related to LAAO are either acute or delayed and most of them can be detect by peri-procedural imaging. Table 3 summarizes LAAO related complications, their incidence and treatment options.

**Conclusion**

LAAO is a reasonable alternative approach that is used for preventing embolic events in patients with AF who are deemed not eligible for anticoagulation. While AHA/ACC guidelines have no recommendations for LAAO, the ESC guidelines have class IIb (B) recommendation for LAAO in patients with AF and contra-indications for long-term anticoagulation. Similarly, Australian guidelines recommend considering LAAO in patients with non-valvular AF at moderate to high risk of stroke and with contraindications to OAC. Watchman is the only FDA approved device for LAAO and indicated to reduce the risk of thromboembolism from the LAA in patients with non-valvular AF who: are at increased risk for stroke and systemic embolism based on CHADS2 or CHA2DS2-VASc scores; are deemed by their physicians to be suitable for warfarin; and have an appropriate rationale to seek a non-pharmacologic alternative to warfarin, taking into account the safety and effectiveness of the device compared to warfarin. Amplatzer is another successful LAAO device that has CE mark and is waiting for FDA approval. Optimal antithrombotic regimen post LAAO is highly patient-specific and recommended to prevent device associated thrombus. Due to wide variety of shapes, sizes, indications, and implantation techniques in different LAAO devices, there is a need for further research to identify the best type of LAAO device that suites each patient profile. We believe that the development of established clinical guidelines and expert consensus supporting the use of LAAO in the foreseeable future will ultimately improve patient outcomes.

**References**

1 **Go AS**, Hylek EM, Phillips KA, Chang Y, Henault LE, Selby JV, Singer DE. Prevalence of diagnosed atrial fibrillation in adults: national implications for rhythm management and stroke prevention: the AnTicoagulation and Risk Factors in Atrial Fibrillation (ATRIA) Study. *JAMA* 2001; **285**: 2370-2375 [PMID: 11343485 doi: 10.1001/jama.285.18.2370]

2 **Miyasaka Y**, Barnes ME, Gersh BJ, Cha SS, Bailey KR, Abhayaratna WP, Seward JB, Tsang TS. Secular trends in incidence of atrial fibrillation in Olmsted County, Minnesota, 1980 to 2000, and implications on the projections for future prevalence. *Circulation* 2006; **114**: 119-125 [PMID: 16818816 DOI: 10.1161/CIRCULATIONAHA.105.595140]

3 **Chugh SS**, Havmoeller R, Narayanan K, Singh D, Rienstra M, Benjamin EJ, Gillum RF, Kim YH, McAnulty JH Jr, Zheng ZJ, Forouzanfar MH, Naghavi M, Mensah GA, Ezzati M, Murray CJ. Worldwide epidemiology of atrial fibrillation: a Global Burden of Disease 2010 Study. *Circulation* 2014; **129**: 837-847 [PMID: 24345399 DOI: 10.1161/CIRCULATIONAHA.113.005119]

4 **Colilla S**, Crow A, Petkun W, Singer DE, Simon T, Liu X. Estimates of current and future incidence and prevalence of atrial fibrillation in the U.S. adult population. *Am J Cardiol* 2013; **112**: 1142-1147 [PMID: 23831166 DOI: 10.1016/j.amjcard.2013.05.063]

5 **Benjamin EJ**, Blaha MJ, Chiuve SE, Cushman M, Das SR, Deo R, de Ferranti SD, Floyd J, Fornage M, Gillespie C, Isasi CR, Jiménez MC, Jordan LC, Judd SE, Lackland D, Lichtman JH, Lisabeth L, Liu S, Longenecker CT, Mackey RH, Matsushita K, Mozaffarian D, Mussolino ME, Nasir K, Neumar RW, Palaniappan L, Pandey DK, Thiagarajan RR, Reeves MJ, Ritchey M, Rodriguez CJ, Roth GA, Rosamond WD, Sasson C, Towfighi A, Tsao CW, Turner MB, Virani SS, Voeks JH, Willey JZ, Wilkins JT, Wu JH, Alger HM, Wong SS, Muntner P; American Heart Association Statistics Committee and Stroke Statistics Subcommittee. Heart Disease and Stroke Statistics-2017 Update: A Report From the American Heart Association. *Circulation* 2017; **135**: e146-e603 [PMID: 28122885 DOI: 10.1161/CIR.0000000000000485]

6 **Wolf PA**, Abbott RD, Kannel WB. Atrial fibrillation as an independent risk factor for stroke: the Framingham Study. *Stroke* 1991; **22**: 983-988 [PMID: 1866765 DOI: 10.1161/01.STR.22.8.983]

7 **Frost L**, Engholm G, Johnsen S, Møller H, Henneberg EW, Husted S. Incident thromboembolism in the aorta and the renal, mesenteric, pelvic, and extremity arteries after discharge from the hospital with a diagnosis of atrial fibrillation. *Arch Intern Med* 2001; **161**: 272-276 [PMID: 11176743 doi: 10.1001/archinte.161.2.272]

8 **Wang TJ**, Massaro JM, Levy D, Vasan RS, Wolf PA, D'Agostino RB, Larson MG, Kannel WB, Benjamin EJ. A risk score for predicting stroke or death in individuals with new-onset atrial fibrillation in the community: the Framingham Heart Study. *JAMA* 2003; **290**: 1049-1056 [PMID: 12941677 DOI: 10.1001/jama.290.8.1049]

9 **Hart RG**, Benavente O, McBride R, Pearce LA. Antithrombotic therapy to prevent stroke in patients with atrial fibrillation: a meta-analysis. *Ann Intern Med* 1999; **131**: 492-501 [PMID: 10507957 DOI: 10.7326/0003-4819-131-7-199910050-00003]

10 Risk factors for stroke and efficacy of antithrombotic therapy in atrial fibrillation. Analysis of pooled data from five randomized controlled trials. *Arch Intern Med* 1994; **154**: 1449-1457 [PMID: 8018000 doi: 10.1001/archinte.1994.00420130036007]

11 **Al-Saady NM**, Obel OA, Camm AJ. Left atrial appendage: structure, function, and role in thromboembolism. *Heart* 1999; **82**: 547-554 [PMID: 10525506]

12 **Wunderlich NC**, Beigel R, Swaans MJ, Ho SY, Siegel RJ. Percutaneous interventions for left atrial appendage exclusion: options, assessment, and imaging using 2D and 3D echocardiography. *JACC Cardiovasc Imaging* 2015; **8**: 472-488 [PMID: 25882576 DOI: 10.1016/j.jcmg.2015.02.002]

13 **Shirani J**, Alaeddini J. Structural remodeling of the left atrial appendage in patients with chronic non-valvular atrial fibrillation: Implications for thrombus formation, systemic embolism, and assessment by transesophageal echocardiography. *Cardiovasc Pathol* 2000; **9**: 95-101 [PMID: 10867359 doi: 10.1016/S1054-8807(00)00030-2]

14 **Manning WJ**, Silverman DI, Keighley CS, Oettgen P, Douglas PS. Transesophageal echocardiographically facilitated early cardioversion from atrial fibrillation using short-term anticoagulation: final results of a prospective 4.5-year study. *J Am Coll Cardiol* 1995; **25**: 1354-1361 [PMID: 7722133 DOI: 10.1016/0735-1097(94)00560-D]

15 **Leung DY**, Black IW, Cranney GB, Hopkins AP, Walsh WF. Prognostic implications of left atrial spontaneous echo contrast in nonvalvular atrial fibrillation. *J Am Coll Cardiol* 1994; **24**: 755-762 [PMID: 8077549 doi: 10.1016/0735-1097(94)90025-6]

16 **January CT**, Wann LS, Alpert JS, Calkins H, Cigarroa JE, Cleveland JC Jr, Conti JB, Ellinor PT, Ezekowitz MD, Field ME, Murray KT, Sacco RL, Stevenson WG, Tchou PJ, Tracy CM, Yancy CW; American College of Cardiology/American Heart Association Task Force on Practice Guidelines. 2014 AHA/ACC/HRS guideline for the management of patients with atrial fibrillation: a report of the American College of Cardiology/American Heart Association Task Force on Practice Guidelines and the Heart Rhythm Society. *J Am Coll Cardiol* 2014; **64**: e1-76 [PMID: 24685669 DOI: 10.1016/j.jacc.2014.03.022]

17 **Klein AL**, Grimm RA, Murray RD, Apperson-Hansen C, Asinger RW, Black IW, Davidoff R, Erbel R, Halperin JL, Orsinelli DA, Porter TR, Stoddard MF; Assessment of Cardioversion Using Transesophageal Echocardiography Investigators. Use of transesophageal echocardiography to guide cardioversion in patients with atrial fibrillation. *N Engl J Med* 2001; **344**: 1411-1420 [PMID: 11346805 DOI: 10.1056/NEJM200105103441901]

18 **Blackshear JL**, Odell JA. Appendage obliteration to reduce stroke in cardiac surgical patients with atrial fibrillation. *Ann Thorac Surg* 1996; **61**: 755-759 [PMID: 8572814 DOI: 10.1016/0003-4975(95)00887-X]

19 Predictors of thromboembolism in atrial fibrillation: II. Echocardiographic features of patients at risk. The Stroke Prevention in Atrial Fibrillation Investigators. *Ann Intern Med* 1992; **116**: 6-12 [PMID: 1727097 DOI: 10.1016/0736-4679(92)90169-T]

20 **Manning WJ**, Weintraub RM, Waksmonski CA, Haering JM, Rooney PS, Maslow AD, Johnson RG, Douglas PS. Accuracy of transesophageal echocardiography for identifying left atrial thrombi. A prospective, intraoperative study. *Ann Intern Med* 1995; **123**: 817-822 [PMID: 7486462 DOI: 10.7326/0003-4819-123-11-199512010-00001]

21 **Müller S**, Feuchtner G, Bonatti J, Müller L, Laufer G, Hiemetzberger R, Pachinger O, Barbieri V, Bartel T. Value of transesophageal 3D echocardiography as an adjunct to conventional 2D imaging in preoperative evaluation of cardiac masses. *Echocardiography* 2008; **25**: 624-631 [PMID: 18652008 DOI: 10.1111/j.1540-8175.2008.00664.x]

22 **Asch FM**, Bieganski SP, Panza JA, Weissman NJ. Real-time 3-dimensional echocardiography evaluation of intracardiac masses. *Echocardiography* 2006; **23**: 218-224 [PMID: 16524392 DOI: 10.1111/j.1540-8175.2006.00196.x]

23 **Perk G**, Lang RM, Garcia-Fernandez MA, Lodato J, Sugeng L, Lopez J, Knight BP, Messika-Zeitoun D, Shah S, Slater J, Brochet E, Varkey M, Hijazi Z, Marino N, Ruiz C, Kronzon I. Use of real time three-dimensional transesophageal echocardiography in intracardiac catheter based interventions. *J Am Soc Echocardiogr* 2009; **22**: 865-882 [PMID: 19647156 DOI: 10.1016/j.echo.2009.04.031]

24 **Berti S**, Paradossi U, Meucci F, Trianni G, Tzikas A, Rezzaghi M, Stolkova M, Palmieri C, Mori F, Santoro G. Periprocedural intracardiac echocardiography for left atrial appendage closure: a dual-center experience. *JACC Cardiovasc Interv* 2014; **7**: 1036-1044 [PMID: 25234677 DOI: 10.1016/j.jcin.2014.04.014]

25 **Martinez MW**, Kirsch J, Williamson EE, Syed IS, Feng D, Ommen S, Packer DL, Brady PA. Utility of nongated multidetector computed tomography for detection of left atrial thrombus in patients undergoing catheter ablation of atrial fibrillation. *JACC Cardiovasc Imaging* 2009; **2**: 69-76 [PMID: 19356536 DOI: 10.1016/j.jcmg.2008.09.011]

26 **Sakamoto I**, Hayashi K, Matsunaga N, Ogawa Y, Matsuoka Y, Okimoto T, Takagi M, Yano K, Toda G, Miyahara Y. Coronary angiographic finding of thrombus in the left atrial appendage. *Acta Radiol* 1996; **37**: 749-753 [PMID: 8915287 DOI: 10.1177/02841851960373P264]

27 **Kirchhof P**, Benussi S, Kotecha D, Ahlsson A, Atar D, Casadei B, Castella M, Diener HC, Heidbuchel H, Hendriks J, Hindricks G, Manolis AS, Oldgren J, Popescu BA, Schotten U, Van Putte B, Vardas P; ESC Scientific Document Group. 2016 ESC Guidelines for the management of atrial fibrillation developed in collaboration with EACTS. *Eur Heart J* 2016; **37**: 2893-2962 [PMID: 27567408 DOI: 10.1093/eurheartj/ehw210]

28 **NHFA CSANZ Atrial Fibrillation Guideline Working Group.**, Brieger D, Amerena J, Attia J, Bajorek B, Chan KH, Connell C, Freedman B, Ferguson C, Hall T, Haqqani H, Hendriks J, Hespe C, Hung J, Kalman JM, Sanders P, Worthington J, Yan TD, Zwar N. National Heart Foundation of Australia and the Cardiac Society of Australia and New Zealand: Australian Clinical Guidelines for the Diagnosis and Management of Atrial Fibrillation 2018. *Heart Lung Circ* 2018; **27**: 1209-1266 [PMID: 30077228 DOI: 10.1016/j.hlc.2018.06.1043]

29 **Holmes DR Jr**, Lakkireddy DR, Whitlock RP, Waksman R, Mack MJ. Left atrial appendage occlusion: opportunities and challenges. *J Am Coll Cardiol* 2014; **63**: 291-298 [PMID: 24076495 DOI: 10.1016/j.jacc.2013.08.1631]

30 **Kanderian AS**, Gillinov AM, Pettersson GB, Blackstone E, Klein AL. Success of surgical left atrial appendage closure: assessment by transesophageal echocardiography. *J Am Coll Cardiol* 2008; **52**: 924-929 [PMID: 18772063 DOI: 10.1016/j.jacc.2008.03.067]

31 **Friedman DJ**, Piccini JP, Wang T, Zheng J, Malaisrie SC, Holmes DR, Suri RM, Mack MJ, Badhwar V, Jacobs JP, Gaca JG, Chow SC, Peterson ED, Brennan JM. Association Between Left Atrial Appendage Occlusion and Readmission for Thromboembolism Among Patients With Atrial Fibrillation Undergoing Concomitant Cardiac Surgery. *JAMA* 2018; **319**: 365-374 [PMID: 29362794 DOI: 10.1001/jama.2017.20125]

32 **Salzberg SP**, Plass A, Emmert MY, Desbiolles L, Alkadhi H, Grünenfelder J, Genoni M. Left atrial appendage clip occlusion: early clinical results. *J Thorac Cardiovasc Surg* 2010; **139**: 1269-1274 [PMID: 19880144 DOI: 10.1016/j.jtcvs.2009.06.033]

33 **Ailawadi G**, Gerdisch MW, Harvey RL, Hooker RL, Damiano RJ Jr, Salamon T, Mack MJ. Exclusion of the left atrial appendage with a novel device: early results of a multicenter trial. *J Thorac Cardiovasc Surg* 2011; **142**: 1002-1009, 1009.e1 [PMID: 21906756 DOI: 10.1016/j.jtcvs.2011.07.052]

34 **Slater AD**, Tatooles AJ, Coffey A, Pappas PS, Bresticker M, Greason K, Slaughter MS. Prospective clinical study of a novel left atrial appendage occlusion device. *Ann Thorac Surg* 2012; **93**: 2035-8; discussion 2038-40 [PMID: 22632497 DOI: 10.1016/j.athoracsur.2011.12.077]

35 **Ventosa-Fernandez G**, Quintana E, Castellá M, Pereda D. Exclusion of the left atrial appendage with the TigerPaw II system: a word of caution. *Interact Cardiovasc Thorac Surg* 2015; **21**: 803-804 [PMID: 26395944 DOI: 10.1093/icvts/ivv256]

36 **Bartus K**, Han FT, Bednarek J, Myc J, Kapelak B, Sadowski J, Lelakowski J, Bartus S, Yakubov SJ, Lee RJ. Percutaneous left atrial appendage suture ligation using the LARIAT device in patients with atrial fibrillation: initial clinical experience. *J Am Coll Cardiol* 2013; **62**: 108-118 [PMID: 23062528 DOI: 10.1016/j.jacc.2012.06.046]

37 **Massumi A**, Chelu MG, Nazeri A, May SA, Afshar-Kharaghan H, Saeed M, Razavi M, Rasekh A. Initial experience with a novel percutaneous left atrial appendage exclusion device in patients with atrial fibrillation, increased stroke risk, and contraindications to anticoagulation. *Am J Cardiol* 2013; **111**: 869-873 [PMID: 23312129 DOI: 10.1016/j.amjcard.2012.11.061]

38 **Dar T**, Afzal MR, Yarlagadda B, Kutty S, Shang Q, Gunda S, Samanta A, Thummaluru J, Arukala KS, Kanmanthareddy A, Reddy M, Atkins D, Bommana S, Dawn B, Lakkireddy D. Mechanical function of the left atrium is improved with epicardial ligation of the left atrial appendage: Insights from the LAFIT-LARIAT Registry. *Heart Rhythm* 2018; **15**: 955-959 [PMID: 29477973 DOI: 10.1016/j.hrthm.2018.02.022]

39 **Giedrimas E**, Lin AC, Knight BP. Left atrial thrombus after appendage closure using LARIAT. *Circ Arrhythm Electrophysiol* 2013; **6**: e52-e53 [PMID: 23962862 DOI: 10.1161/CIRCEP.113.000532]

40 **Briceno DF**, Fernando RR, Laing ST. Left atrial appendage thrombus post LARIAT closure device. *Heart Rhythm* 2014; **11**: 1600-1601 [PMID: 24184785 DOI: 10.1016/j.hrthm.2013.10.053]

41 **Truesdell AG**, Patel CP, Maini BS. Late-occurring left atrial appendage thrombus after ligation using LARIAT. *J Interv Card Electrophysiol* 2014; **41**: 101 [PMID: 24928486 DOI: 10.1007/s10840-014-9916-9]

42 **Pillai AM**, Kanmanthareddy A, Earnest M, Reddy M, Ferrell R, Nath J, Pillarisetti J, Vallakati A, Lakkireddy D. Initial experience with post Lariat left atrial appendage leak closure with Amplatzer septal occluder device and repeat Lariat application. *Heart Rhythm* 2014; **11**: 1877-1883 [PMID: 24993460 DOI: 10.1016/j.hrthm.2014.06.035]

43 **Sievert H**, Lesh MD, Trepels T, Omran H, Bartorelli A, Della Bella P, Nakai T, Reisman M, DiMario C, Block P, Kramer P, Fleschenberg D, Krumsdorf U, Scherer D. Percutaneous left atrial appendage transcatheter occlusion to prevent stroke in high-risk patients with atrial fibrillation: early clinical experience. *Circulation* 2002; **105**: 1887-1889 [PMID: 11997272 DOI: 10.1161/01.CIR.0000015698.54752.6D]

44 **Ostermayer SH**, Reisman M, Kramer PH, Matthews RV, Gray WA, Block PC, Omran H, Bartorelli AL, Della Bella P, Di Mario C, Pappone C, Casale PN, Moses JW, Poppas A, Williams DO, Meier B, Skanes A, Teirstein PS, Lesh MD, Nakai T, Bayard Y, Billinger K, Trepels T, Krumsdorf U, Sievert H. Percutaneous left atrial appendage transcatheter occlusion (PLAATO system) to prevent stroke in high-risk patients with non-rheumatic atrial fibrillation: results from the international multi-center feasibility trials. *J Am Coll Cardiol* 2005; **46**: 9-14 [PMID: 15992628 DOI: 10.1016/j.jacc.2005.03.042]

45 **Bayard YL**, Omran H, Neuzil P, Thuesen L, Pichler M, Rowland E, Ramondo A, Ruzyllo W, Budts W, Montalescot G, Brugada P, Serruys PW, Vahanian A, Piéchaud JF, Bartorelli A, Marco J, Probst P, Kuck KH, Ostermayer SH, Büscheck F, Fischer E, Leetz M, Sievert H. PLAATO (Percutaneous Left Atrial Appendage Transcatheter Occlusion) for prevention of cardioembolic stroke in non-anticoagulation eligible atrial fibrillation patients: results from the European PLAATO study. *EuroIntervention* 2010; **6**: 220-226 [PMID: 20562072]

46 **Park JW**, Leithäuser B, Gerk U, Vrsansky M, Jung F. Percutaneous left atrial appendage transcatheter occlusion (PLAATO) for stroke prevention in atrial fibrillation: 2-year outcomes. *J Invasive Cardiol* 2009; **21**: 446-450 [PMID: 19726815]

47 **Sick PB**, Schuler G, Hauptmann KE, Grube E, Yakubov S, Turi ZG, Mishkel G, Almany S, Holmes DR. Initial worldwide experience with the WATCHMAN left atrial appendage system for stroke prevention in atrial fibrillation. *J Am Coll Cardiol* 2007; **49**: 1490-1495 [PMID: 17397680 DOI: 10.1016/j.jacc.2007.02.035]

48 **Holmes DR**, Reddy VY, Turi ZG, Doshi SK, Sievert H, Buchbinder M, Mullin CM, Sick P; PROTECT AF Investigators. Percutaneous closure of the left atrial appendage versus warfarin therapy for prevention of stroke in patients with atrial fibrillation: a randomised non-inferiority trial. *Lancet* 2009; **374**: 534-542 [PMID: 19683639 DOI: 10.1016/S0140-6736(09)61343-X]

49 **Reddy VY**, Doshi SK, Sievert H, Buchbinder M, Neuzil P, Huber K, Halperin JL, Holmes D; PROTECT AF Investigators. Percutaneous left atrial appendage closure for stroke prophylaxis in patients with atrial fibrillation: 2.3-Year Follow-up of the PROTECT AF (Watchman Left Atrial Appendage System for Embolic Protection in Patients with Atrial Fibrillation) Trial. *Circulation* 2013; **127**: 720-729 [PMID: 23325525 DOI: 10.1161/CIRCULATIONAHA.112.114389]

50 **Reddy VY**, Sievert H, Halperin J, Doshi SK, Buchbinder M, Neuzil P, Huber K, Whisenant B, Kar S, Swarup V, Gordon N, Holmes D; PROTECT AF Steering Committee and Investigators. Percutaneous left atrial appendage closure vs warfarin for atrial fibrillation: a randomized clinical trial. *JAMA* 2014; **312**: 1988-1998 [PMID: 25399274 DOI: 10.1001/jama.2014.15192]

51 **Holmes DR Jr**, Kar S, Price MJ, Whisenant B, Sievert H, Doshi SK, Huber K, Reddy VY. Prospective randomized evaluation of the Watchman Left Atrial Appendage Closure device in patients with atrial fibrillation versus long-term warfarin therapy: the PREVAIL trial. *J Am Coll Cardiol* 2014; **64**: 1-12 [PMID: 24998121 DOI: 10.1016/j.jacc.2014.04.029]

52 **Reddy VY**, Doshi SK, Kar S, Gibson DN, Price MJ, Huber K, Horton RP, Buchbinder M, Neuzil P, Gordon NT, Holmes DR Jr; PREVAIL and PROTECT AF Investigators. 5-Year Outcomes After Left Atrial Appendage Closure: From the PREVAIL and PROTECT AF Trials. *J Am Coll Cardiol* 2017; **70**: 2964-2975 [PMID: 29103847 DOI: 10.1016/j.jacc.2017.10.021]

53 **Reddy VY**, Holmes D, Doshi SK, Neuzil P, Kar S. Safety of percutaneous left atrial appendage closure: results from the Watchman Left Atrial Appendage System for Embolic Protection in Patients with AF (PROTECT AF) clinical trial and the Continued Access Registry. *Circulation* 2011; **123**: 417-424 [PMID: 21242484 DOI: 10.1161/CIRCULATIONAHA.110.976449]

54 **Holmes DR Jr**, Doshi SK, Kar S, Price MJ, Sanchez JM, Sievert H, Valderrabano M, Reddy VY. Left Atrial Appendage Closure as an Alternative to Warfarin for Stroke Prevention in Atrial Fibrillation: A Patient-Level Meta-Analysis. *J Am Coll Cardiol* 2015; **65**: 2614-2623 [PMID: 26088300 DOI: 10.1016/j.jacc.2015.04.025]

55 **Boersma LV**, Schmidt B, Betts TR, Sievert H, Tamburino C, Teiger E, Stein KM, Bergmann MW. EWOLUTION: Design of a registry to evaluate real-world clinical outcomes in patients with AF and high stroke risk-treated with the WATCHMAN left atrial appendage closure technology. *Catheter Cardiovasc Interv* 2016; **88**: 460-465 [PMID: 26719158 DOI: 10.1002/ccd.26358]

56 **Boersma LV**, Ince H, Kische S, Pokushalov E, Schmitz T, Schmidt B, Gori T, Meincke F, Protopopov AV, Betts T, Foley D, Sievert H, Mazzone P, De Potter T, Vireca E, Stein K, Bergmann MW; EWOLUTION Investigators. Efficacy and safety of left atrial appendage closure with WATCHMAN in patients with or without contraindication to oral anticoagulation: 1-Year follow-up outcome data of the EWOLUTION trial. *Heart Rhythm* 2017; **14**: 1302-1308 [PMID: 28577840 DOI: 10.1016/j.hrthm.2017.05.038]

57 **Boersma LV**, Schmidt B, Betts TR, Sievert H, Tamburino C, Teiger E, Pokushalov E, Kische S, Schmitz T, Stein KM, Bergmann MW; EWOLUTION investigators. Implant success and safety of left atrial appendage closure with the WATCHMAN device: peri-procedural outcomes from the EWOLUTION registry. *Eur Heart J* 2016; **37**: 2465-2474 [PMID: 26822918 DOI: 10.1093/eurheartj/ehv730]

58 **Reddy VY**, Möbius-Winkler S, Miller MA, Neuzil P, Schuler G, Wiebe J, Sick P, Sievert H. Left atrial appendage closure with the Watchman device in patients with a contraindication for oral anticoagulation: the ASAP study (ASA Plavix Feasibility Study With Watchman Left Atrial Appendage Closure Technology). *J Am Coll Cardiol* 2013; **61**: 2551-2556 [PMID: 23583249 DOI: 10.1016/j.jacc.2013.03.035]

59 **Holmes DR**, Reddy VY, Buchbinder M, Stein K, Elletson M, Bergmann MW, Schmidt B, Saw J. The Assessment of the Watchman Device in Patients Unsuitable for Oral Anticoagulation (ASAP-TOO) trial. *Am Heart J* 2017; **189**: 68-74 [PMID: 28625383 DOI: 10.1016/j.ahj.2017.03.007]

60 **Gloekler S**, Shakir S, Doblies J, Khattab AA, Praz F, Guerios Ê, Koermendy D, Stortecky S, Pilgrim T, Buellesfeld L, Wenaweser P, Windecker S, Moschovitis A, Jaguszewski M, Landmesser U, Nietlispach F, Meier B. Early results of first versus second generation Amplatzer occluders for left atrial appendage closure in patients with atrial fibrillation. *Clin Res Cardiol* 2015; **104**: 656-665 [PMID: 25736061 DOI: 10.1007/s00392-015-0828-1]

61 **Abualsaud A**, Freixa X, Tzikas A, Chan J, Garceau P, Basmadjian A, Ibrahim R. Side-by-Side Comparison of LAA Occlusion Performance With the Amplatzer Cardiac Plug and Amplatzer Amulet. *J Invasive Cardiol* 2016; **28**: 34-38 [PMID: 26716593]

62 **Landmesser U**, Schmidt B, Nielsen-Kudsk JE, Lam SCC, Park JW, Tarantini G, Cruz-Gonzalez I, Geist V, Della Bella P, Colombo A, Zeus T, Omran H, Piorkowski C, Lund J, Tondo C, Hildick-Smith D. Left atrial appendage occlusion with the AMPLATZER Amulet device: periprocedural and early clinical/echocardiographic data from a global prospective observational study. *EuroIntervention* 2017; **13**: 867-876 [PMID: 28649053 DOI: 10.4244/EIJ-D-17-00493]

63 **Tzikas A**, Gafoor S, Meerkin D, Freixa X, Cruz-Gonzalez I, Lewalter T, Saw J, Berti S, Nielsen-Kudsk JE, Ibrahim R, Lakkireddy D, Paul V, Arzamendi D, Nietlispach F, Worthley SG, Hildick-Smith D, Thambo JB, Tondo C, Aminian A, Kalarus Z, Schmidt B, Sondergaard L, Kefer J, Meier B, Park JW, Sievert H, Omran H. Left atrial appendage occlusion with the AMPLATZER Amulet device: an expert consensus step-by-step approach. *EuroIntervention* 2016; **11**: 1512-1521 [PMID: 27107315 DOI: 10.4244/EIJV11I13A292]

64 **Korsholm K**, Nielsen KM, Jensen JM, Jensen HK, Andersen G, Nielsen-Kudsk JE. Transcatheter left atrial appendage occlusion in patients with atrial fibrillation and a high bleeding risk using aspirin alone for post-implant antithrombotic therapy. *EuroIntervention* 2017; **12**: 2075-2082 [PMID: 27973336 DOI: 10.4244/EIJ-D-16-00726]

65 **Berti S**, Pastormerlo LE, Rezzaghi M, Trianni G, Paradossi U, Cerone E, Ravani M, De Caterina AR, Rizza A, Palmieri C. Left atrial appendage occlusion in high-risk patients with non-valvular atrial fibrillation. *Heart* 2016; **102**: 1969-1973 [PMID: 27492943 DOI: 10.1136/heartjnl-2015-309150]

66 **Kleinecke C**, Park JW, Gödde M, Zintl K, Schnupp S, Brachmann J. Twelve-month follow-up of left atrial appendage occlusion with Amplatzer Amulet. *Cardiol J* 2017; **24**: 131-138 [PMID: 28198520 DOI: 10.5603/CJ.a2017.0017]

67 **Urena M**, Rodés-Cabau J, Freixa X, Saw J, Webb JG, Freeman M, Horlick E, Osten M, Chan A, Marquis JF, Champagne J, Ibrahim R. Percutaneous left atrial appendage closure with the AMPLATZER cardiac plug device in patients with nonvalvular atrial fibrillation and contraindications to anticoagulation therapy. *J Am Coll Cardiol* 2013; **62**: 96-102 [PMID: 23665098 DOI: 10.1016/j.jacc.2013.02.089]

68 **Osmancik P**, Tousek P, Herman D, Neuzil P, Hala P, Stasek J, Haman L, Kala P, Poloczek M, Branny M, Chovancik J, Cervinka P, Holy J, Vancura V, Rokyta R, Taborsky M, Kovarnik T, Zemanek D, Peichl P, Haskova S, Jarkovsky J, Widimsky P; PRAGUE-17 Investigators. Interventional left atrial appendage closure vs novel anticoagulation agents in patients with atrial fibrillation indicated for long-term anticoagulation (PRAGUE-17 study). *Am Heart J* 2017; **183**: 108-114 [PMID: 27979034 DOI: 10.1016/j.ahj.2016.10.003]

69 **Xu H**, Xie X, Wang B, Ma S, Wang F. Efficacy and Safety of Percutaneous Left Atrial Appendage Occlusion for Stroke Prevention in Nonvalvular Atrial Fibrillation: A Meta-analysis of Contemporary Studies. *Heart Lung Circ* 2016; **25**: 1107-1117 [PMID: 27199213 DOI: 10.1016/j.hlc.2016.03.016]

70 **Figini F**, Mazzone P, Regazzoli D, Porata G, Ruparelia N, Giannini F, Stella S, Ancona F, Agricola E, Sora N, Marzi A, Aurelio A, Trevisi N, Della Bella P, Colombo A, Montorfano M. Left atrial appendage closure: A single center experience and comparison of two contemporary devices. *Catheter Cardiovasc Interv* 2017; **89**: 763-772 [PMID: 27567013 DOI: 10.1002/ccd.26678]

71 **Betts TR**, Leo M, Panikker S, Kanagaratnam P, Koa-Wing M, Davies DW, Hildick-Smith D, Wynne DG, Ormerod O, Segal OR, Chow AW, Todd D, Cabrera Gomez S, Kirkwood GJ, Fox D, Pepper C, Foran J, Wong T. Percutaneous left atrial appendage occlusion using different technologies in the United Kingdom: A multicenter registry. *Catheter Cardiovasc Interv* 2017; **89**: 484-492 [PMID: 27651124 DOI: 10.1002/ccd.26782]

72 **Fauchier L**, Cinaud A, Brigadeau F, Lepillier A, Pierre B, Abbey S, Fatemi M, Franceschi F, Guedeney P, Jacon P, Paziaud O, Venier S, Deharo JC, Gras D, Klug D, Mansourati J, Montalescot G, Piot O, Defaye P. Device-Related Thrombosis After Percutaneous Left Atrial Appendage Occlusion for Atrial Fibrillation. *J Am Coll Cardiol* 2018; **71**: 1528-1536 [PMID: 29622159 DOI: 10.1016/j.jacc.2018.01.076]

73 **Bergmann MW**, Betts TR, Sievert H, Schmidt B, Pokushalov E, Kische S, Schmitz T, Meincke F, Stein KM, Boersma LVA, Ince H. Safety and efficacy of early anticoagulation drug regimens after WATCHMAN left atrial appendage closure: three-month data from the EWOLUTION prospective, multicentre, monitored international WATCHMAN LAA closure registry. *EuroIntervention* 2017; **13**: 877-884 [PMID: 28606886 DOI: 10.4244/EIJ-D-17-00042]

74 **Tilz RR**, Potpara T, Chen J, Dobreanu D, Larsen TB, Haugaa KH, Dagres N. Left atrial appendage occluder implantation in Europe: indications and anticoagulation post-implantation. Results of the European Heart Rhythm Association Survey. *Europace* 2017; **19**: 1737-1742 [PMID: 29016910 DOI: 10.1093/europace/eux254]

75 **Meier B**, Blaauw Y, Khattab AA, Lewalter T, Sievert H, Tondo C, Glikson M. EHRA/EAPCI expert consensus statement on catheter-based left atrial appendage occlusion. *EuroIntervention* 2015; **10**: 1109-1125 [PMID: 25169595]

76 **Reddy VY**, Gibson DN, Kar S, O'Neill W, Doshi SK, Horton RP, Buchbinder M, Gordon NT, Holmes DR. Post-Approval U.S. Experience With Left Atrial Appendage Closure for Stroke Prevention in Atrial Fibrillation. *J Am Coll Cardiol* 2017; **69**: 253-261 [PMID: 27816552 DOI: 10.1016/j.jacc.2016.10.010]

77 **Singh SM**, Douglas PS, Reddy VY. The incidence and long-term clinical outcome of iatrogenic atrial septal defects secondary to transseptal catheterization with a 12F transseptal sheath. *Circ Arrhythm Electrophysiol* 2011; **4**: 166-171 [PMID: 21248245 DOI: 10.1161/CIRCEP.110.959015]

**P-Reviewer:** Barik R, Dai X, Kharlamov AN, Korosoglou G, Ueda H **S-Editor:** Ma YJ **L-Editor:** A **E-Editor:** Wu YXJ

**Specialty type:** Cardiac and Cardiovascular Systems

**Country of origin:** United States

**Peer-review report classification**


Grade A (Excellent): A

Grade B (Very good): B, B, B

Grade C (Good): 0

Grade D (Fair): D

Grade E (Poor): 0

****

**Figure 1 Surgical and percutaneous devices that are used for left atrial appendage occlusion.**

**Table 1 Comparison of left atrial appendage occlusion devices**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Device | Study | Year of Introduction | Description | Approach | Approval |
| Atriclip Device System (Atricure) | EXCLUDE study[33] | 2008 | Self-closing, sterile, implantable clip, with a reusable deployment tool applied pericardially. | Epicardial | CE Mark |
| Tiger Paw System (Terumo Cardiovascular Systems) | Slater et al[34] | Introduced in 2009 and recalled in 2015 | Implantable fastener of titanium connectors that staples the LAA tissue plus rims of silicone that seal the puncture sites. | Epicardial | Recalled |
| Lariat device (SentreHEART) | Bartus *et al*[36]  Massumi *et al*[37]  Dar *et al*[38] | 2009 | Multicomponent system including: transvenous and epicardial balloon catheters, magnet tipped guidewires, and suture delivery system | Epicardial and transvenous | FDA approval for soft tissue closure not LAAO  CE mark |
| PLAATO  (Appriva Medical) | Sievert *et al*[43]  Ostermaye *et al*[44]  Bayard *et al*[45]  Park *et al*[46] | Introduced in 2001and discontinued in 2007 | Self-expanding nitinol cage covered with polymeric membrane (ePTFE) designed to be placed in the orifice of the LAA | Transvenous, trans-septal | Discontinued |
| Watchman  (Boston Scientific) | Pilot study[47]  PROTECT AF study[48-50]  PREVAIL study[51,52]  CAP 1 registry[53]  CAP 2 registry[54]  EWOLUTION registry[55-57]    RELEXAO  Registry[72]  ASAP study[58]  ASAP TOO study[59] | 2005 | Self-expanding nitinol frame structure with fixation barbs and a permeable polyester fabric that covers the atrial facing surface of the device | Transvenous, trans-septal | FDA approved and CE Mark |
| ACP (St. Jude Medical) | Urena *et al*[67]  Gloekler *et al*[60]  Abualsaud *et al*[61]  Korsholm *et al*[64]  Berti *et al*[65]  RELEXAO  Registry[72] | 2008 | Self-expanding distal lobe (6.5mm in length) and proximal disc (4-6mm larger than distal lobe) nitinol mesh with articulating waist | Transvenous, trans-septal | CE Mark |
| Amplatzer Amulet  (St. Jude Medical) | Gloekler *et al*[60]  Abualsaud *et al*[61]  Landmesser *et al*[62]  Tzikas *et al*[63]  Korsholm *et al*[64]  Berti *et al*[65]  Kleinecke *et al*[66]  RELEXAO  Registry[72] | 2013 | Self-expanding distal lobe and proximal disc nitinol mesh with articulating waist, and more anchors | Transvenous, trans-septal | CE Mark |

LAAO: left atrial appendage occlusion; FDA: Food and drug administration; ACP: Amplatzer cardiac plug.

**Table 2 Antithrombotic therapy regimens following left atrial appendage occlusion**

|  |  |
| --- | --- |
| Study/reference | Regimen |
| PROTECT AF trial[48-50] | Warfarin for 45 d followed by aspirin and clopidogrel for 6 mo, then aspirin indefinitely |
| PREVAIL trial[51] | Warfarin for 45 d followed by aspirin and clopidogrel for 6 mo, then aspirin indefinitely |
| EWOLUTION registry[56] | warfarin in 16%, NOAC in 11%, DAPT in 60%, SAPT in 7%, and no therapy in 6% |
| RELEXAO registry[72] | OACs in 28.8%, SAPT in 36.2%, DAPT in 23.2%, OACs plus DAPT in 4.3%, and no therapy in 7.5%. |
| ASAP trial[58] | DAPT for 6 mo followed by aspirin indefinitely |
| EHRA/EAPCI expert consensus[75] | DAPT for 1 to 6 mo followed by aspirin indefinitely |

OAC: Oral anticoagulant; NOAC: Non-vitamin K oral anticoagulant; SAPT: Single antiplatelet; DAPT: Dual antiplatelet; EHRA/EAPCI: European Heart Rhythm Association/European Association of Percutaneous Cardiovascular Interventions.

**Table 3 complications related to left atrial appendage occlusion**

|  |  |  |
| --- | --- | --- |
| Complication | Incidence | Treatment |
| Pericardial effusion/tamponade that require intervention[47,48,50,51,53,56-58,63,76] | 1.2% to 5% | Pericardiocentesis |
| Device embolization[47,48,50,51,56-58,63,76] | 0% to 3.7% | Transcatheter removal or surgery |
| Device related thrombus[47,56-58,63] | Up to 14% | Anticoagulation |
| Persistent ASD[77] | 11% at 6 mo and 7% at 12 mo | Usually small no need for treatment |
| Cardiac perforation[51] | 0% to 0.4% | surgery |
| Procedure related stroke[47,48,51, 53,56-58,76] | 0% to 1.1% | Stroke management |

ASD: atrial septal defect.