Name of Journal: World Journal of Orthopedics

Manuscript NO: 56927

Manuscript Type: ORIGINAL ARTICLE

Randomized Controlled Trial

Highly cross-linked versus conventional polyethylene inserts in total hip arthroplasty, a five-year Roentgen stereophotogrammetric analysis randomised controlled trial

van Loon J et al. HXLPE in THA: A RSA RCT

Justin van Loon, Daniël Hoornenborg, Inger Sierevelt, Kim T.M. Opdam, Gino M.M.J. Kerkhoffs, Daniël Haverkamp

Justin van Loon, Daniël Hoornenborg, Inger Sierevelt, Daniël Haverkamp, Department of Orthopaedic Surgery, Xpert Orthopedie Amsterdam/Specialized Center of Orthopedic Research and Education, Amsterdam 1101EA, Netherlands

Justin van Loon, Kim T.M. Opdam, Gino M.M.J. Kerkhoffs, Department of Orthopaedic Surgery, Amsterdam University Medical Centres, Location Academic Medical Center, Amsterdam 1105AZ, Netherlands

Author contributions: All authors contributed to the study conception and design. Haverkamp D and Hoornenborg D were involved in the initial surgery procedures; writing of the study protocol was performed by Haverkamp D; data collection and analysis were performed by van Loon J and Sierevelt IN; the first draft of the manuscript was written by van Loon J; all authors commented on previous versions of the manuscript. All authors read and approved the final manuscript.

Corresponding author: Daniël Haverkamp, MD, PhD, Surgeon, Department of Orthopaedic Surgery, Xpert Orthopedie Amsterdam/Specialized Center of Orthopedic Research and Education, Laarderhoogtweg 12, Amsterdam 1101EA, Netherlands. d.haverkamp@xpertorthopedie.nl

Received: May 19, 2020

Revised: June 3, 2020

Accepted: August 25, 2020

Published online:

Abstract

BACKGROUND

Polyethylene (PE) particles produced by wear of the acetabular insert are thought to cause osteolysis and thereby aseptic loosening of the implant in total hip arthroplasty (THA). As highly cross-linked polyethylene (HXLPE) is presumed to give lower wear rates, in vivo studies are needed to confirm this.

AIM

To compare the wear of REXPOL, a HXPLE, with conventional PE within the first five years after implantation using Roentgen stereophotogrammetric analysis (RSA).

METHODS

Patients were randomised to receive either a HXLPE (REXPOL) or a conventional PE insert during primary THA. RSA images were obtained directly postoperative and after 6 wk, 12 wk, 6 mo, 12 mo, 24 mo and five years. Functional outcomes were assessed using the Hip Injury and Osteoarthritis Outcome Score and Harris Hip Score at baseline and five years after surgery.

RESULTS

The HXLPE (REXPOL) showed less wear in the latero-medial direction. Significant wear rates of conventional PE were seen in the latero-medial and center-proximal direction and in volume and corrected volume, whereas the REXPOL did not show these outcomes over time. Improvement from baseline in functional outcome did not significantly differ.

CONCLUSION

Total 3D wear is less in THAs inserted with a REXPOL inlay than a conventional PE inlay after five years. This study confirms, for the first, that the REXPOL HXLPE inlay is preferred to standard PE.

Key words: Highly cross-linked polyethylene; Primary total hip arthroplasty; Polyethylene wear; Roentgen stereophotogrammetric analysis

van Loon J, Hoornenborg D, Sierevelt I, Opdam KT, Kerkhoffs GM, Haverkamp D. Highly cross-linked versus conventional polyethylene inserts in total hip arthroplasty, a five-year Roentgen stereophotogrammetric analysis randomised controlled trial. World J Orthop 2020; In press

Core tip: Polyethylene (PE) particles produced by wear of the acetabular insert can cause osteolysis and thereby aseptic loosening of the implant in total hip arthroplasty (THA). As highly cross-linked polyethylene (HXLPE) is presumed to result in lower wear rates, we performed a randomised controlled trial (RCT) using Roentgen stereophotogrammetric analysis with a five-year follow-up period. This RCT showed, for the first time, that five-year total 3D wear was less in THAs inserted with REXPOL HXLPE and hereby confirms that this inlay is preferred to standard PE.

INTRODUCTION

Since the introduction of total hip arthroplasty (THA) in the 1960s, the incidence of this procedure has been increasing. Although THA is one of the most successful orthopaedic procedures, the main causes of late revisions are wear, and the resulting osteolysis causing aseptic loosening of the implant[1].

Therefore, the search to minimise wear continues and several bearing couplings over time have been tried, of which polyethylene (PE) with a ceramic head still remains the best option[2]. However, wear still occurs due to existing friction, resulting in progressive loss of material and the presence of microparticles. These PE particles induce a foreign-body reaction, which results in osteolysis[3]. The number of wear particles produced, the material used, and its morphological form determine the severity of the aforementioned reaction[4]. In response to this problem of PE wear, a highly cross-linked PE (HXLPE) has been developed. Following irradiation, free-radicals are formed, creating cross-links in the PE, which are increased by heating and reduce wear[5]. Depending on the type and dose of irradiation and the type of PE used as the control group, wear can be decreased by 42%-100% compared to traditional PE[6].

To determine the performance of an implant, a standardised and reliable method is required to measure wear. Stilling et al[7] demonstrated that wear in different directions combined with volume wear of acetabular inserts can be calculated accurately using Roentgen stereophotogrammetric analysis (RSA). As wear is one of the most important reasons for revision in THA, and therefore an indicator of long-term survival, HXLPE could reduce the number of revisions needed in the future. To prove this, in vivo analyses with RSA are needed to confirm that the in vitro results are confirmed in the real setting.

The objective of this randomised controlled trial (RCT) was to compare the wear of two different inlays, the HXLPE (REXPOL) and conventional PE acetabular inserts with similar ceramic head articulation, within the first five years after implantation.

Our hypothesis was that total 3D wear after five years in the REXPOL group would be less than that in the conventional PE group.

MATERIALS AND METHODS

Ethical approval/registration

This single center RCT was granted ethical approval by the local ethics committee review board of the Slotervaart Medical Center (registration number: NL23524.048.08; Dutch trial register: NL5605). The design and reporting of this study were conducted in accordance with the Consolidated Standards of Reporting Trials (CONSORT) principles.

Study design

This was a single center, double-blind RCT comparing HXLPE (REXPOL, Smith and Nephew) to standard PE acetabular inserts (Standard PE, Smith and Nephew) with similar ceramic head (Biolox, Smith and Nephew) articulation. Both the patients and investigators were blinded with regard to the group patients were assigned to. RSA analysis was performed in a blinded mode. Randomisation was performed by the use of numbered opaque envelopes, containing the prescribed PE insert. The orthopaedic surgeon randomly received those envelopes and opened them prior to the procedure.

Eligibility

Between January 2011 and January 2014, patients undergoing THA in the Slotervaart Medical Center were included in this study after completing an informed consent if they met the inclusion criteria (Table 1).

Sample size

Previous RSA studies showed a high degree of sensitivity and accuracy of measurements of migration; relatively small patient groups showed a statistically significant outcome[8]. Standard PE has a linear wear rate of around 0.06-0.08 mm/year, whereas REXPOL is expected to show almost no wear over five years. A recent publication on five-year wear results in THA measured by RSA, revealed a mean 3D wear of 0.23 mm (95%CI: 0.17-0.29) for HXLPE vs 0.41 mm (95%CI: 0.32-0.50) for conventional PE[9]. Based on this difference in wear of 0.18 mm, a SD of 0.21 and a power of 80%, a sample size of 21 patients was required in each group, to identify a statistically significant difference at the 0.05 significance level.

Surgical procedure

All THAs were performed in the Slotervaart Medical Center in the standardised way using a straight lateral approach, according to the surgical technique described by the manufacturer of the implants. All patients received the same uncemented acetabular cup (EP-FIT PLUS, Smith and Nephew) and a titanium uncemented Zweymuller femoral stem implant (SL-PLUS, Smith and Nephew) with the same ceramic head articulation (Biolox, Smith and Nephew). As inclination of > 45˚ gives more wear, the navigated position of the cup is aimed to be between 40 and 45˚ of inclination and 15 to 25˚ of anteversion[10,11]. In these series, computer navigation was used to determine this position (CT free navigation Galileo, Plus Orthopedic AG, Switzerland). The liner used was either a HXLPE liner (REXPOL, Smith and Nephew) or a standard PE liner (Standard PE, Smith and Nephew). Leg length and femoral offset were aimed to be identical to the contralateral side. In addition to this procedure, at least five well-scattered tantalum markers were installed (ø 1.0 mm) with a specially designed insertion instrument into the bone around the stem component to obtain skeletal landmarks.

RSA outcomes

Patient demographics were recorded at baseline. RSA evaluations were performed postoperatively, after receiving the same standard rehabilitation program, within one week, at 6 wk, 3 mo, 6 mo, 12 mo, 24 mo and 60 mo after implantation. RSA measurements were performed as described in the guidelines of Valstar et al[8], in the supine position using a uniplanar calibration box (Medis CarbonBox nr. 011, Medis Specials, Leiden, Netherlands). Analysis of the radiographic images was carried out with the model-based RSA Software, version 4.1 (RSAcore, Dept. of Orthopaedics, LUMC, Netherlands). The RSA system resulted in anteroposterior and lateral views of the hip simultaneously. The RSA at four to seven days postoperatively was used as a baseline. By using the implanted tantalum balls that were fixed in the bone around the implant, the position of the implant relative to the bone was accurately assessed using a model-based RSA technique (Figure 1). With this technique the 2D head penetration as a measure of linear wear was measured in millimetres by the proximal-distal migration (A-axis) and medial-lateral migration (B-axis). Using this penetration, the thickness of the inlay could be calculated in millimetres. Additionally, the anterior-posterior migration (C-axis) was measured to calculate 3D head penetration, to determine the volume. The volume of the PE inlay was determined (in mm3/year) to measure the number of millimetres of linear wear/year. Normally wear occurs in the upward direction, in a cylindrical shape. However, as the wear is not only in a neat upwards direction, but also in other angles or directions, a corrected volume was also calculated, according to the formula of Hashimoto[12]. This formula has been validated as the most accurate way to determine volume wear from linear wear[13].

Functional outcomes

The pain and activity of daily living (ADL) domains of the Hip Injury and Osteoarthritis Outcome Score (HOOS) were assessed pre-operatively, and after five years by a research nurse[14]. The HOOS was constructed to assess patient-relevant outcomes in five separate subscales: pain, symptoms, ADL, sport and recreation function and hip-related quality of life. The sum scores of the domains in this questionnaire are transformed into a zero to 100, worst to best scale. Another functional questionnaire assessed, was the Harris Hip Score (HHS)[15]. This questionnaire was focussed on pain and function, completed by range of motion and deformity. The maximum of 100 points is the best possible outcome.

Statistical analysis

Statistical analyses were performed with IBM SPSS Statistics version 26.0 (IBM Corp., Armonk, NY, United States). After confirmation of normal distribution, continuous variables are presented as mean ± SD. Categorical data are described as numbers with accompanying proportions. A mixed model analysis was performed to evaluate the amount of wear between both groups during follow-up. The effect of the different inlay was considered as a model factor and interaction with the follow-up time was evaluated to assess the differences in progression of wear in both inlays. The difference in wear at final follow-up was assessed by the Student’s t-test. To assess the differences of the PROMs between the inlay groups after five years, univariate as well as multivariate regression analyses were performed to adjust for potential confounders such as demographics. The differences were significant if the P-values were less than 0.05. All statistical methods in this study were performed by a biomedical statistical expert (Inger N Sierevelt).

RESULTS

A total of 51 consecutive patients were included in this study at baseline. Figure 2 shows a flow chart of the patients during this study. Seven patients were excluded, and the remaining 44 patients were included in our analysis; 22 in the REXPOL and 22 in the Standard PE insert group. During follow-up, five patients in the REXPOL group and three in the Standard PE group were lost to follow-up. The patient demographics and baseline characteristics of both groups were comparable and are shown in Table 2. No significant differences were seen in cup sizes between the two groups and no revisions were needed during follow-up in either group.

RSA migration

The total wear of the inlay measured from baseline showed less wear in all directions in the REXPOL group, which was significant in the REXPOL group in the latero-medial direction. All results of total wear measured from baseline are shown in Table 3. Due to a significant interaction between cup type and follow-up time, the wear pattern during follow-up of the REXPOL and Standard PE inlay were analysed separately. These wear patterns over the years showed greater wear in all directions in the conventional PE group, which is shown in Figures 3-6. The corresponding wear rates over this time period in Table 4 show that in all directions and volumes calculated, conventional PE had significant wear rates, whereas REXPOL did not show this outcome over time. The RSA images showed no signs of osteolysis.

Functional outcomes

The functional questionnaires were obtained at five years, to detect potential differences in functional outcomes. These results are shown in Table 5, with no significant differences observed.

DISCUSSION

The main finding of this study is that total 3D wear was less in the REXPOL group than in the standard PE group, with significant less wear in the medial direction after five years. Moreover, the wear rates in the medial and proximal direction and in both volume and corrected volume were significant in the standard PE group, but not in the REXPOL group.

Several in vivo studies have shown that HXLPE can reduce wear in comparison with normal PE inlays in THA[16]. However, only one study investigated the results of REXPOL in vivo, without randomisation and RSA analysis[17]. The outcomes in that study supported our findings of reduced wear in the REXPOL group, with approximately 70% less wear at the five-year follow-up. Therefore, this study is the first to present randomised clinical RSA data regarding the REXPOL liner.

In other in vivo studies using RSA, a systematic review performed by Callary et al[18] showed that only 12 cohorts comprising 260 THAs have compared the outcomes of HXLPE vs normal PE. Their recommendations on standardisation of reporting RSA outcomes are applied in our study. However, the studies included in their review assessed different inlays and not all of them were randomised. Thus, our study contributes to their statement that more longer-term standardised studies are needed to improve our understanding of the factors related to wear. Moreover, this will provide a better indication of the chance of osteolysis and as a result loosening of the cup and revision in the longer term.

A literature review by Dumbleton et al[19] showed that a threshold for wear of 0.05 mm/year would eliminate osteolysis. Although both standard PE and REXPOL showed wear rates below this threshold in our study, the long-term wear of REXPOL is still unclear. Long-term results were reported in the study by Broomfield et al[20] using another brand of HXPLE with the same low wear rates at 12 years. Rates of 0.03 mm/year were seen in the standard PE group and 0.003 mm/year in the HXLPE group, with higher wear rates in patients with osteolysis. Their long-term outcomes were supported by several studies showing ten-year or longer wear rates in favour of HXLPE[21-23]. Moreover, the study by Oparaugo et al[24] clarified the correlation between wear debris-induced osteolysis, volumetric wear-rates and revision.

Despite this correlation, subsequent concern was raised that HXLPE microparticles would show increased bioactivity in vivo as this had been observed in in vitro studies, since these particles are smaller than conventional PE[25-27]. However, the in vivo study by Lachiewicz et al[28] showed that at 10 to 14 years, small osteolytic lesions were also seen with HXLPE. Broomfield et al[20] supported this outcome and showed 50% osteolysis after 12-years with conventional PE vs 4% with HXPLE, which was statistically significant. This shows that HXLPE wear particles are not more biologically active than conventional PE and may not elevate the risk of osteolysis.

As osteolysis is one of the main reasons for loosening of the cup, the aforementioned results on long-term reduction of osteolysis become even more clinically relevant if a reduction in revisions is seen over time. The study by Hanna et al[29] showed less wear in the HXLPE group and as a result no osteolysis or revisions at 13-years in the HXPLE group with an implant survival rate of 100% vs 86% in the conventional PE group. De Steiger et al[30] also confirmed this in a large observational study and showed a 16-year cumulative revision rate of 11.7% with conventional PE vs 6.2% with HXLPE. The aforementioned outcomes confirm that the lower wear rates of HXLPE as seen in our study can reduce the risk of osteolysis when compared to conventional PE and can also reduce revision rates in the longer term.

By measuring wear as the slope of the amount of penetration in the different directions, some negative results on wear are seen in this study. These negative wear rate outcomes have been reported in previous studies[1,31-35]. However, this may be due to lower wear rates of HXLPE being harder to accurately measure compared to conventional PE. Although RSA is considered the best way to measure wear of a prothesis, it has an accuracy range of 0.022 mm to 0.086 mm, depending on the direction of measurement. In the case of HXPLE with even lower wear rates, it becomes more challenging to determine small amounts of wear[36]. Therefore, it becomes more important to have large cohorts to detect significant differences. As our study was carried out with small cohorts, the wear results should be interpreted while bearing this in mind. The expectation is that in the longer term these wear rates can be calculated more accurately for HXLPE, because they will be determined outwith the threshold. Since our RSA analysis was performed while in the supine position, another explanation of the negative wear rates is subluxation of the femoral head while lying. However, this was not confirmed by the review conducted by Callary et al[18] who showed no differences between the studies on supine or standing RSA. To overcome problems of negative wear results, long-term results of HXLPE wear are needed.

According to the outcomes in favour of HXLPE, this study confirms that the use of inlays such as the REXPOL, is preferred in THAs.

Limitations
The clinical outcomes of our study were measured to assess any major drawbacks of standard PE or HXLPE at the five-year follow-up. As improvement from baseline was seen in both groups with no significant differences between the groups, no practical disadvantages were seen by preferring one inlay over the other. As the study was not powered by clinical outcomes, further research is needed to investigate these outcomes.

Implications for further research

This study showed, for the first time, that REXPOL resulted in less wear in the short-term in a randomised setting by RSA. Therefore, further investigation of wear over a longer period should be performed, to confirm that REXPOL can reduce the risk of osteolysis and consequently reduce revision rates in THA. Also, more research needs to be carried out to overcome problems of minimal differences in wear rates. In addition, research on other variables that influence wear such as activity, weight and surgical factors such as inclination of the acetabular component should be performed.

CONCLUSION

Total 3D wear is less with REXPOL inlay than with conventional PE inlay in THAs after five years. This study confirmed, for the first time, that the REXPOL HXLPE inlay is preferred to the standard PE inlay.

ARTICLE HIGHLIGHTS

Research background

Highly cross-linked polyethylene (HXLPE) inlay in total hip arthroplasty (THA) is thought to result in lower wear rates in vivo, compared to conventional polyethylene.

Research motivation

More in vivo studies are needed, especially those using Roentgen stereophotogrammetric analysis (RSA), to confirm the advantage of HXLPE over conventional polyethylene (PE).

Research objectives

The objective of the study was to compare wear of the HXLPE (REXPOL) and conventional PE acetabular inlay with similar ceramic head articulation, within the first five years after implantation.

Research methods

A double-blind randomised controlled trial was performed to compare wear of REXPOL, a HXLPE, with conventional PE within the first five years after implantation using RSA.

Research results

The HXLPE (REXPOL) inlay showed less wear in the latero-medial direction. Significant wear rates of the conventional PE inlay were seen in the latero-medial and center-proximal direction and in volume and corrected volume, whereas the REXPOL inlay did not show this outcome over time.

Research conclusions

Total 3D wear is less with the REXPOL (HXLPE) inlay than with the conventional PE inlay in THAs after five years.

Research perspectives

Further investigations into the long-term wear and factors that might influence wear rates should be conducted, to confirm that HXLPE (REXPOL) can reduce the risk of osteolysis and consequently reduce revision rates in THA. In addition, an investigation of the impact of wear reduction by HXPLE on the functional outcomes of patients is required.

REFERENCES

1 McCalden RW, MacDonald SJ, Rorabeck CH, Bourne RB, Chess DG, Charron KD. Wear rate of highly cross-linked polyethylene in total hip arthroplasty. A randomized controlled trial. J Bone Joint Surg Am 2009; 91: 773-782 [PMID: 19339560 DOI: 10.2106/JBJS.H.00244]

2 Haverkamp D. The latest information on total hip arthroplasty bearing surfaces. Minerva Ortopedica e Traumatologica 2009; 60: 233-240

3 Green TR, Fisher J, Stone M, Wroblewski BM, Ingham E. Polyethylene particles of a 'critical size' are necessary for the induction of cytokines by macrophages in vitro. Biomaterials 1998; 19: 2297-2302 [PMID: 9884043 DOI: 10.1016/s0142-9612(98)00140-9]

4 Chuter GS, Cloke DJ, Mahomed A, Partington PF, Green SM. Wear analysis of failed acetabular polyethylene: a comparison of analytical methods. J Bone Joint Surg Br 2007; 89: 273-279 [PMID: 17322451 DOI: 10.1302/0301-620X.89B2.18158]

5 McKellop H, Shen FW, Lu B, Campbell P, Salovey R. Development of an extremely wear-resistant ultra high molecular weight polyethylene for total hip replacements. J Orthop Res 1999; 17: 157-167 [PMID: 10221831 DOI: 10.1002/jor.1100170203]

6 Atienza C Jr, Maloney WJ. Highly cross-linked polyethylene bearing surfaces in total hip arthroplasty. J Surg Orthop Adv 2008; 17: 27-33 [PMID: 18284901]

7 Stilling M, Kold S, de Raedt S, Andersen NT, Rahbek O, Søballe K. Superior accuracy of model-based radiostereometric analysis for measurement of polyethylene wear: A phantom study. Bone Joint Res 2012; 1: 180-191 [PMID: 23610688 DOI: 10.1302/2046-3758.18.2000041]

8 Valstar ER, Gill R, Ryd L, Flivik G, Börlin N, Kärrholm J. Guidelines for standardization of radiostereometry (RSA) of implants. Acta Orthop 2005; 76: 563-572 [PMID: 16195075 DOI: 10.1080/17453670510041574]

9 Digas G, Kärrholm J, Thanner J, Malchau H, Herberts P. The Otto Aufranc Award. Highly cross-linked polyethylene in total hip arthroplasty: randomized evaluation of penetration rate in cemented and uncemented sockets using radiostereometric analysis. Clin Orthop Relat Res 2004; 6-16 [PMID: 15577460]

10 Hirakawa K, Mitsugi N, Koshino T, Saito T, Hirasawa Y, Kubo T. Effect of acetabular cup position and orientation in cemented total hip arthroplasty. Clin Orthop Relat Res 2001; : 135-142 [PMID: 11451112 DOI: 10.1097/00003086-200107000-00020]

11 Patil S, Bergula A, Chen PC, Colwell CW Jr, D'Lima DD. Polyethylene wear and acetabular component orientation. J Bone Joint Surg Am 2003; 85-A Suppl 4: 56-63 [PMID: 14652394 DOI: 10.2106/00004623-200300004-00007]

12 Hashimoto Y. Polyethylene wear in total hip arthroplasty: volumetric wear measurement of retrieved acetabular components. 41th Annual Meeting Orthopaedic Research Society 1995; 1995: 116-120

13 Mizoue T, Yamamoto K, Masaoka T, Imakiire A, Akagi M, Clarke IC. Validation of acetabular cup wear volume based on direct and two-dimensional measurements: hip simulator analysis. J Orthop Sci 2003; 8: 491-499 [PMID: 12898299 DOI: 10.1007/s00776-003-0674-2]

14 Nilsdotter AK, Lohmander LS, Klässbo M, Roos EM. Hip disability and osteoarthritis outcome score (HOOS)--validity and responsiveness in total hip replacement. BMC Musculoskelet Disord 2003; 4: 10 [PMID: 12777182 DOI: 10.1186/1471-2474-4-10]

15 Harris WH. Traumatic arthritis of the hip after dislocation and acetabular fractures: treatment by mold arthroplasty. An end-result study using a new method of result evaluation. J Bone Joint Surg Am 1969; 51: 737-755 [PMID: 5783851]

16 Kurtz SM, Gawel HA, Patel JD. History and systematic review of wear and osteolysis outcomes for first-generation highly crosslinked polyethylene. Clin Orthop Relat Res 2011; 469: 2262-2277 [PMID: 21431461 DOI: 10.1007/s11999-011-1872-4]

17 Orradre Burusco I, Romero R, Brun M, López Blasco JJ. Cross-linked ultra-high-molecular weight polyethylene liner and ceramic femoral head in total hip arthroplasty: a prospective study at 5 years follow-up. Arch Orthop Trauma Surg 2011; 131: 1711-1716 [PMID: 21739116 DOI: 10.1007/s00402-011-1340-3]

18 Callary SA, Solomon LB, Holubowycz OT, Campbell DG, Munn Z, Howie DW. Wear of highly crosslinked polyethylene acetabular components. Acta Orthop 2015; 86: 159-168 [PMID: 25301435 DOI: 10.3109/17453674.2014.972890]

19 Dumbleton JH, Manley MT, Edidin AA. A literature review of the association between wear rate and osteolysis in total hip arthroplasty. J Arthroplasty 2002; 17: 649-661 [PMID: 12168184 DOI: 10.1054/arth.2002.33664]

20 Broomfield JA, Malak TT, Thomas GE, Palmer AJ, Taylor A, Glyn-Jones S. The Relationship Between Polyethylene Wear and Periprosthetic Osteolysis in Total Hip Arthroplasty at 12 Years in a Randomized Controlled Trial Cohort. J Arthroplasty 2017; 32: 1186-1191 [PMID: 27998657 DOI: 10.1016/j.arth.2016.10.037]

21 Johanson PE, Digas G, Herberts P, Thanner J, Kärrholm J. Highly crosslinked polyethylene does not reduce aseptic loosening in cemented THA 10-year findings of a randomized study. Clin Orthop Relat Res 2012; 470: 3083-3093 [PMID: 22669546 DOI: 10.1007/s11999-012-2400-x]

22 Röhrl SM, Nivbrant B, Nilsson KG. No adverse effects of submelt-annealed highly crosslinked polyethylene in cemented cups: an RSA study of 8 patients 10 yaers after surgery. Acta Orthop 2012; 83: 148-152 [PMID: 22248172 DOI: 10.3109/17453674.2011.652889]

23 Teeter MG, Lanting BA, Naudie DD, McCalden RW, Howard JL, MacDonald SJ. Highly crosslinked polyethylene wear rates and acetabular component orientation: a minimum ten-year follow-up. Bone Joint J 2018; 100-B: 891-897 [PMID: 29954202 DOI: 10.1302/0301-620X.100B7.BJJ-2017-1457.R3]

24 Oparaugo PC, Clarke IC, Malchau H, Herberts P. Correlation of wear debris-induced osteolysis and revision with volumetric wear-rates of polyethylene: a survey of 8 reports in the literature. Acta Orthop Scand 2001; 72: 22-28 [PMID: 11327409 DOI: 10.1080/000164701753606644]

25 Fisher J, McEwen HM, Tipper JL, Galvin AL, Ingram J, Kamali A, Stone MH, Ingham E. Wear, debris, and biologic activity of cross-linked polyethylene in the knee: benefits and potential concerns. Clin Orthop Relat Res 2004; : 114-119 [PMID: 15534530 DOI: 10.1097/01.blo.0000148783.20469.4c]

26 Ingram JH, Stone M, Fisher J, Ingham E. The influence of molecular weight, crosslinking and counterface roughness on TNF-alpha production by macrophages in response to ultra high molecular weight polyethylene particles. Biomaterials 2004; 25: 3511-3522 [PMID: 15020125 DOI: 10.1016/j.biomaterials.2003.10.054]

27 Ries MD, Scott ML, Jani S. Relationship between gravimetric wear and particle generation in hip simulators: conventional compared with cross-linked polyethylene. J Bone Joint Surg Am 2001; 83-A Suppl 2 Pt 2: 116-122 [PMID: 11712832 DOI: 10.2106/00004623-200100022-00009]

28 Lachiewicz PF, Soileau ES, Martell JM. Wear and Osteolysis of Highly Crosslinked Polyethylene at 10 to 14 Years: The Effect of Femoral Head Size. Clin Orthop Relat Res 2016; 474: 365-371 [PMID: 25903943 DOI: 10.1007/s11999-015-4319-5]

29 Hanna SA, Somerville L, McCalden RW, Naudie DD, MacDonald SJ. Highly cross-linked polyethylene decreases the rate of revision of total hip arthroplasty compared with conventional polyethylene at 13 years' follow-up. Bone Joint J 2016; 98-B: 28-32 [PMID: 26733512 DOI: 10.1302/0301-620X.98B1.36527]

30 de Steiger R, Lorimer M, Graves SE. Cross-Linked Polyethylene for Total Hip Arthroplasty Markedly Reduces Revision Surgery at 16 Years. J Bone Joint Surg Am 2018; 100: 1281-1288 [PMID: 30063590 DOI: 10.2106/JBJS.17.01221]

31 Bragdon CR, Barrett S, Martell JM, Greene ME, Malchau H, Harris WH. Steady-state penetration rates of electron beam-irradiated, highly cross-linked polyethylene at an average 45-month follow-up. J Arthroplasty 2006; 21: 935-943 [PMID: 17027534 DOI: 10.1016/j.arth.2006.01.006]

32 Callary SA, Field JR, Campbell DG. Low wear of a second-generation highly crosslinked polyethylene liner: a 5-year radiostereometric analysis study. Clin Orthop Relat Res 2013; 471: 3596-3600 [PMID: 23893361 DOI: 10.1007/s11999-013-3188-z]

33 Campbell DG, Field JR, Callary SA. Second-generation highly cross-linked X3™ polyethylene wear: a preliminary radiostereometric analysis study. Clin Orthop Relat Res 2010; 468: 2704-2709 [PMID: 20151231 DOI: 10.1007/s11999-010-1259-y]

34 Engh CA Jr, Stepniewski AS, Ginn SD, Beykirch SE, Sychterz-Terefenko CJ, Hopper RH Jr, Engh CA. A randomized prospective evaluation of outcomes after total hip arthroplasty using cross-linked marathon and non-cross-linked Enduron polyethylene liners. J Arthroplasty 2006; 21: 17-25 [PMID: 16950057 DOI: 10.1016/j.arth.2006.05.002]

35 Manning DW, Chiang PP, Martell JM, Galante JO, Harris WH. In vivo comparative wear study of traditional and highly cross-linked polyethylene in total hip arthroplasty. J Arthroplasty 2005; 20: 880-886 [PMID: 16230239 DOI: 10.1016/j.arth.2005.03.033]

36 Bragdon CR, Malchau H, Yuan X, Perinchief R, Kärrholm J, Börlin N, Estok DM, Harris WH. Experimental assessment of precision and accuracy of radiostereometric analysis for the determination of polyethylene wear in a total hip replacement model. J Orthop Res 2002; 20: 688-695 [PMID: 12168656 DOI: 10.1016/S0736-0266(01)00171-1]

Footnotes

Institutional review board statement: This single center RCT was granted ethical approval by the local ethics committee review board of Slotervaart Hospital (NL23524.048.08).

Clinical trial registration statement: This study is registered at the Dutch Trial Register. The registration identification number is NL5605.

Informed consent statement: All study participants, or their legal guardian, provided informed written consent prior to study enrollment.

Conflict-of-interest statement: The authors declare that there is no conflict of interest.

Data sharing statement: No additional data are available.

CONSORT 2010 statement: The design and reporting of this study were carried out in accordance with the Consolidated Standards of Reporting Trials (CONSORT) principles (see attached CONSORT checklist).

Open-Access: This article is an open-access article that was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution NonCommercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

Manuscript source: Unsolicited Manuscript

Corresponding Author's Membership in Professional Societies: ESSKA-AFAS.

Peer-review started: May 19, 2020

First decision: May 26, 2020

Article in press:

Specialty type: Orthopedics

Country/Territory of origin: Netherlands

Peer-review report’s scientific quality classification

Grade A (Excellent): 0

Grade B (Very good): B

Grade C (Good): C

Grade D (Fair): 0

Grade E (Poor): 0

P-Reviewer: Cho MR, Cankaya D S-Editor: Zhang H L-Editor: Webster JR E-Editor:

Figure Legends

Figure 1 Model of Roentgen stereophotogrammetric analysis technique. Model of Roentgen stereophotogrammetric analysis technique on right-sided acetabular component after insertion of tantalum markers, by measuring the penetration of the head in the proximal-distal (A-axis), medial-lateral (B-axis) and anterior-posterior migration (C-axis) direction.

Figure 2 Flow chart of follow-up. PE: Polyethylene; RSA: Roentgen stereophotogrammetric analysis.

Figure 3 Wear of the inlay in the medial direction in mm over time (months).

Figure 4 Wear of the inlay in the proximal direction in mm over time (months).

Figure 5 Volumetric wear of the inlay in mm3 over time (months).

Figure 6 Corrected volumetric wear of the inlay in mm3 over time (months).

Table 1 Inclusion and exclusion criteria

| Inclusion | Exclusion |
| --- | --- |
| Primary arthroplasty due to: | Patients who recently suffered: |
| Primary osteoarthritis | Post-operative osteoarthritis |
| Avascular necrosis | Charnley C osteoarthritis |
| Femoral neck fracture | Infection of the hip |
| Hip dysplasia |  |
| Age between 60 - 75 yr at surgery | Prior osteotomy or arthroplasty of the affected hip |
| Willing to comply with the post-operative review program | Under treatment for osteoporosis |
|  | Body mass index > 35 kg/m2 |
|  | Requiring cortisone medication |

Table 2 Patient and baseline characteristics

|  | Conventional PE | REXPOL |
| --- | --- | --- |
| Number of patients, n (%) | 25 (49) | 26 (51) |
| Gender, n (%) |  |  |
| Male | 10 (40) | 10 (39) |
| Female | 15 (60) | 16 (59) |
| BMI, mean ± SD kg/m2 | 26.7 ± 2.9 | 27.2 ± 3.4 |
| Age at operation in years, mean ± SD | 68.5 ± 4.6 | 68.6 ± 5.1 |
| HOOS pain, mean ± SD | 46.5 ± 21.6 | 51.1 ± 17.7 |
| HOOS ADL, mean ± SD | 41.1 ± 16.6 | 46.0 ± 17.2 |
| HHS, mean ± SD | 50.6 ± 12.7 | 54.6 ± 10.9 |

PE: Polyethylene; BMI: Body mass index; HOOS: Hip injury and osteoarthritis outcome score; ADL: Activity of daily living; HHS: Harris hip score.

Table 3 Total wear in all directions at the five-year follow-up, presented as mean with ranges

|  | Conventional PE | REXPOL | P value |
| --- | --- | --- | --- |
| Medial (mm) | -0.128 (-0.202 to -0.054) | 0.013 (-0.054 to 0.081) | 0.006 |
| Proximal (mm) | 0.196 (0.054 to 0.338) | 0.017 (-0.120 to 0.154) | 0.07 |
| Volume (mm3) | 113.39 (-2.48 to 229.26) | -30.59 (-167.16 to 105.98) | 0.10 |
| Corrected volume (mm3) | 121.6 (17.76 to 225.46) | -12.45 (-122.94; 98.03) | 0.07 |

PE: Polyethylene.

Table 4 Mean wear rates per year, presented as mean with ranges

|  | Conventional PE | P value | REXPOL | P value |
| --- | --- | --- | --- | --- |
| Medial (mm/yr) | -0.021 (-0.028 to -0.015) | < 0.001 | 0.004 (-0.002 to 0.009) | 0.21 |
| Proximal (mm/yr) | 0.033 (0.018 to 0.047) | < 0.001 | 0.003 (-0.011 to 0.017) | 0.66 |
| Volume (mm3/yr) | 15.94 (2.729 to 29.15) | 0.02 | -5.545 (-20.36 to 9.269) | 0.46 |
| Corrected volume (mm3/yr) | 18.53 (7.188 to 29.86) | 0.002 | -2.142 (-14.13 to 9.841) | 0.72 |

PE: Polyethylene.

| Table 5 Functional outcomes at the five-year follow-up, presented as mean with 95%CI |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Univariate |  |  | Multivariate |  |
|  | Conventional PE (n = 17) | REXPOL (n = 17) | P value | Adjusted β-coefficient | P value |
| HOOS pain | 93.8 (86.8; 100) | 85.9 (77.1; 94.7) | 0.15 | -3.3 (-14.9; 8.3) | 0.57 |
| HOOS ADL | 89.0 (81.2: 96.8) | 77.6 (66.6; 88.5) | 0.08 | -8.6 (-23.3; 6.1) | 0.24 |
| HHS | 89.7 (83.0; 96.4) | 86.5 (78.8; 94.2) | 0.51 | 0.15 (-10.1; 10.4) | 0.98 |

PE: Polyethylene; HOOS: Hip injury and osteoarthritis outcome score; ADL: Activity of daily living; HHS: Harris hip score.