Name of Journal: World Journal of Biological Chemistry

Manuscript NO: 84265

Manuscript Type: ORIGINAL ARTICLE

Basic Study

Comparison of the conventional tube and erythrocyte-magnetized technology in titration of red blood cell alloantibodies

He XH et al. EMT in red blood cell alloantibody titration

Xue-Hua He, Hong Yan, Chun-Yan Wang, Xue-Yun Duan, Jia-Jia Qiao, Xiao-Jun Guo, Hong-Bin Zhao, Dong Ren, Jian-She Li, Qiang Zhang

Xue-Hua He, Jia-Jia Qiao, Xiao-Jun Guo, Hong-Bin Zhao, Dong Ren, Department of Blood Transfusion, Shanxi Bethune Hospital, Shanxi Academy of Medical Sciences, Tongji Shanxi Hospital, Third Hospital of Shanxi Medical University, Taiyuan 030032, Shanxi Province, China

Hong Yan, Jian-She Li, Department of Blood Transfusion, The Second Hospital of Shanxi Medical University, Taiyuan 030001, Shanxi Province, China

Chun-Yan Wang, Department of Blood Transfusion, Shanxi Cancer Hospital, Taiyuan 030013, Shanxi Province, China

Xue-Yun Duan, Department of Blood Transfusion, Shanxi Cardiovascular Hospital, Taiyuan 030024, Shanxi Province, China

Qiang Zhang, Department of Clinical Laboratory, Taiyuan Blood Center, Institute of Blood Transfusion Technology, Taiyuan 030024, Shanxi Province, China

Author contributions: He XH and Yan H contributed equally; He XH, Qiao JJ, and Guo XJ designed the study; Zhao HB, and Ren D wrote the manuscript; Yan H, and Li JS reviewed and edited; Wang CY projected administration; Duan XY performed the experiments; Zhang Q analysed the data; All authors have read and agreed to the published version of the manuscript.

Supported by Project of Shanxi Provincial Health Commission, No. 2021144.

Corresponding author: Qiang Zhang, BSc, Chief Physician, Department of Clinical Laboratory, Taiyuan Blood Center, Institute of Blood Transfusion Technology, No. 185 Yingze West Street, Taiyuan 030024, Shanxi Province, China. qi59794383@163.com

Received: March 14, 2023

Revised: April 16, 2023

Accepted: May 15, 2023

Published online: May 27, 2023

Abstract

BACKGROUND

Erythrocyte alloantibodies are mainly produced after immune stimulation, such as blood transfusion, pregnancy, and transplantation, and are the leading causes of severe hemolytic transfusion reactions and difficulty in blood grouping and matching. Therefore, antibody screening is critical to prevent and improve red cell alloantibodies. Routine tube assay is the primary detection method of antibody screening. Recently, erythrocyte-magnetized technology (EMT) has been increasingly used in clinical practice. This study intends to probe the application and efficacy of the conventional tube and EMT in red blood cell alloantibody titration to provide a reference for clinical blood transfusion.

AIM

To investigate the application value of conventional tube and EMT in red blood cell alloantibody titration and enhance the safety of blood transfusion practice.

METHODS

A total of 1298 blood samples were harvested from blood donors at the Department of Blood Transfusion of our hospital from March 2021 to December 2022. A 5 mL blood sample was collected in tubing, which was then cut, and the whole blood was put into a test tube for centrifugation to separate the serum. Different red blood cell blood group antibody titers were simultaneously detected using the tube polybrene test, tube antiglobulin test (AGT), and EMT screening irregular antibody methods to determine the best test method.

RESULTS

Simultaneous detection was performed through the tube polybrene test, tube AGT and EMT screening irregular antibodies. It was discovered that the EMT screening irregular antibody method could detect all immunoglobulin G (IgG) and immunoglobulin M (IgM) irregular antibodies, and the results of manual tube AGT were satisfactory, but the operation time was lengthy, and the equipment had a large footprint. The EMT screening irregular antibody assay was also conducted to determine its activity against type O Rh (D) red blood cells, and the outcomes were satisfactory. Furthermore, compared to the conventional tube method, the EMT screening irregular antibody method was more cost-effective and had significantly higher detection efficiency.

CONCLUSION

With a higher detection rate, the EMT screening irregular antibody method can detect both IgG and IgM irregular antibodies faster and more effectively than the conventional tube method.

Key Words: Erythrocyte-magnetized technology; Conventional tube; Red blood cell alloantibodies; Transfusion reactions; Application

©The Author(s) 2023. Published by Baishideng Publishing Group Inc. All rights reserved.

Citation: He XH, Yan H, Wang CY, Duan XY, Qiao JJ, Guo XJ, Zhao HB, Ren D, Li JS, Zhang Q. Comparison of the conventional tube and erythrocyte-magnetized technology in titration of red blood cell alloantibodies. World J Biol Chem 2023; 14(3): 62-71

URL: https://www.wjgnet.com/1949-8454 /full/v14/i3/62.htm

DOI: https://dx.doi.org/10.4331/wjbc.v14.i3.62

Core Tip: Irregular antibody screening has long been a routine blood test for blood donors in numerous developed countries. However, only a few blood stations in China have tried using a saline medium for this type of screening. Monoclonal anti-A (B) is a standard reagent for ABO blood grouping, but false positive or false negative reactions can occur, reducing the accuracy of the test. With the improvement of diagnostic techniques and medical levels, the erythrocyte-magnetized technology (EMT) screening irregular antibody method has been gradually applied in a range of clinical settings. This study analyzed blood samples from voluntary blood donors to explore the application value and effect of the conventional tube method and EMT in red blood cell alloantibody titration, with the goal of providing valuable references to improve the safety of clinical blood transfusion.

INTRODUCTION

Antibody screening aims to detect the presence of antibodies other than the ABO blood group in patients' serum or plasma by serum determination. This includes autoantibodies, drug antibodies, and allospecific antibodies to red blood cell blood groups[1,2]. Red blood cell alloantibodies, which are blood group antibodies other than anti-A and anti-B, are referred to as irregular antibodies. These antibodies are mainly produced after immune stimulation, such as blood transfusion, pregnancy, and transplantation, and are the primary cause of severe hemolytic transfusion reactions, difficulties in blood grouping, and complex matching[3-5]. Therefore, improving the screening rate of red blood cell alloantibodies is crucial, as it has significant implications for the safety of clinical blood transfusion.

Although many developed countries have been including irregular antibody screening in routine blood testing for blood donors, only a few blood stations in China have experimented with irregular antibody screening using saline media[6]. False positive or negative reactions sometimes occur using monoclonal anti-A (B) as a standard ABO blood grouping reagent[7]. The Collies automatic blood grouping system has recently been utilized in clinical practice employing erythrocyte-magnetized technology (EMT), an international advanced technique. This technique magnetizes red blood cells, causing rapid sedimentation of red blood cells under the attraction of a magnetic field, thus replacing the serological centrifuge[8]. Irregular antibody screening using this instrument has been reported to be an indirect antiglobulin test (AGT) method that combines solid-phase packaging antiglobulin with EMT. The plasma to be tested was incubated with ready-to-use irregular antibody screening red blood cells, which can be sensitized by irregular antibodies in the plasma. In the presence of a magnetic field, magnetized red blood cells migrate to the bottom of the microplate and react with antiglobulin, forming an evenly distributed red blood cell layer[9,10]. Earlier studies have demonstrated that EMT can effectively mitigate the impact of centrifugation on experimental results, making the results more reliable and saving the costs associated with centrifuge calibration and maintenance[11,12].

Based on this, this study analyzed blood samples from voluntary blood donors to explore the application value and effect of the conventional tube and EMT in red blood cell alloantibody titration, hoping to provide valuable references for the safety of clinical blood transfusion.

MATERIALS AND METHODS

General information

We randomly selected 1298 blood donors from the Department of Blood Transfusion at our hospital between March 2021 and December 2022. The donors were between 18 and 55, with 612 males and 686 females. Among the donors, 897 (69%) were between 25 and 40. We separated serum from 5 mL of whole blood using centrifugation.

Materials and reagents

Reagents and instruments: Anti-human globulin cards (Jiangsu Libo Pharmaceutical Biotechnology Co., Ltd., batch number: 202208001), red blood cell antibody screening cell kits (batch number: 20187012, 20187014), and red blood cell antibody identification spectrum cells (REAGENS, batch number: 726000, 733000) were provided by Shanghai Blood Biomedicine Co., Ltd. Special centrifuge for blood immunology 2005-2 Zhuhai Bezo Biotechnology Co., Ltd. Polybrene came from reagents and instruments company of Zhuhai Bezo Biotechnology Co., Ltd. Collies 3 (QWALYS3) automatic blood analyzer (DIAGAST, France), Deakin (EvO-2) immunoassay sample adding system (TECAN, Switzerland). Shanghai Medical Equipment Co., Ltd. Digital explicit water bath, Heidolph oscillator (Titramax101) made in Germany, and HeShi centrifuge with hanging micro reaction plates in the German obstetrics department.

Eleven known irregular antibodies with known specificity against D, E, C, c, e, S, Fy', JK', M, N, and phosphatidylinositol antibody were prepared by the Institute of Blood Transfusion Technology, Taiyuan, China.

The TSFT puncture kit is mainly composed of low ionic medium, polybrene application solution, neutralization solution, and positive control (1gG anti-D), which are provided by Shanghai Blood Bio-Pharmaceutical Co., Ltd. and Zhuhai Bezo Biotechnology Co., Ltd., respectively.

Methods

The following procedures were used to screen and identify antibodies in blood donors using the all-tube method: Equal amounts of sera from 10 blood donors were mixed and identified for cell reactions with antibodies using the tube method Treponema pallidum particle test for antibody screening. If a positive reaction was observed, antibody screening was repeated on the 10 original samples before mixing using the traditional AGT. For those who tested positive, irregular antibodies were identified using antibody identification cells, and an absorption and diffusion test was performed if necessary to determine antibody specificity.

The tube polybrene test was conducted as follows[13]: One drop (50 μL) of reagent red blood cells was mixed with two drops (100 μL) of serum or plasma and incubated at room temperature for 15 s at 3000 rpm. The results were double-checked and recorded. Agglutination indicated a positive reaction, while the absence of agglutination indicated a negative reaction.

The tube AGT method was performed as follows[14]: After observing the results using the above methods, the red blood cell serum or plasma mixture was incubated at 37 °C for 30 min. The red blood cells were then washed three times with saline and finally, the packed red blood cells were suspended with two drops (100 mL) of antiglobulin and centrifuged. The outcomes were evaluated according to the same criteria as the upper water tube polybrene test method.

EMT screening for irregular antibodies[15]: With the use of the Collies automatic blood grouping system, with a special program matched with the original reagent, the instrument loads a 96-well ScreenLys microplate coated with antiglobulin, adds SONL isolation solution, diluent, 151 L magnetized screening cells, and 15 μL specimen plasma, completes the closed incubation at 37 °C after sample addition according to the set program, and completes the irregular antibody screening determination of the specimen by magnetization shaking, photography, and interpretation. Figure 1 shows the test results.

Antibody potency assay: The irregular antibody was diluted by doubling with antibody diluent and reacted with the reagent red blood cells with corresponding positive antigens, and its titer was measured. The titer endpoint is the highest dilution with agglutination ≥ 1 +. These reagent red blood cells were also reacted with normal donor sera as a negative control test.

Statistical analysis

All data from this study were processed and analyzed using IBM SPSS 21.0 software (SPSS Inc., Chicago, IL, United States), and a t-test was harnessed to compare groups. If P < 0.05, statistical significance was confirmed.

RESULTS

Screening and identification results of irregular antibodies in blood donors

A total of 1298 random blood donors were screened using a commercial kit, and the polybrene test performed irregular antibody detection. The results revealed that 18 cases were positive, with a positive rate of 1.39%. Further identification showed the presence of Rh, MN, and other blood group system antibodies, as detailed in Table 1.

Comparison of sensitivity for detecting irregular antibodies with known specificity

A total of 11 known irregular antibody titers were detected in parallel by tube polybrene test, tube AGT, and EMT screening irregular antibody methods, respectively, while negative quality control tests were performed. It was found that all immunoglobulin G (IgG) and immunoglobulin M (IgM) irregular antibodies could be detected by the EMT screening irregular antibody method, and the manual tube AGT results in the control group were satisfactory, but the operation time was long, and the equipment occupied a large space (such as tubes and racks), as shown in Table 2.

Comparison of irregular antibody screening positive samples from the blood donor

The 18 positive samples selected from the blood donor spectrum were detected in parallel by the tube polybrene test, tube ACT, and EMT screening irregular antibody method. It was discovered that the reaction pattern of all sera and commercial screening cells or spectrum cells was consistent.

Effect of hemolysis and lipemia on irregular antibodies for EMT screening

The physical examination of blood donors discovered macroscopically significant severe hemolysis and lipemia samples. In equal proportion, the supernatant sera from these samples were mixed with low concentrations (titer 1:8) of IgG anti-D and AB sera (negative samples). In addition, their activities against type O Rh (D) red blood cells were determined by EMT screening irregular antibodies. The results indicated that the staining failed, but the cell membrane permeability remained normal, suggesting that the cells were active.

Comparison of cost of reagents, equipment, and manpower consumption

The consumption cost of screening irregular antibodies in the aforementioned blood donors was analyzed using the tube polybrene test, tube ACT and EMT screening methods. The cost was calculated by dividing the cost of each method by the number of blood donors tested. The outcomes showed that the EMT screening irregular antibody method cost less than 0.5 RMB per blood donor, while the tube polybrene test and tube AGT methods cost more than 5 RMB and 5.5 RMB per blood donor, respectively. These findings demonstrated that the EMT screening method was cost-effective and had advanced detection technology. Additionally, it allowed for sample concentration, resulting in improved work efficiency (see Table 3 for details).

DISCUSSION

At present, blood transfusion therapy, as a clinical treatment, is still irreplaceable in the treatment of some diseases. Although some advanced treatment options have significantly reduced the dependence on blood products in the treatment after being applied in clinical practice, hemolytic transfusion reactions and cross-matching difficulties caused by blood group antibodies other than ABO, that is, irregular antibodies, also occur from time to time[16,17]. Blood group antibodies destroy mismatched red blood cells in transfused blood or shorten their lifespan, producing hemolytic transfusion reactions, which may affect treatment outcomes or endanger the patient's life. Studies have unveiled that neonatal hemolytic disease caused by prenatal irregular blood group antibodies in pregnant women, especially Rh-HDN, has severe symptoms and often leads to serious harm[18,19]. Furthermore, monoclonal hyperactivity-A (B) reagent gas reportedly reacts with most red blood cells with high sensitivity and specificity[20]. Malformed hemolytic transfusion reactions caused by ABO blood group incompatibility are typically uncommon. In contrast, adverse transfusion reactions, neonatal hemolysis, and difficulty in blood grouping caused by irregular antibodies to ABO blood group accidents of red blood cells occasionally occur[21,22]. In addition, it has been internationally reported that monoclonal reagents can lead to false agglutination or false negative phenomenon during ABO blood group detection, and there have been corresponding reports in China in recent years[23]. Therefore, blood grouping alone is insufficient before transfusion, and antibody screening is necessary to ensure blood transfusion safety.

In recent years, with the development of medical technology, automatic blood grouping has been more and more widely used in the blood transfusion departments of major hospitals and blood collection institutions at all levels. It also plays a vital role in blood grouping tests, irregular antibody screening tests, and clinical cross-matching tests[24]. EMT is a new technology based on red blood cell magnetization, which has been maturely applied in immunodiagnosis, cell separation, protein purification, and nucleic acid extraction, and has also improved the automation rate of laboratories[25,26]. Based on this, this study aimed to investigate the effect of the conventional tube and EMT in red blood cell alloantibody titration, hoping to provide some help for clinically safe blood transfusion.

This study collected 5 mL of whole blood from 1298 blood donors. The serum was separated by centrifugation. Different red blood cell blood group antibody titers were detected in parallel using the tube polybrene test, tube AGT, and EMT screening irregular antibody. Usually, voluntary blood donor samples are centrally collected. In order to save labor and reagents, 5-10 samples can be mixed for primary screening, and positive samples can be reexamined one by one to detect antibody-positive samples. However, because a few irregular antibodies with low titer in the mixed plasma still have the possibility of missed detection, the mixed plasma primary screening method is not suitable for the pretransfusion test of patients. In addition, it was found that all irregular antibodies of IgG and IgM nature were detected by EMT screening irregular antibody method. Nevertheless, the tube polybrene test method and tube AGT method did not detect irregular antibodies of low concentration, which may be related to the lack of sensitivity of IgM active blood group antibodies in low concentrations and at room temperature in the polybrene method and AGT.

Red cell alloantibodies are classified based on their properties as macromolecular IgM or small molecular IgG[27]. There are many methods to detect red blood cell antibodies, but standard tube methods can only detect IgM antibodies in most cases[28,29]. The results showed that the EMT screening irregular antibody method had clear advantages in determining the activity of IgG O Rh (D) red blood cell antibodies. In recent years, the use of traditional in vitro AGT has become more popular in China[30]. The findings in this paper indicate that the EMT screening irregular antibody method, the in vitro polybrene test method, and the in vitro AGT method can be utilied for routine screening of irregular antibodies in blood donors. Among these methods, the EMT screening irregular antibody method is more rapid, convenient, efficient, and less expensive.

In summary, compared with the in vitro polybrene test method and in vitro AGT method, the Collies automatic blood group system based on EMT for irregular antibody screening has the characteristics of more robust antibody specificity, more convenient operation, more accurate results, and complete lake source, which is suitable for large-scale screening of irregular antibodies in blood stations and is worthy of being widely popularized in clinical practice. Of course, more sensitive test techniques help to improve the antibody detection rate, but there may still be some antibodies that are not easy. In daily testing, the recorded history of checking past antibodies should be listed as a part of routine compatibility tests. Hospitals at all levels should actively understand the patient's blood transfusion history, pregnancy history, and current history and perform blood group and irregular antibody detection in advance so that the matching blood components can be selected for patients in time to ensure the safety and effectiveness of clinical blood transfusion.

CONCLUSION

Compared with the conventional tube method, the EMT-based Collies automatic blood group system for irregular antibody screening can more accurately screen its irregular antibody and has more substantial antibody specificity. In addition, it is more convenient to operate, with more accurate results, less cost consumed, and other characteristics, and is worthy of broader promotion in clinical practice.

ARTICLE HIGHLIGHTS

Research background

Magnetic red cell immunoseparation is a biomedical technique for separating and detecting small numbers of targeted cells. It has the advantages of high sensitivity, high precision, and easy operation and has been widely applied in the field of in vitro diagnostics and therapy.

Research motivation

Conventional separation methods can be time-consuming and involve complicated procedures, while magnetic red cell immunoseparation has the advantages of ease of use, high sensitivity, and precision. Therefore, this technology has been widely applied in in vitro diagnostics and therapy, attracting much attention from researchers and medical professionals.

Research objectives

This study aims to explore the application value of conventional test tubes and erythrocyte-magnetized technology (EMT) in red blood cell alloantibody titration to improve the safety of clinical blood transfusion.

Research methods

Parallel detection of antibody titers for different red blood cell blood groups using in vitro polyene test, tube antiglobulin test (AGT) and EMT screening for irregular antibodies.

Research results

The irregular antibody method for EMT screening could detect all immunoglobulin G and immunoglobulin M irregular antibodies, and the operation time is shorter than manual tube AGT. Furthermore, the EMT screening irregular antibody test was performed to detect its activity on O-type R (D) red blood cells, and the results reflected that it was normal. In addition, compared to the conventional tube method, the EMT screening method for irregular antibodies had lower costs and significantly higher detection efficiency.

Research conclusions

Compared with traditional in vitro methods, EMT screening for irregular antibodies has lower costs and significantly higher detection efficiency.

Research perspectives

The broad application prospects of red blood cell magnetization technology in medicine are evident. Technological advancements will further expand the application scope of this technology. For example, red blood cell magnetization technology can diagnose diseases like early cancer and cardiovascular diseases. In addition, this technology can be used for in vivo and in vitro research on the movement, interaction, and molecular processes of cells and pathogenic microorganisms.

REFERENCES

1 Fox T, Geppert J, Dinnes J, Scandrett K, Bigio J, Sulis G, Hettiarachchi D, Mathangasinghe Y, Weeratunga P, Wickramasinghe D, Bergman H, Buckley BS, Probyn K, Sguassero Y, Davenport C, Cunningham J, Dittrich S, Emperador D, Hooft L, Leeflang MM, McInnes MD, Spijker R, Struyf T, Van den Bruel A, Verbakel JY, Takwoingi Y, Taylor-Phillips S, Deeks JJ; Cochrane COVID-19 Diagnostic Test Accuracy Group. Antibody tests for identification of current and past infection with SARS-CoV-2. Cochrane Database Syst Rev 2022; 11: CD013652 [PMID: 36394900 DOI: 10.1002/14651858.CD013652.pub2]

2 Strand V, Balsa A, Al-Saleh J, Barile-Fabris L, Horiuchi T, Takeuchi T, Lula S, Hawes C, Kola B, Marshall L. Immunogenicity of Biologics in Chronic Inflammatory Diseases: A Systematic Review. BioDrugs 2017; 31: 299-316 [PMID: 28612180 DOI: 10.1007/s40259-017-0231-8]

3 Shastry S, Chenna D, Basavarajegowda A, Das S, Chaudhary RK. Red blood cell alloimmunization among recipients of blood transfusion in India: A systematic review and meta-analysis. Vox Sang 2022; 117: 1057-1069 [PMID: 35608911 DOI: 10.1111/vox.13296]

4 Ngoma AM, Mutombo PB, Ikeda K, Nollet KE, Natukunda B, Ohto H. Red blood cell alloimmunization in transfused patients in sub-Saharan Africa: A systematic review and meta-analysis. Transfus Apher Sci 2016; 54: 296-302 [PMID: 26597314 DOI: 10.1016/j.transci.2015.10.017]

5 Wong K, Lai WK, Jackson DE. HLA Class II regulation of immune response in sickle cell disease patients: Susceptibility to red blood cell alloimmunization (systematic review and meta-analysis). Vox Sang 2022; 117: 1251-1261 [PMID: 36102140 DOI: 10.1111/vox.13351]

6 Alshehri AA, Jackson DE. Non-Invasive Prenatal Fetal Blood Group Genotype and Its Application in the Management of Hemolytic Disease of Fetus and Newborn: Systematic Review and Meta-Analysis. Transfus Med Rev 2021; 35: 85-94 [PMID: 33781630 DOI: 10.1016/j.tmrv.2021.02.001]

7 Boateng LA, Ngoma AM, Bates I, Schonewille H. Red Blood Cell Alloimmunization in Transfused Patients With Sickle Cell Disease in Sub-Saharan Africa; a Systematic Review and Meta-Analysis. Transfus Med Rev 2019; 33: 162-169 [PMID: 31345590 DOI: 10.1016/j.tmrv.2019.06.003]

8 Fortin PM, Hopewell S, Estcourt LJ. Red blood cell transfusion to treat or prevent complications in sickle cell disease: an overview of Cochrane reviews. Cochrane Database Syst Rev 2018; 8: CD012082 [PMID: 30067867 DOI: 10.1002/14651858.CD012082.pub2]

9 Estcourt LJ, Fortin PM, Hopewell S, Trivella M, Doree C, Abboud MR. Interventions for preventing silent cerebral infarcts in people with sickle cell disease. Cochrane Database Syst Rev 2017; 5: CD012389 [PMID: 28500860 DOI: 10.1002/14651858.CD012389.pub2]

10 Fisher SA, Brunskill SJ, Doree C, Gooding S, Chowdhury O, Roberts DJ. Desferrioxamine mesylate for managing transfusional iron overload in people with transfusion-dependent thalassaemia. Cochrane Database Syst Rev 2013: CD004450 [PMID: 23963793 DOI: 10.1002/14651858.CD004450.pub3]

11 Estcourt LJ, Kimber C, Hopewell S, Trivella M, Doree C, Abboud MR. Interventions for preventing silent cerebral infarcts in people with sickle cell disease. Cochrane Database Syst Rev 2020; 4: CD012389 [PMID: 32250453 DOI: 10.1002/14651858.CD012389.pub3]

12 Cirillo P, Gold AK, Nardi AE, Ornelas AC, Nierenberg AA, Camprodon J, Kinrys G. Transcranial magnetic stimulation in anxiety and trauma-related disorders: A systematic review and meta-analysis. Brain Behav 2019; 9: e01284 [PMID: 31066227 DOI: 10.1002/brb3.1284]

13 Lee YF, Hsu TW, Liang CS, Yeh TC, Chen TY, Chen NC, Chu CS. The Efficacy and Safety of Tube Feeding in Advanced Dementia Patients: A Systemic Review and Meta-Analysis Study. J Am Med Dir Assoc 2021; 22: 357-363 [PMID: 32736992 DOI: 10.1016/j.jamda.2020.06.035]

14 Segel GB, Lichtman MA. Direct antiglobulin ("Coombs") test-negative autoimmune hemolytic anemia: a review. Blood Cells Mol Dis 2014; 52: 152-160 [PMID: 24411920 DOI: 10.1016/j.bcmd.2013.12.003]

15 Selvaraj EA, Mózes FE, Jayaswal ANA, Zafarmand MH, Vali Y, Lee JA, Levick CK, Young LAJ, Palaniyappan N, Liu CH, Aithal GP, Romero-Gómez M, Brosnan MJ, Tuthill TA, Anstee QM, Neubauer S, Harrison SA, Bossuyt PM, Pavlides M; LITMUS Investigators. Diagnostic accuracy of elastography and magnetic resonance imaging in patients with NAFLD: A systematic review and meta-analysis. J Hepatol 2021; 75: 770-785 [PMID: 33991635 DOI: 10.1016/j.jhep.2021.04.044]

16 Stussi G, Huggel K, Lutz HU, Schanz U, Rieben R, Seebach JD. Isotype-specific detection of ABO blood group antibodies using a novel flow cytometric method. Br J Haematol 2005; 130: 954-963 [PMID: 16156865 DOI: 10.1111/j.1365-2141.2005.05705.x]

17 Nadarevic T, Colli A, Giljaca V, Fraquelli M, Casazza G, Manzotti C, Štimac D, Miletic D. Magnetic resonance imaging for the diagnosis of hepatocellular carcinoma in adults with chronic liver disease. Cochrane Database Syst Rev 2022; 5: CD014798 [PMID: 35521901 DOI: 10.1002/14651858.CD014798.pub2]

18 McQuilten ZK, Crighton G, Brunskill S, Morison JK, Richter TH, Waters N, Murphy MF, Wood EM. Optimal Dose, Timing and Ratio of Blood Products in Massive Transfusion: Results from a Systematic Review. Transfus Med Rev 2018; 32: 6-15 [PMID: 28803752 DOI: 10.1016/j.tmrv.2017.06.003]

19 Saadah NH, van Hout FMA, Schipperus MR, le Cessie S, Middelburg RA, Wiersum-Osselton JC, van der Bom JG. Comparing transfusion reaction rates for various plasma types: a systematic review and meta-analysis/regression. Transfusion 2017; 57: 2104-2114 [PMID: 28766723 DOI: 10.1111/trf.14245]

20 Dodd JM, Windrim RC, van Kamp IL. Techniques of intrauterine fetal transfusion for women with red-cell isoimmunisation for improving health outcomes. Cochrane Database Syst Rev 2010: CD007096 [PMID: 20556774 DOI: 10.1002/14651858.CD007096.pub2]

21 Lin Y, Stanworth S, Birchall J, Doree C, Hyde C. Recombinant factor VIIa for the prevention and treatment of bleeding in patients without haemophilia. Cochrane Database Syst Rev 2011: CD005011 [PMID: 21328270 DOI: 10.1002/14651858.CD005011.pub3]

22 Pecoraro V, Negro A, Pirotti T, Trenti T. Estimate false-negative RT-PCR rates for SARS-CoV-2. A systematic review and meta-analysis. Eur J Clin Invest 2022; 52: e13706 [PMID: 34741305 DOI: 10.1111/eci.13706]

23 Kohli M, Schiller I, Dendukuri N, Dheda K, Denkinger CM, Schumacher SG, Steingart KR. Xpert(®) MTB/RIF assay for extrapulmonary tuberculosis and rifampicin resistance. Cochrane Database Syst Rev 2018; 8: CD012768 [PMID: 30148542 DOI: 10.1002/14651858.CD012768.pub2]

24 Severe Covid-19 GWAS Group, Ellinghaus D, Degenhardt F, Bujanda L, Buti M, Albillos A, Invernizzi P, Fernández J, Prati D, Baselli G, Asselta R, Grimsrud MM, Milani C, Aziz F, Kässens J, May S, Wendorff M, Wienbrandt L, Uellendahl-Werth F, Zheng T, Yi X, de Pablo R, Chercoles AG, Palom A, Garcia-Fernandez AE, Rodriguez-Frias F, Zanella A, Bandera A, Protti A, Aghemo A, Lleo A, Biondi A, Caballero-Garralda A, Gori A, Tanck A, Carreras Nolla A, Latiano A, Fracanzani AL, Peschuck A, Julià A, Pesenti A, Voza A, Jiménez D, Mateos B, Nafria Jimenez B, Quereda C, Paccapelo C, Gassner C, Angelini C, Cea C, Solier A, Pestaña D, Muñiz-Diaz E, Sandoval E, Paraboschi EM, Navas E, García Sánchez F, Ceriotti F, Martinelli-Boneschi F, Peyvandi F, Blasi F, Téllez L, Blanco-Grau A, Hemmrich-Stanisak G, Grasselli G, Costantino G, Cardamone G, Foti G, Aneli S, Kurihara H, ElAbd H, My I, Galván-Femenia I, Martín J, Erdmann J, Ferrusquía-Acosta J, Garcia-Etxebarria K, Izquierdo-Sanchez L, Bettini LR, Sumoy L, Terranova L, Moreira L, Santoro L, Scudeller L, Mesonero F, Roade L, Rühlemann MC, Schaefer M, Carrabba M, Riveiro-Barciela M, Figuera Basso ME, Valsecchi MG, Hernandez-Tejero M, Acosta-Herrera M, D'Angiò M, Baldini M, Cazzaniga M, Schulzky M, Cecconi M, Wittig M, Ciccarelli M, Rodríguez-Gandía M, Bocciolone M, Miozzo M, Montano N, Braun N, Sacchi N, Martínez N, Özer O, Palmieri O, Faverio P, Preatoni P, Bonfanti P, Omodei P, Tentorio P, Castro P, Rodrigues PM, Blandino Ortiz A, de Cid R, Ferrer R, Gualtierotti R, Nieto R, Goerg S, Badalamenti S, Marsal S, Matullo G, Pelusi S, Juzenas S, Aliberti S, Monzani V, Moreno V, Wesse T, Lenz TL, Pumarola T, Rimoldi V, Bosari S, Albrecht W, Peter W, Romero-Gómez M, D'Amato M, Duga S, Banales JM, Hov JR, Folseraas T, Valenti L, Franke A, Karlsen TH. Genomewide Association Study of Severe Covid-19 with Respiratory Failure. N Engl J Med 2020; 383: 1522-1534 [PMID: 32558485 DOI: 10.1056/NEJMoa2020283]

25 Ilic D, Djulbegovic M, Jung JH, Hwang EC, Zhou Q, Cleves A, Agoritsas T, Dahm P. Prostate cancer screening with prostate-specific antigen (PSA) test: a systematic review and meta-analysis. BMJ 2018; 362: k3519 [PMID: 30185521 DOI: 10.1136/bmj.k3519]

26 Nelson HD, Fu R, Cantor A, Pappas M, Daeges M, Humphrey L. Effectiveness of Breast Cancer Screening: Systematic Review and Meta-analysis to Update the 2009 U.S. Preventive Services Task Force Recommendation. Ann Intern Med 2016; 164: 244-255 [PMID: 26756588 DOI: 10.7326/M15-0969]

27 Macedo ACL, Prestes GDS, Colonetti T, Candido ACR, Uggioni MLR, Gomes AC, Silva FR, Ceretta LB, Grande AJ, da Rosa MI. A systematic review and meta-analysis of the accuracy of SARS-COV-2 IGM and IGG tests in individuals with COVID-19. J Clin Virol 2022; 148: 105121 [PMID: 35245882 DOI: 10.1016/j.jcv.2022.105121]

28 Zhang ZL, Hou YL, Li DT, Li FZ. Diagnostic efficacy of anti-SARS-CoV-2 IgG/IgM test for COVID-19: A meta-analysis. J Med Virol 2021; 93: 366-374 [PMID: 32568413 DOI: 10.1002/jmv.26211]

29 Rio-Aige K, Azagra-Boronat I, Castell M, Selma-Royo M, Collado MC, Rodríguez-Lagunas MJ, Pérez-Cano FJ. The Breast Milk Immunoglobulinome. Nutrients 2021; 13 [PMID: 34073540 DOI: 10.3390/nu13061810]

30 Jarne-Borràs M, Miró-Mur F, Anunciación-Llunell A, Alijotas-Reig J. Antiphospholipid antibodies in women with recurrent embryo implantation failure: A systematic review and meta-analysis. Autoimmun Rev 2022; 21: 103101 [PMID: 35452853 DOI: 10.1016/j.autrev.2022.103101]

Footnotes

Institutional review board statement: The study was approved by Ethics Committee of the Third Hospital of Shanxi Medical University, No. 2018-003.

Institutional animal care and use committee statement: This study was not involved any animal experiments.

Informed consent statement: All study participants or their legal guardian provided informed written consent about personal and medical data collection prior to study enrolment.

Conflict-of-interest statement: All the authors report no relevant conflicts of interest for this article.

Data sharing statement: No additional data are available.

Open-Access: This article is an open-access article that was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution NonCommercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: https://creativecommons.org/Licenses/by-nc/4.0/

Provenance and peer review: Unsolicited article; Externally peer reviewed.

Peer-review model: Single blind

Peer-review started: March 14, 2023

First decision: April 7, 2023

Article in press: May 15, 2023

Specialty type: Biochemistry and molecular biology

Country/Territory of origin: China

Peer-review report’s scientific quality classification

Grade A (Excellent): 0

Grade B (Very good): B

Grade C (Good): C

Grade D (Fair): 0

Grade E (Poor): 0

P-Reviewer: Mesa F, Spain; Yerke LM, United States S-Editor: Fan JR L-Editor: A P-Editor: Cai YX

Figure Legends

Figure 1 Application of red blood cell magnetization technology in irregular antibody screening. A: Automatic blood group analyzer QWALYS 3 is used for irregular antibody screening; B: Erythrocyte blood group antibody identification cell reaction pattern table.

Table 1 Specific distribution of irregular antibodies in 18 blood donors

| Antibody specificity | Number of cases | Percentage of irregular antibodies (%) | IgG classification | Potency |
| --- | --- | --- | --- | --- |
| Anti-D | 7 | 38.91 | IgG | 2-64 |
| Anti-E | 2 | 11.11 | IgG | 16-32 |
| Anti-cE | 1 | 5.56 | IgG | 8 |
| Anti-C | 1 | 5.56 | IgG | 16 |
| Anti-PI | 4 | 22.22 | IgM or IgG + IgM | 8-16 |
| Anti-M | 1 | 5.56 | IgG | 4-32 |
| Anti-N | 1 | 5.56 | IgG | 8.32 |
| Anti-Lea | 1 | 5.56 | IgG | 4-16 |

Anti-D: Anti-D antibody; Anti-E: Anti-E antibody; Anti-cE: Anti-cE compound antigens antibody; Anti-C: Anti-C antibody; Anti-PI: Anti phosphatidylinositol antibody; Anti-M: Anti M antibody; Anti-N: Anti N antibody; Anti-Lea: Anti-Lea antibody; IgG: Immunoglobulin G; IgM: Immunoglobulin M.

Table 2 Results of irregular antibody titers detected in parallel by three detection methods

| Test method | Specificity of irregular antibodies |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | D | E | C | c | e | S | Fyb | Jka | M | N | P1 |
| Tube polybrene test | Potency | 512 | 256 | 128 | 64 | 32 | 16 | 32 | 64 | 64 | 64 | 32 |
|  | Score | 57 | 53 | 46 | 40 | 33 | 30 | 34 | 40 | 38 | 37 | 35 |
| Tube AGT method | Potency | 256 | 64 | 64 | 32 | 16 | 16 | 8 | 32 | 16 | 16 | 32 |
| 32 | Score | 52 | 59 | 40 | 32 | 28 | 29 | 20 | 30 | 28 | 28 | 32 |
| EMT screening irregular antibody method | Potency | 512 | 256 | 128 | 64 | 16 | 16 | 32 | 64 | 64 | 64 | 32 |
|  | Score | 56 | 54 | 64 | 41 | 31 | 29 | 33 | 40 | 38 | 36 | 34 |

D, E, C, c, e, S, Fyb, Jka, M, N, P1 is all IgG antibodies; Fyb and Jka represent IgG type anti Fyb antibodies and IgG type anti Jka antibodies. EMT: Erythrocyte-magnetized technology; AGT: Antiglobulin test.

Table 3 Cost comparison of different methods

| Methods | Cost (CNY/Donor) |
| --- | --- |
| EMT screening irregular antibody method | < 0.5 |
| Tube polybrene test | > 5 |
| Tube AGT method | > 5.5 |

EMT: Erythrocyte-magnetized technology; AGT: Antiglobulin test.

Published by Baishideng Publishing Group Inc

7041 Koll Center Parkway, Suite 160, Pleasanton, CA 94566, USA

Telephone: +1-925-3991568

E-mail: bpgoffice@wjgnet.com

Help Desk: https://www.f6publishing.com/helpdesk

https://www.wjgnet.com

© 2023 Baishideng Publishing Group Inc. All rights reserved.