Name of journal: World Journal of Gastroenterology

ESPS Manuscript NO: 12276

Columns: PROSPECTIVE STUDY

Measurement system that improves the accuracy of polyp size determined at colonoscopy

Leng Q et al. Measurement system for polyp size estimation

Qiang Leng, Hei-Ying Jin

Qiang Leng, Hei-Ying Jin, National Center of Colorectal Surgery, the Third Affiliated Hospital of Nanjing, University of Traditional Chinese Medicine, Jiangsu Integrate Colorectal Oncology Center, Nanjing 210001, Jiangsu Province, China

Author contributions: Leng Q data collection, analysis and drafting of the manuscript; Jin HY design study, collection data, analysis and revision of the manuscript.

Supported by National Nature Science Foundation of China No. 30973837, No.81273944; and the Nanjing Medical Technology Development Project No. NJYX201203

Correspondence to: Hei-Ying Jin, MD, National Center of Colorectal Surgery, the Third Affiliated Hospital of Nanjing, University of Traditional Chinese Medicine, Jiangsu Integrated Colorectal Oncology Center, 1 Jinling Road, Nanjing 210001, Jiangsu Province, China. jinheiying@hotmail.com

Telephone: +86-25-52276200 	Fax: +86-25-52276200

Received: June 29, 2014		Revised: August 5, 2014

Accepted: September 18, 2014

Published online:

Abstract

AIM: To assess the accuracy of polyp size from an endoscopic lesion measurement system (ELMS).

METHODS: Disposable graduated biopsy forceps (DGBF) were used as a “scale-plate” and ELMS were used to compare the accuracy of polyp size assessment among different methods including: visual estimation, disposable graduated biopsy forceps and the ELMS.

RESULTS: There were 192 polyps from 166 cases included in this study. The mean diameter of the post polypectomy measurement was 0.85 ± 0.53 cm (0.2-3.0 cm). The mean diameter by visual estimation was 1.10 ± 0.53cm; the difference was significant compared to the actual size of the polyp (P < 0.001). The mean diameter using the DGBF was0.87±0.54cm; the difference was not significant when compared to the actual size of the polyp (P = 0.134). The mean diameter using the ELMS was 0.85 ± 0.53 cm; the difference was not significant compared to the actual size of the polyp (P = 0.289). The difference between the measurements from the ELMS and DGBF was not significant (P = 0.147).

CONCLUSION:  The accuracy of estimation of polyp size diameter at colonoscopy that relied on endoscopists’ experience was low. The use of the endoscopic graduated biopsy forceps and the endoscopic lesion measurement system increased the accuracy of estimation of polyp size.

© 2014 Baishideng Publishing Group Inc. All rights reserved.

Key words: Endoscopic lesion measurement system; Polyp size measurement; Colonoscopy; Accuracy

Core tip: In this study, we established an endoscopic lesion measurement system (ELMS) to measure the endoscopic polyp size. We enrolled 192 polyps from 162 patients and found that the accuracy of visual estimation for poly size is low. If the disposable graduated biopsy forceps (DGBF) and ELMS were used, the accuracy of estimation of colonic polyps could be distinctly increased. Especial for polyps over 1cm, the accuracy of estimation by ELMS tend to be higher than that by clinician’s estimation and DGBF.

Leng Q, Jin HY. Measurement system that improves the accuracy of polyp size determined at colonoscopy. World J Gastroenterol 2014; In press

INTRODUCTION

Tumor size determined at colonoscopy is associated with the treatment plan developed for the tumor. However, currently there are no standard criteria used for measuring the size of a tumor identified at colonoscopy. Many clinicians evaluate the size of a tumor based on their experience. Sometimes, this type of assessment has been inconsistent with the actual size of the tumor[1]. Morales et al[2] used open biopsy forceps to evaluate the diameter of colon polyps at colonoscopy, and three quarters of the results were inconsistent with the actual results.  Eichenseer et al[3] compared the estimated polyp size, determined at endoscopy, of 10 to 25-mm polyps, with the histopathology of post fixation polyp measurements, from 15 different endoscopists; the results showed that the mean size variation of polyp size estimation, determined at endoscopy, and the size estimation at histopathology, was 73.6% (range of mean size variation, 13%–127%). About 62.6% (0-91%) of the polyps had their size estimated incorrectly. The diameter of many polyps was over estimated; this might lead to inappropriate advice with regard to the treatment plan. Ping Liu et al[4] developed disposable graduated biopsy forceps that could increase the accuracy of polyp size estimation. However, there continues to be variation, especially for visual estimation of polyps larger than 2 cm. To estimate the size of polyps at colonoscopy more accurately, we designed an endoscopic lesion measurement system (ELMS) using disposable graduated biopsy forceps (DGBF) as a “scale-plate” and assessed its accuracy in the clinic setting.

MATERIALS AND METHODS

Development and use of the ELMS

DGBF were used as a “scale-plate” to design the ELMS. When a polyp was observed at colonoscopy, disposable graduated biopsy forceps parallel to the lesion were pressed onto it and a picture was taken. The ELMS could view the image. First, the length was defined using an endwise graduation scale; then, the endwise diameter of a lesion was measured with the given length as a scale plate. Second, the diameter of the biopsy forceps was used as a scale plate to measure the size of a polyp. According to the results of the up arc of the circumference divided by the length of the arc, if the lesion occupied a quarter of the enteric cavity, then the actual diameter of the polyp was considered to be 1.1 times the measurement. When the lesion occupied one third of the enteric cavity, the size of the polyp was considered to be 1.2 times that of the measurement; the size of the polyp was estimated to be 1.6 times that of the measurement when it occupied half of the enteric cavity (Figure 1).

The ethics committee of the third affiliated hospital of Nanjing University of TCM approved this study.

Study subjects

Patients with colon polyps were enrolled from the National Center of Colorectal Surgery, Nanjing University of Traditional Chinese Medicine. First, when the clinician observed a polyp or tumor, the width of the largest diameter of the tumor was determined; then, a graduated biopsy forceps was inserted and used to measure the width of the largest diameter of the tumor from the vertical view and a picture was obtained according to the ELMS protocol; then, the ELMS was used to measure the size of the lesion; finally, the “gold standard” for the width of the largest diameter of the tumor was accurately measured using Vernier calipers after the tumor was excised by surgery or endoscopy. A correct measurement was considered when the variation between the evaluation size and actual size, as measured by Vernier calipers, was less than 10%. The estimated accuracy rate was equal to the number of polyps accurately evaluated divided by the total number of polyps observed. Five endoscopists with experience of over 2000 colonoscopy cases were involved in this study.

Statistical analyses

Data was analyzed using SPSS 17.0 software for Windows. The paired Student’s t test was used to compare the ratio of the estimated size by the endoscopists, the DGBF and the ELMS, to the actual size measured by Vernier calipers. Analysis of variance was used to compare the differences among the four groups. A P of < 0.05 was considered to be statistically significant.

RESULTS

Basic clinical characteristics of the polyps

From Aug 2013 to Mar 2014, 166 patients (110 male, 56 female, mean age 58.7 ± 11.2 years, range 28 to 84years) with 192 colorectal polyps were included in this study. Among all polyps, there were 69 rectal polyps, 64 sigmoid colonic polyps, 16 descending colon polyps, 15 transverse colon polyps, and 28 cecum or ascending colon polyps. Among all patients, five patients underwent laparoscopic colectomy, and the rest had endoscopic mucosal resection. The pathological results of the all resected lesions included 117 tubular adenomas, 27 villous adenomas (of them, one was mucosal cancer), 46 tubulovillous adenomas, and 2 neuroendocrine tumor (carcinoid).

Comparison of the largest diameter measured by Vernier calipers, the clinicians’ evaluation, the graduated biopsy forceps and the endoscopic lesion measurement system

The mean largest diameter of the resected polyps measured by a vernier caliper was 0.85±0.53 cm (range, 0.2-3.0 cm), and of them, 134 were < 1 cm, 43 were between 1 and 2 cm, and 15 were > 2 cm. The mean largest diameter of the resected polyps by the clinicians’ evaluation was 1.10 ± 0.53 cm and was statistically different from that of the mean largest diameter of the resected polyps measured by a vernier caliper (P < 0.001). The mean largest diameter of the resected polyps measured by the graduated biopsy forceps was 0.87 ± 0.54 cm that was not statistically different from that of the mean largest diameter of the resected polyps measured by a vernier caliper (P = 0.134). The mean largest diameter of the resected polyps measured by the endoscopic lesion measurement system was 0.85 ± 0.53 cm that was not statistically different from that of the mean largest diameter of the resected polyps measured by a vernier caliper (P = 0.289) (Figure 2.) The mean largest diameter of the resected polyps measured by the graduated biopsy forceps and by the endoscopic lesion measurement system was statistically different from that of the mean largest diameter of the resected polyps measured by clinicians’ evaluation (P < 0.001). While the mean largest diameter of the resected polyps measured by the graduated biopsy forceps had no statistical difference compared to the results measured by the endoscopic lesion measurement system (P = 0.147) (Table1).

Comparison of the ratio of the estimated size by endoscopists, graduated biopsy forceps,  the endoscopic lesion measurement system and the actual size measured by Vernier calipers

The ratio of the estimated size by the endoscopists to the actual size measured by a vernier caliper was 1.30±0.33 while the ratio of the size measured by the graduated biopsy forceps to the actual size measured by a vernier caliper was1.11 ± 0.37; the ratio of estimated size by the endoscopic lesion measurement system to the actual size measured by a vernier caliper was 1.04 ± 0.13 (Figure 3, Table 2).

Influence of tumor size on tumor size measurement accuracy

For 134 polyps less than 1 cm, only 20.6% (28/134) could be accurately estimated by the clinicians based on their experience while 61.2% (82/134) could be accurately measured by the graduated biopsy forceps and 75.4% (101/134）could be accurately measured by the endoscopic lesion measurement system, these were statistically different (P < 0.001). The results between DGBL and ELMS was also significantly statistical difference (P = 0.001）.

For 43 polyps between 1 cm and 2 cm, 14.0% (6/43) could be accurately estimated by the clinicians based on their experience while 48.8 % (21/43) could be accurately measured by the graduated biopsy forceps and 74.4% (32/43) could be accurately measured by the endoscopic lesion measurement system; these were statistically different (P < 0.001). The results between DGBL and ELMS had no significantly statistical difference (P = 0.003).

For 15 polyps over 2 cm, none could be accurately estimated by the clinicians based on their experience while 40% (6/15) could be accurately measured by the graduated biopsy forceps and 80.0% (12/15) could be accurately measured by the endoscopic lesion measurement system; these were statistically different (P = 0.008 and P = 0.046, respectively). But the results between DGBL and ELMS had no significantly statistical difference (P = 0.212) (Table3).

DISCUSSION

The size of a polyp determined at colonoscopy is crucial for selecting a treatment plan and assessing the type of polyp present[5-7]. Currently, the size of a polyp is evaluated according to the clinician’s experience. However, questions have been raised about the difference between endoscopists’ evaluation and the actual size of a polyp. Eichenseer et al[3] reported that the mean size variation between the size of a polyp estimated at endoscopy and the size determined by histopathology was 73.6% (range of mean size variation, 13%-127%). About 62.6% (0-91%) of polyps are incorrectly estimated, and the sizes of most are over estimated. As a result, inappropriate advice might be given for follow up. Gopalswamy et al[8] compared the accuracy of a linear probe, visual estimation and forceps for estimating polyp size during colonoscopy and found that the measurement of polyp size, using a linear probe, was most consistent  with  the actual polyp size. The open biopsy forceps method was found to be the least accurate method. However, the linear probe requires special software for estimation, which adds to examination time and cost. Liu et al[4] used a disposable graduated biopsy forceps to estimate the size of polyps determined at colonoscopy. They found when one only relies on the experience of the endoscopist the accuracy of size estimation is very low. However, using a disposable graduated biopsy forceps as a “scale-plate” can increase the accuracy of estimation of polyp size.  However, when using graduated biopsy forceps, to measure the diameter of a polyp, due to visual deviation, the results are inaccurate, especially for the transverse diameter of the side of the developing tumor. In order to reduce visual deviation, we developed an endoscopic lesion measurement system, which uses graduated biopsy forceps as a “scale-plate” to measure the diameter of a lesion, and thereby increase the measurement accuracy.

In this study, the average largest diameter of the resected polyps measured by vernier caliper was 0.85 ± 0.69 cm, and the average largest diameter of the polyps estimated by endoscopists’ experience was 1.10 ± 0.53 cm, there was statistical difference. And this proved that the accuracy of estimation of polyps’ size by endoscopists’ experience was low. The mean diameter of the polyps measured by endoscopic graduated biopsy forceps was 0.87 ± 0.54 cm, which had no significant difference compared to the actual measurement. And this indicated that the use of endoscopic graduated biopsy forceps could increase the accuracy of estimation of polyps, which was consistent with the report of Liu et al[4]. The mean diameter of the polyps measured by endoscopic lesion measurement system was 0.85 ± 0.53cm, without significant difference with the actual measurement, which indicated that the endoscopic lesion measurement system could estimate the polyps’ size correctly. The coincident measure results of endoscopic graduated biopsy forceps and endoscopic lesion measurement system suggested that both endoscopic graduated biopsy forceps and endoscopic lesion measurement system could improve the accuracy of endoscopic estimation of polyps.

When evaluating the accuracy for different polyps’ size, of the 134 polyps, which less than 1 cm, only 20.6% was correctly estimated by endoscopists’ experience; while by using endoscopic graduated biopsy forceps, 61.2% was correctly estimated, which conformed to the previous study[4]. And the accuracy rate of the endoscopic lesion measurement system was 75.4%. The accuracy rates of endoscopic graduated biopsy forceps was lower than that by endoscopic lesion measurement system.But there are still 25% error rate in less than 1cm group. The reasons maybe, first, the less visual deviation of the polyps less than 1 cm increased the accuracy of biopsy forceps. Second, many polyps which less than 1 cm were protrude type, when taking pictures,  due to the potential deviation of visual angle, when the measure deviation was more than 10%, the measurement would be judged to be inaccuracy. But for the deviation less than 1 mm would not affect the choice of the treatment. For the 43 polyps which were 1 to 2 cm, rely on the experience of endoscopists’, only 14.0% was estimated correctly. However the accuracy rates were increased to 48.8 % and 74.4% by using endoscopic graduated biopsy forceps and endoscopic lesion measurement system, respectively. The accuracy rates of endoscopic graduated biopsy forceps and endoscopic lesion measurement system had significant difference, the accuracy of endoscopic lesion measurement system for lesions was trend to be increased.  Especially for the 15 polyps which were larger than 2 cm, none of the polyps was accurately estimated by endosocopists’ visual estimation , while the accuracy rates of endoscopic graduated biopsy forceps and endoscopic lesion measurement system were 40% and 80%, which suggested that the polyps larger than 2 cm would increase visual deviation, then the accuracy of endoscopic lesion measurement system was trend to increase. But in this study, the number of the polyps that larger than 2 cm was not enough to reflect statistical difference. In future, more polyps bigger than 2 cm will be included in this study. Besides, this study only estimate the biggest diameter of the lesions, and did not compare the transverse or endwise diameter of the lesions. But all these will be investigated in future studies.

In conclusion, the accuracy of estimation of polyp size diameter at  colonoscopy that relied on endoscopists’ experience was low. The use of the endoscopic graduated biopsy forceps and the endoscopic lesion measurement system increased the accuracy of estimation of polyp size.

ACKNOWLEDGEMENTS

Many thanks to: Mrs Ping Liu, Mrs Xiu Zhang, Mrs Huiping Lin for the acquisition of data and Dr. Hang Yao, Dr. Kunlan Wu, Dr. Jinhao Zhang for the analysis and interpretation of data.

COMMENTS

Background

Tumor size via colonoscopy is related to the nature and treatment of the tumor. However there are still no standardized criteria for measuring the size of the tumor via colonoscopy, and many clinicians evaluated the size based on their experience. Sometimes, this kind of judgment was inconsistent with the actual size of the tumor. We invented a disposable graduated biopsy forceps which could increase the accuracy of estimating the size of the polyps, while there is still visual variations especially for the polyps which larger than 2 cm. To estimate the size of polyps via colonoscopy more accuracy, we designed an endoscopic lesion measurement system (ELMS) by using disposable graduated biopsy forceps (DGBF) as “scale-plate” and investigated its evaluation in clinic.

Research frontiers

In order to establish an ELMS and to study of accuracy of endoscopic polyp size measurement. Using the DGBF as “scale-plate” and an ELMS was designed. The accuracy of polyps size was compared among different methods by visual estimation, disposable graduated biopsy forceps and ELMS.

Innovations and breakthroughs

In this study, we established an ELMS to measure the endoscopic polyp size. We found that the accuracy of visual estimation for poly size is low. The accuracy of visual estimation for polyp size is low. If the DGBF and ELMS were used, the accuracy of estimation of colonic polyps could be distinctly increased. For polyps less than 1cm, the accuracy of estimation by DGBF tend to be higher and for polyps over 1cm, the accuracy of estimation by ELMS tend to be higher.

Applications

This system is very easy to learn and provides objective estimation of polyps size under endoscopy. It can be used in clinical practice and do not increase the medical cost.

Peer review

An endoscopic lesion measurement system can improve the accuracy of endoscopic polyp size measurement. The investigators have compared a new technique versus standard to grade endoscopic polyp size. They find significant differences between the old and new methods. This is potentially a compelling study. However, there must be some improvements.

REFERENCES

1 Wang HM, Huang CM, Zheng CH, Li P, Xie JW, Wang JB, Lin JX, Lu J. Tumor size as a prognostic factor in patients with advanced gastric cancer in the lower third of the stomach. World J Gastroenterol 2012; 18: 5470-5475 [PMID: 23082065 DOI: 10.3748/wjg.v18.i38.5470]

2 Morales TG, Sampliner RE, Garewal HS, Fennerty MB, Aickin M. The difference in colon polyp size before and after removal. Gastrointest Endosc 1996; 43: 25-28 [PMID: 8903813 DOI: 10.1016/S0016-5107(96)70255-9]

3 Eichenseer PJ, Dhanekula R, Jakate S, Mobarhan S, Melson JE. Endoscopic mis-sizing of polyps changes colorectal cancer surveillance recommendations. Dis Colon Rectum 2013; 56: 315-321 [PMID: 23392145 DOI: 10.1097/DCR.0b013e31826dd138]

4 Liu P, Zhang X, Lin HP, Jin HJ, Leng Q, Zhang JH, Zhang Y, Yao H, Wu KL. [Study on accuracy of endoscopic polyp size measurement by disposable graduated biopsy forceps]. Zhonghua Wei Chang Wai Ke Za Zhi 2013; 16: 1142-1145 [PMID: 24369393]

5 Tseng MY, Lin JC, Huang TY, Shih YL, Chu HC, Chang WK, Hsieh TY, Chen PJ. Endoscopic submucosal dissection for early colorectal neoplasms: clinical experience in a tertiary medical center in Taiwan. Gastroenterol Res Pract 2013; 2013: 891565

6 Toyonaga T, Man-i M, Chinzei R, Takada N, Iwata Y, Morita Y, Sanuki T, Yoshida M, Fujita T, Kutsumi H, Hayakumo T, Inokuchi H, Azuma T. Endoscopic treatment for early stage colorectal tumors: the comparison between EMR with small incision, simplified ESD, and ESD using the standard flush knife and the ball tipped flush knife. Acta Chir Iugosl 2010; 57: 41-46 [PMID: 21066982]

7 Horiuchi Y, Chino A, Matsuo Y, Kishihara T, Uragami N, Fujimoto Y, Ueno M, Tamegai Y, Hoshino E, Igarashi M. Diagnosis of laterally spreading tumors (LST) in the rectum and selection of treatment: characteristics of each of the subclassifications of LST in the rectum. Dig Endosc 2013; 25: 608-614 [PMID: 23369130 DOI: 10.1111/den.12040]

8 Gopalswamy N, Shenoy VN, Choudhry U, Markert RJ, Peace N, Bhutani MS, Barde CJ. Is in vivo measurement of size of polyps during colonoscopy accurate? Gastrointest Endosc 1997; 46: 497-502 [PMID: 9434215 DOI: 10.1016/S0016-5107(97)70003-8]

P-Reviewer: Amornyotin S, Ciaccio EJ,  Mentes O S-Editor: Qi Y

L-Editor:   E-Editor:

Figure 1 Interface of the endoscopic lesion measurement system. A: Sketch of the endoscopic lesion measurement system (ELMS); B: Entering interface.

Figure 2 Comparison of the largest diameter of lesions between the actual measurement and the other three estimation methods.

Figure 3 Comparison of the accuracy for the largest diameter according to: visual estimation, disposable graduated biopsy forceps and endoscopic lesion measurement system compared to the actual diameter.

Table 1 P value of comparison of the largest diameter measured by different methods (n = 192)

| Method for measurement | Mean(cm) | VC | CE | DGBF |
| --- | --- | --- | --- | --- |
| VC | 0.85 ± 0.69 | - | - | - |
| CE | 1.10 ± 0.53 | < 0.001 | - | - |
| DGBF | 0.87 ± 0.54 | 0.134 | < 0.001 | - |
| ELMS | 0.85 ± 0.53 | 0.289 | < 0.001 | 0.027 |

VC: Vernier caliper; DGBF: Disposable graduated biopsy forceps; ELMS: Endoscopic lesion measurement system; CE: Clinicians’ evaluation.

Table 2 P value of comparison of the ratio of difference estimation to actual measurement (n = 192)

| Ratio | Mean(cm) | DGBF to VC | ELMS to VC |
| --- | --- | --- | --- |
| CE to VC | 1.30 ± 0.33 | < 0.001 | < 0.001 |
| DGBF to VC | 1.11 ± 0.37 | - | 0.021 |
| ELMS to VC | 1.04 ± 0.13 | - | - |

VC: Vernier caliper; DGBF: Disposable graduated biopsy forceps; ELMS: Endoscopic lesion measurement system; CE: Clinicians’ evaluation.

Table 3 Influence of different tumor sizes on the tumor size accurate estimated rate (n = 192)

| Polyps size | Rate of accurate estimation | CE | DGBF |
| --- | --- | --- | --- |
| < 1cm |  |  |  |
| CE | 20.6% (28/134) | - | - |
| DGBF | 61.2% (82/134) | < 0.001 | - |
| ELMS | 75.4%（101/134） | <0.001 | 0.001 |
| 1-2cm |  |  |  |
| CE | 14.0% (6/43) |  |  |
| DGBF | 48.8 % (21/43) | < 0.001 |  |
| ELMS | 74.4%（32/43） | < 0.001 | 0.003 |
| > 2cm |  |  |  |
| CE | 0 |  |  |
| DGBF | 40% (6/15) | 0.008 |  |
| ELMS | 80.0%(12/15） | 0.046 | 0.212 |

VC: Vernier caliper; DGBF: Disposable graduated biopsy forceps; ELMS: Endoscopic lesion measurement system: CE: Clinicians’ evaluation.