Name of journal: World Journal of Gastroenterology

ESPS Manuscript NO: 16072

Columns: EDITORIAL

Wireless endoscopy in 2020: Still a capsule and what form?

Koulaouzidis A et al. Wireless endoscopy

Anastasios Koulaouzidis, Dimitris K Iakovidis, Alexandros Karargyris, Emanuele Rondonotti

Anastasios Koulaouzidis, Endoscopy Unit, The Royal Infirmary of Edinburgh, EH16 4SA Edinburgh, Scotland, United Kingdom

Dimitris K Iakovidis, Department of Computer Engineering, Technological Educational Institute of Central Greece, 35100 Lamia, Greece

Alexandros Karargyris, The National Institute of Health, Bethesda, MD 20810, United States

Emanuele Rondonotti, Gastroenterology Department, Ospedale Valduce, 22100 Como, Italy

Author contributions: All authors contributed to this manuscript.

Conflict-of-interest: Koulaouzidis A has received research support from Given Imaging and SynMed UK, lecture honoraria from Dr Falk Pharma UK, and travel support from Abbott, Dr Falk Pharma UK, Almirall, and MSD. The rest of the authors declare no competing interests.

Open-Access: This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/licenses/by-nc/4.0/

Corresponding to: Anastasios Koulaouzidis, MD, FEBGH, FRSPH, FRCPE, DM, The Royal Infirmary of Edinburgh, Endoscopy Unit, 51 Little France Crescent, EH16 4SA Edinburgh, Scotland, United Kingdom. akoulaouzidis@hotmail.com

Telephone: +44-131-2421126

Fax: +44-131-2421618

Received: December 24, 2014

Peer-review started: December 31, 2014

First decision: January 8, 2015

Revised: January 26, 2015

Accepted: March 18, 2015

Article in press:

Published online:

Abstract

Currently, the major problem of all existing commercial capsule devices is the lack of control of movement. In the future, with an interface application, the clinician will be able to stop and direct the device into points of interest for detailed inspection/diagnosis, and therapy delivery. This editorial presents current commercially-available new designs, European projects and delivery capsule and gives an overview of the progress required and progress that will be achieved -according to the opinion of the authors- in the next 5 year leading to 2020.

Key words: Capsule endoscopy; Hardware; Development; Future

© The Author(s) 2015. Published by Baishideng Publishing Group Inc. All rights reserved.

Core tip: Currently, the major problem of all existing commercial capsule devices is the lack of control of movement. In the future, with an interface application, the clinician will be able to stop and direct the device into points of interest for detailed inspection/diagnosis, and therapy delivery. This editorial presents current commercially-available new designs, European projects and delivery capsule and gives an overview of the progress required and progress that will be achieved -according to the opinion of the authors- in the next 5 year leading to 2020.

Koulaouzidis A, Iakovidis DK, Karargyris A, Rondonotti E. Wireless endoscopy in 2020: Still a capsule and what form? World J Gastroenterol 2015; In press

CAPSULE ENDOSCOPY

The swallowable-capsule concept first appeared in 1957 in Mackay and Jacobson’s groundbreaking paper on radiofrequency (RF) transmission of temperature and pressure from within the human body[1,2]. Nowadays, capsule endoscopy (CE) is the prime mode of non-invasive and discomfort-free endoscopic exploration of the small-bowel[3,4]. It became available to clinicians at the dawn of the millennium[5]. The fact that its realization was the brainchild of Garviel Iddan D.Sc., an Israeli electro-optical engineer - initially working at the RAFAEL Armament Development Authority on guided missile technology, is to an extent multi-semantic. Parallelly, Professor Paul Swain, a British gastroenterologist, led a research group in London and published the first conceptual studies on a wireless capsule in 1994[6,7]. Since then, a great number of clinical and scientific papers have studied the clinical use of this fascinating new technology[8,9] in the investigation of small-bowel diseases. Nowadays, capsule platforms are also available for the non-invasive exploration of the oesophagus and the colon[10,11]. The full spectrum of the commercial capsule models is available elsewhere and remains beyond the scope of this editorial[8,12].

Currently, the major problem of all existing commercial capsule devices is the lack of control of movement; the latter is achieved by the peristaltic contractions of the small bowel. Therefore, as bowel peristalsis is a complex event of five contractile patterns, i.e., peristaltic waves, stationary contractions, clusters of contractions (Phase III), giant contractions and anti-peristalsis waves[13], its impact on capsule locomotion (speed, position and orientation) is unpredictable[14]. Consequently, the first significant step for a forward leap in wireless technology is achieving sustainable, active movement and control of the device’s locomotion[4,15]. However, when extrapolating from the advancements made in conventional (flexible) endoscopy since its introduction -more than 60 years ago- in regular clinical practice, the prospect for wireless devices should be considered as anything but optimistic. Manual push, perhaps the most primitive actuation method, is still used as the mainstream advancement option of conventional/flexible endoscopes[15]. Hence, embarrassment, discomfort and/or frank pain during the advancement phase of a flexible endoscope is the major drawback of this technique[15,16].

Nevertheless, the absence of wire in capsule-like endoscopy systems requires drastic measures and has pushed research forward in a more radical manner. Sliker and Ciuti[15] argue that in the next few years, an integrated, magnetically actuated, automated locomotion system will be available for regular clinical use. Furthermore, possibly with an interface application, the clinician will be able to stop and direct the device into points of interest for detailed inspection/diagnosis, and therapy delivery. Hence, passive locomotion will be replaced by externally (magnetically) controlled actuation or a combination of a miniaturization, new battery type or complete battery elimination[17]. Especially the battery-less and/or miniaturization version will eliminate any concerns about the chemicals in the batteries, can produce potential hazards and contaminate the environment when the capsules are improperly disposed of[18,19].

One of the major problems of the capsule platform for digestive endoscopy, especially when there is no locomotion control is the adoption of the single-pole lens model; however, integrating multiple cameras has still problems with size and power consumption[20]. Therefore, some ideas around this have led to few experimental developmental platforms that, at this point, should receive particular mention:

NORIKA AND SAYAKA

RF SYSTEM lab. (RF Co., Ltd., Nagano, Japan) was established in 1993 in Nagano. The deep-sea submarine-like capsule, named NORIKA 3 (Figure 1A), was developed in 1997 when RF Co. joined the Japanese Experimental Module Project[21]; the latter was a venture operated by the National Space Development Agency of Japan to monitor the growth of plants in a manned spacecraft[21]. SAYAKA, the first "battery-less capsule", superseding the basic technology of NORIKA 3, was announced to the international market on December 2001; however, it remains still under development (Figure 1B)[22]. The dimension of a capsule camera is 9 mm × 23 mm. A 0.6 mm color lens and a 410000-pixel charge-coupled device (CCD) camera can obtain up to 30 frames per sec (fps), as long as the patient is wearing a vest that transmits microwaves to the capsule[23]. Two tanks with valves are positioned in the center and a there is a capacitor to store electric power and a microwave video signal transmitter[22,23]. Around the camera lens, four white LEDs and magnetic coils for focus adjustment are placed[23]. More importantly, 40% of the volume is free space, which can be used for surgical purposes such as medication spray, laser treatment, pH sensing and more[24]. Mosaicing technology was developed for use with the experimental model Norika 3[25]. Mosaicing is a process by which the final image on the computer display is made by combining multiple images taken from various angles[23,25]. However, no human studies have been conducted with this capsule, which has been described on the web for almost a decade now[12].

SELF-STABILIZING CAPSULE ENDOSCOPE

An innovative method of imaging of the GI tract by using a self-stabilizing capsule endoscope (SsCE) was first proposed in 2006 (Figure 2A)[26]. SsCE allows the use of a single-pole imager devise for visualizing the wider diameter segments of the digestive tract such as the colon without tumbling and with the ability to passively distend colon walls[27,28]. The capsule coating is designed to dissolve in the colon, which exposes a semipermeable, expandable container attached to the back of the capsule endoscope, while simultaneously turning on the camera (Figure 2B)[27,28]. In a recent experimental study with live canine models, 4 mongrel dogs underwent laparotomy and the implantation of 5-8 suture markers to approximate colon lesions. Each dog had both single-dome CE and ScCE in random order. The average percentages of the marker detection rate for unmodified capsule endoscopy, self-stabilizing capsule endoscopy, and colonoscopy, respectively, were 31.1%, 86%, and 100% (P < 0.01), with both self-stabilizing capsule endoscopy and colonoscopy performing significantly better than the unmodified capsule endoscopy[28].

ODOCAPSULE

Odocapsule, a capsule device to achieve accurate localization of small-bowel lesions and endoscopic video stabilization in capsule endoscopy was proposed in 2010[29] (Figure 3A). It is equipped with 3 miniature legs, each carrying a wheel. These legs are extendable and retractable thanks to a micro-motor and three custom-made torsion springs[30,31]. The wheels are specifically designed to function as micro-odometers by registering each rotation they perform. As the Odocapsule traverses through the small bowel, the wheels turn and their rotations are translated to distance covered by the device from the point of duodenal entry to each area/point of interest. As the legs are expandable in a tripod formation, they allow the device to stabilize itself without obstructing its locomotion and thus offering smooth video capture without missing any pathologies (Figure 3B).

This editorial does not aim to be exhaustive or indeed cover areas presented in details in previous excellent reviews[2,14,15,17,22]. Its main aim is to present the authors’ views on interesting current projects in CE and the vision for potential convergence/integration of various diagnostic/therapeutic platforms.

DRUG DELIVERY CAPSULE PLATFORMS

Coagulation capsule

Various chitosan- and mineral-based haemostatic granules or powders are used for the control of compressible, external haemorrhage in combat casualties and are incorporated in first-aid kits used by the military[32]. Hemospray (TC-325) is a novel hemostatic agent (COOK Medical, Winston-Salem, NC), CE-marked for use in the endoscopic treatment of high-risk, non-variceal upper gastrointestinal (GI) bleeding in Europe and Canada. It is a proprietary inorganic mineral, absorbent powder that has no known allergens and rapidly concentrates clotting factors at the target site, thereby forming an adherent coagulum[32,33]. Currently, its hand-held container - consisting a pressurized CO2 canister, a through-the-scope delivery catheter, and a reservoir for the 21-g powder cartridge- is operated push button in 1- to 2-s bursts[32]. A prototype coagulation capsule (CoCap) has been built and tested which employs an exothermic chemical reaction to generate heat using the interaction of calcium oxide and water[34]. Professor Swain pointed out that it is possible for other thermal/therapeutic applications to be added in the future[35]. Recently, the feasibility of a novel method of controlled colonic insufflation via an untethered capsule in vivo was demonstrated (Figure 4)[36]. It is truly almost self-intuitive to combine the capabilities of various capsule designs (powder carrier, insufflator and imaging/controllability) in one CoCap for use in cases of obscure GI bleeding. These authors believe that the aforementioned ideas will be transformative on how the deal with OGIB in 2020 and allow CE to claim, once more, the role of “disruptive” technology in the medical field. Indeed, there are few capsule products nowadays that can serve as medication carriers to sites in the small-bowel.

Enterion

The Enterion capsule was developed by Phaeton Research, Nottingham, United Kingdom (in association with the Nottingham University spin-off Pharmaceutical Profiles) for targeted delivery of a wide range of different drug formulations into any region of the gut[37,38]. It is a 32-mm long capsule device that contains a drug reservoir with a volume capacity of approximately 1 mL (Figure 5). Enterion has the capability to deliver, through a 9-mm opening, any type of drug formulation (including dry powder, semisolid, suspension or solution formulations) when the spring is released, forcing a piston to move along the shaft of the capsule, thus, emptying the reservoir to specific locations of the GI tract[2,39]. A radioactive marker is placed inside a separate sealed tracer port to allow real-time visualization of the capsule location using gamma scintigraphy[37]. The movement of the piston also operates a switch, which transmits a weak radio signal at a precise frequency. Detection of this signal externally confirms that the capsule has opened successfully[37,38]. The piston motion is stopped near the end of the capsule, which maintains a seal and prevents contact of the internal electronic components with the GI fluids.

InteliSite® Capsule

The InteliSite capsule (Innovative Devices, LLC, Raleigh, North Carolina, USA) is a 10 mm x 35 mm, radiofrequency activated, non-disintegrating drug delivery device (Figure 6)[37]. The use of radiomarkers (gamma scintigraphy) allows determination of the capsule location[2,37]. The capsule packaging consists of inner and outer sleeves, which have portholes; when the holes align under a magnetic field straighten shape-memory alloy wires, there is release of the drugs[2,37]. To date, limited information is available in the literature on the use of the capsule[2,37].

INTELLICAP SYSTEM

Philips Research in Eindhoven, Netherlands, has developed a prototype called intelliCap[2,37,38]. IntelliCap has a drug reservoir, a drug pump, a pH sensor, a wireless transceiver, and a microcontroller. The combination allows us to measure the local pH, report the data in real time, and have full control over drug delivery, based on a combination of data such as time, pH, or operator intervention[2,37,38].

MICROROBOT FOR TARGETED DRUG DELIVERY

With a stabilizing/holding mechanism that resembles daVinci-like drawings, the micro-robot platform developed by Woods and Constandinou[40,41], allows targeted drug delivery in the next-generation wireless CE. The micro-positioning mechanism allows a needle to be positioned within a 22.5° segment of a cylindrical capsule and be extendible by up to 1.5 mm outside the capsule body (Figure 7A). This is controlled by a single micro-motor and occupying a total volume of just 200 mm3. The holding mechanism is depicted in Figure 7B. Perhaps adrenaline injection directly into a bleeding small bowel angioectasia is not a bream after all.

Another micro-robot that could be used for drug or other therapy delivery employs actuated legs with compliant feet lined with micro-pillar adhesives - inspired by gecko and beetle foot hairs - to be pressed into the intestine wall to anchor the device at a fixed location[41]. The investigators envisage the use of this capsule as second-line procedure, “after images of the conventional capsule examination are analyzed by a clinician to map areas of interest in the intestine and identify any problematic regions that may require a closer look”[41].

MAGNETICALLY ACTUATED SOFT CAPSULE ENDOSCOPE

A magnetically actuated soft capsule endoscope (MASCE) has been developed by Yim and Sitti (Yim S, Sitti M. 3-D Localization Method for a Magnetically Actuated Soft Capsule Endoscope and Its Applications. Robotics, IEEE Transactions 2013 Oct; 1139-1151. Yim S, Sehyuk, and Sitti M. Design and rolling locomotion of a magnetically actuated soft capsule endoscope. Robotics, IEEE Transactions 2012; 28: 183-194), for diagnostic and therapeutic medical applications in the stomach. It has a characteristic ability to deform both passively and actively. Passive deformation is possible because its cover is made from elastomer-based compliant structures. This prevents any possible tissue damage during swallowing and/or the peristaltic movement. Active deformation is possible on the axial direction by external magnetic actuation. This provides an extra degree of freedom that enables various advanced functions such as axial position control, drug releasing, drug injection, or biopsy. For example in the developed prototype the axial magnetic attraction compresses a drug chamber between two internal magnets so as to release a drug through holes at a critical pressure.

Furthermore, the magnetic actuation enables the soft capsule endoscope to roll on the surface of the stomach. Advantages of the rolling ability include smooth and continuous locomotion and stable steering because the capsule is always anchored to the tissue surface while moving, it allows improved tissue surface diagnosis and tissue targeting capability for therapeutic procedures, and it provides a more complete view of the 3-D stomach tissue wall; thus reducing the possibility of missing problematic tissue areas of the stomach wall.

EUROPEAN PROJECTS

Nemo

A consortium funded by EU 6th Framework Programme (FP) during 2006-2010, worked on NEMO (Nano based capsule-Endoscopy with Molecular Imaging and Optical biopsy) project (http://fcs.itc.it/NEMO/) with main aim to develop an advance cancer screening method friendly enough to significantly increase compliance, simplify the diagnostic pathway and increase the sensitivity and specificity of early detection[35,42]. The concept of the NEMO approach was to combine capsule endoscopy with nano-based molecular recognition that would highlight cancerous and precancerous lesions in the GI tract, thus considerably increasing the accuracy and ease of diagnosis[43]. The consortium aimed to merge few technological platforms: nanotechnology for targeting and marking possible lesion; capsule endoscopy for detecting the marked disorder; capsule manoeuvring technologies to move the autonomous NEMO capsule backwards and forwards in the gastrointestinal tract as well as miniaturization and low power technologies[20].

A major task of the project was to develop nanocontainers, labeled with targeting agents and filled with dyeing material. The administered nanocontainers will be tailored to react with the target and mark the intestinal lesion. Another task was to develop a capsule based on narrow band imaging, combining specific optical filters with light emitting diodes (LEDs). The overall clinical outcomes from the NEMO project are summarized in ftp://ftp.cordis.europa.eu/pub/lifescihealth/docs/nemo.pdf[43].

Versatile Endoscopy Capsule for gastrointestinal TuMor Recognitions and therapy

Versatile Endoscopy Capsule for gastrointestinal TumOr Recognitions and therapy (VECTOR) (http://www.vector-project.com) was another project funded by EU FP6 during 2006-2010, aiming to develop a miniaturized robotic wireless endoscope for both diagnosis and therapy in the human digestive tract, with particular focus on the diagnosis and treatment of gastrointestinal cancer and its precursors[44]. VECTOR delivered different versions of prototype capsule endoscopes, for different clinical needs. An impressive prototype was a tele-operated robotic capsule with eight legs, capable of walking within the colon at a speed of about 3 cm per minute against peristalsis[45]. Main limitation of this capsule was the difficulty to include a power source, because most of the space within the capsule was covered by the electro-mechanic components used for moving the robotic legs. The fabrication of such a robotic spider-like capsule is described in a meso-scale engineering case study[46].

Another prototype included a submersible capsule prototype equipped with four propellers for locomotion. This capsule was intended mainly for stomach inspection. Reliable locomotion and steering within the stomach was possible after ingestion of clear water[47]. A wirelessly powered version of this prototype, based on wireless energy harvesting techniques, was later proposed for longer lasting operation. Another alternative investigated for in-vivo capsule navigation was based on externally applied magnetic fields. The magnetic field was generated by a permanent magnet attached to a robotic arm operated above the body[48]. In the context of VECTOR project magnetic control was also applied in the context of therapeutic capsules, designed to apply a customized haemostatic endoscopic clip (OTSC® clip, Ovesco Endoscopy AG, Germany) to the bowel wall[49].

TROY

Another project funded by EU FP6 during 2006-2009 was TROY (endoscope capsule using ultrasound technology) (http://cordis.europa.eu/documents/documentlibrary/125822911EN6.pdf). Its objective was to prove the concept of a diagnostic system for prevention and early warning of superficial cancer and pre-malignant precursor lesions in gastrointestinal tract using ultrasound technology. The outcome of the project was a WCE platform that instead of a camera with an optical imaging sensor, uses miniaturized ultrasound probes to create 3D computer generated images. The validation of this platform was performed with a synthetic bowel simulator with realistic tissue response.

SUPCAM

The SUPCAM project (http://www.supcam.eu) was funded by EU FP7 during 2012-2014 for the development of a spherical endoscopic capsule which can be safely - and accurately - guided along the colonic lumen from the outside, through an electromagnet, in completely wireless mode[50]. Being designed for the evaluation of the colon, this capsule is not for ingestion, but to be administered as a suppository. The endoscopic device will have the appearance of a small (~2 cm in diameter) spherical capsule (there are in fact 2 layers, the inner sphere and the outer sphere) with a compact external control system, easily manageable by the medical operator in order to support him in the diagnosis of colon rectal diseases[51]. The external device will be compact and adapted to be transported and made suitable the room of the majority of common outpatient settings, similarly to an ultrasound scanner. So far, this new system has been presented in international scientific meetings and there is data on the effectiveness of the magnetic control system with laboratory simulations[51] as well as with a phantom colon model[52]. To the best of our knowledge, no data on animal models are available yet.

OTHER CAPSULE ENDOSCOPE DESIGNS

Mermaid capsule

A magnetically propelled fish-inspired capsule endoscope has been proposed as another navigable solution[53]. This capsule is called Mini Mermaid (MM1) it has a size of 12 mm × 45 mm and a flexible silicone fin at its back with a small magnet attached at its end. It can change velocity and direction by adjusting the waveform of the electric current running in magnetic coils controlling the motion of the fin. Its velocity can reach dozens of cm per second, while its battery can last for eight to 10 h. It has been applied safely for the examination of the whole digestive tract of a human, after providing him with doses of polyethylene glycol and water.

Multiple cameras endoscopy capsule

In order to maximize the diagnostic yield of CE, Gu et al[54] proposed recently an ingenious spherical device – called the MicroBall – with multiple (six) cameras (Figure 8). Although this platform allows no controllability of the capsule movement, it promises a more accurate revision of the GI tract and reduction of the false negatives by multiple cameras with ‘smart’ image capture strategy, employing not only movements sensitive control but camera selection as well[54].

Robotic biopsy

The concept of a robotic biopsy device was certainly influenced by the Crosby capsule. To date, the proposed devices consist of modules for the complete process of biopsy, which includes monitoring the intestinal wall by a tissue monitoring module, aligning onto a polyp by an anchor, and sampling of the polyp tissue by a biopsy module. The latter is constructed to remember the guillotine opening of the Crosby capsule and its control/actuation –in the prototypes and conceptual models- is left to the control of the physician[14,15,35,55].

Motility tracking system

Motility Tracking System (MTS) is swallowable, small magnetic pill (6 mm × 18 mm, density 1.8 g/cm3) coated with a plastic capsule[56]. The small electronic pill, when ingested by the patient, emits a signal during its entire transit through the GI tract. Its movements are tracked by a detector plate of 16 magnetic field sensors (4 × 4) with a frequency of 10 Hz. Data from each sensor with an iterative algorithm are used to calculate the position and orientation of the magnetic pill[56,57]. The first generation, radiation-free pill (MTS-1), has been introduced for description of gastrointestinal motility in adults. The current version, Motility Tracking System-2 (MTS2), is also known as 3D-Transit (Figure 9) and is able to track multiple pills simultaneously[57]. This signal is recorded and subsequently downloaded to a data management station.

Wireless motility/pH capsule - SmartPill

The wireless motility/pH capsule (WMC or SmartPill) is data recording device that enables the simultaneous assessment of regional and whole gut transit. It has been approved by the FDA for the evaluation of patients with suspected gastroparesis and chronic idiopathic constipation[58,59]. SmartPill measures the temperature, pH, and pressure while traveling through the GI until exiting the body through the anus. Validated patterns in pH and temperature recordings allow for accurate measurement of gastric emptying, small bowel transit, colonic transit, and whole gut transit times[60]. Clinical trials showed it to be a suitable alternative to scintigraphy and radiopaque marker studies in measuring gastric emptying, small bowel, colonic, and whole gut transit times[59,61]. The WMC should be considered the transit study of choice for individuals suspected of having altered transit in more than one region of the gastrointestinal tract.

Lab-in-a-pill

This capsule has resulted by a collaboration between the Universities of Glasgow, Edinburgh and Strathclyde, and the Institute for System Level Integration in Scotland[2,37]. The capsule device, called Lab-on-a-Pill, measures 16 mm × 55 mm and weighs 13.5 g[37,62], with temperature, pH, conductivity, and dissolved oxygen sensors. The initial aim is to develop new tools to help the Government in its objective for screening colon cancers in all people aged over 50 years. The eventual device will allow wireless detection of markers for bowel cancer, a disease that is common in many countries. However, there have been no trials carried out using this capsule with the exception of artificial gut trials[37].

Tethered capsule endomicroscopy

The system developed by Gora et al[63,64] involves a capsule containing optical frequency domain imaging (OFDI) technology - a rapidly rotating laser tip emitting a beam of near-infrared light - and sensors that record light reflected back from the oesophageal mucosa. The capsule is attached to a string-like tether that connects to the imaging console and allows a physician or other health professional to control the system[65]. Following ingestion, the capsule is carried down the oesophagus by natural peristalsis. OFDI images are taken throughout the capsule’s transit down and up the esophagus.

Check-Cap

Check-Cap (Check-Cap, Mount Carmel, Israel) is a new technology in development for colon imaging[66]. It is proposed that tiny X-ray Radar device (the exposure to radiation is described as minimal, i.e., total dose equivalent to a single plain abdominal radiograph) creates a 3-D reconstructed image of the colon[67]. The major advantage of this platform is that no prior colon cleansing is needed, although the patient is still required to drink an oral contrast solution to label faecal material[67,68]. Eventually, the capsule transmits data to a wrist-worn recorder[66]. It is understood that the clinical performance of Check-Cap device is under investigation.

Sonopill

A £5-million grant by the Engineering and Physical Sciences Research Council (EPSRC) went to Sonopill, an ambitious Dundee University-led project aim to incorporate diagnostic and therapeutic ultrasound along with optical imaging, all deployed in capsule format[69]. It is aspired that this will complemented by other sensors, e.g., for pH measurement and auto-fluorescence imaging (AFI) for multimodal diagnosis[70]. In the longer term, the investigators aim to produce a family of diagnostic and therapeutic capsules, with the patient pathway defined according to diagnostic information assembled. Essentially, the Sonopill aims to push that technology further by proving in practice the complementary role of ultrasound and visual imaging with studies of multimodal diagnosis and therapy.

Endoscopic capsule robots using reconfigurable modular assembly

In order to enhance the locomotion capabilities and the functionalities of current capsule endoscopy platforms, Yoo et al (Yoo SS, Rama S, Szewczyk B, Pui JW, Lee W, Kim L. Endoscopic capsule robots using reconfigurable modular assembly: A pilot study. International Journal of Imaging Systems and Technology 2014; 24: 359-365) proposed a modular system of robotic capsules containing steerable locomotive elements that are capable of assembling into a larger and more complex robot via mutual docking. Docking is based on permanent magnets for energy efficiency purposes. Such robots could provide abilities that are lacking from current capsule endoscopes. A prototype was developed to demonstrate how a camera-enabled wireless capsule endoscope can be enhanced with navigation capabilities by docking on it three capsules with angular and rotational servos acting like legs. Such a modular robot can be used for thorough inspection of regions of the GI tract. Issues to be addressed before such a capsule is applicable in clinical practice, include the use of biocompatible polymer materials for enhanced friction in the slippery environment of the GI tract, energy conservation, e.g., using wireless power transmission, water proofing, and miniaturization.

Inchworm-like micro-robot

A prototype miniaturized navigational capsule endoscope capable of sliding like an inchworm was developed by Kim et al (Kim B, Lee S, Park JH, Park JO. Inchworm-like microrobot for capsule endoscope. In: Robotics and Biomimetics, ROBIO 2004. IEEE International Conference; IEEN, 2004: 458-460), and validated in vitro. The capsule is equipped with four actuators and clampers that mimic insect claws (Scherge M and Gorb S. Biological Micro- and Nano- tribology, Nature’s Solutions. Springer, 2000: 79-150) It clamps the contact surface and the body moves forward during the contraction of a rear linear actuator. The clamper releases the contact surface and slides forward as the front linear actuator is contracting. In this spirit, another locomotion mechanism inspired by earthworms, is based on repeated elongations and retractions of a clamping device (An earthworm-like locomotive mechanism for capsule endoscopes. Intelligent Robots and Systems. 2005 IEEE/RSJ International Conference; 2005: 2997-3002).

NANOTECHNOLOGY

Microrobots are 10 μm - 1 mm scale robots offering great new potential in medicine and nanotechnology is defined as any process performed at nanoscale. Electronics are constantly miniaturized and increase performance. Inevitably capsule endoscopy devices will follow the same trend as they did in the past leading to integration of multiple sensors (i.e., temperature, pH, iron detector, etc.) while maintaining power efficiency. Furthermore, such miniature endoscopy devices could attach to the intestinal mucosa while getting power support from human body heat allowing them to operate for long monitoring periods. Finally with the advent of nanotechnology medicine in recent years it will slowly be incorporated to capsule endoscopy devices allowing them to deliver smart drugs (i.e., targeted drug therapies), diagnosis and treatment to specific parts of the digestive tract.

CONCLUSION

In conclusion, several analogues could be drawn for the field of medicine/gastroenterology and other sciences as well. Therefore, we believe that the mainstream small bowel endoscopy in the third decade of the new millennium should be provided by an enhanced version of the capsule-based based platform, such as the one proposed by Iakovidis et al[71]. Just like aeronautical engineering – for more than a century now – has not significantly moved away from the conceptual design/idea of its most famous aviation pioneers, the Wright brothers, these authors believe that the shell of the capsule device – with some optimizations, e.g., spherical capsules would go easier through stenotic parts – will prevail over time. Furthermore, the speed (and control of the device), the functional characteristics (definition, 3-D application and therapy) and the indications for obtaining it will change over time.

Bio-mimetic/bio-inspired approaches, such as the spider and fish -like capsules[46,53], are promising approaches not only with respect to navigation within the digestive tract, but also for treatment, not only with application of haemostatic clips or drug delivery[2,37,38,40,41], but also for small operations, including biopsies or small dissections.

In conclusion, we should relish the challenges that lie ahead; in an environment like the small-bowel, the requirement is not to make micro-capsules; instead the attention should focus on how to utilize the existing carrier shape and size, but with miniaturized components such as microscopic batteries, that will leave internal space and provide enough power for internal lens rotation (just like the deep sea submarines), space for microscopic labs (lab-in-a-pill) and other sensing capacities and –hopefully- deliver powder medication for bleeding.

REFERENCES

1 Jacobson B, Mackay RS. A pH-endoradiosonde. Lancet 1957; 272: 1224 [PMID: 13440011] 2 McCaffrey C, Chevalerias O, O'Mathuna C, Twomey K. Swallowable-Capsule Technology. Pervasive Computing, IEEE 2008; 7: 23-29 [DOI: 10.1109/MPRV.2008.17] 3 Romero-Vázquez J, Argüelles-Arias F, García-Montes JM, Caunedo-Álvarez Á, Pellicer-Bautista FJ, Herrerías-Gutiérrez JM. Capsule endoscopy in patients refusing conventional endoscopy. World J Gastroenterol 2014; 20: 7424-7433 [PMID: 24966612 DOI: 10.3748/wjg.v20.i23.7424] 4 Bouchard S, Ibrahim M, Van Gossum A. Video capsule endoscopy: perspectives of a revolutionary technique. World J Gastroenterol 2014; 20: 17330-17344 [PMID: 25516644 DOI: 10.3748/wjg.v20.i46.17330] 5 Iddan G, Meron G, Glukhovsky A, Swain P. Wireless capsule endoscopy. Nature 2000; 405: 417 [PMID: 10839527] 6 Adler SN, Bjarnason I. What we have learned and what to expect from capsule endoscopy. World J Gastrointest Endosc 2012; 4: 448-452 [PMID: 23189215 DOI: 10.4253/wjge.v4.i10.448] 7 Gong F, Swain CP, Mills TN. An endorobot for gastrointestinal endoscopy. Gut 1994; 35: S52 [PMID: 8125392] 8 Koulaouzidis A, Rondonotti E, Karargyris A. Small-bowel capsule endoscopy: a ten-point contemporary review. World J Gastroenterol 2013; 19: 3726-3746 [PMID: 23840112 DOI: 10.3748/wjg.v19.i24.3726] 9 Iakovidis DK, Koulaouzidis A. Software for enhanced video capsule endoscopy: Challenges for essential progress. Nat Rev Gastroenterol Hepatol 2015; In press 10 Parker CE, Spada C, McAlindon M, Davison C, Panter S. Capsule endoscopy--not just for the small bowel: a review. Expert Rev Gastroenterol Hepatol 2015; 9: 79-89 [PMID: 25484107 DOI: 10.1586/17474124.2014.934357] 11 Hosoe N, Naganuma M, Ogata H. Current status of capsule endoscopy through a whole digestive tract. Dig Endosc 2015; 27: 205-215 [PMID: 25208463 DOI: 10.1111/den.12380] 12 Eliakim R. Video capsule endoscopy of the small bowel. Curr Opin Gastroenterol 2013; 29: 133-139 [PMID: 23221650 DOI: 10.1097/MOG.0b013e32835bdc03] 13 Slawinski PR, Oleynikov D, Terry BS. Intestinal biomechanics simulator for robotic capsule endoscope validation. J Med Eng Technol 2015; 39: 54-59 [PMID: 25367667 DOI: 10.3109/03091902.2014.973619] 14 Ciuti G, Menciassi A, Dario P. Capsule endoscopy: from current achievements to open challenges. IEEE Rev Biomed Eng 2011; 4: 59-72 [PMID: 22273791 DOI: 10.1109/RBME.2011.2171182] 15 Sliker LJ, Ciuti G. Flexible and capsule endoscopy for screening, diagnosis and treatment. Expert Rev Med Devices 2014; 11: 649-666 [PMID: 25148269 DOI: 10.1586/17434440.2014.941809] 16 Sharma VK. The future is wireless: advances in wireless diagnostic and therapeutic technologies in gastroenterology. Gastroenterology 2009; 137: 434-439 [PMID: 19545570 DOI: 10.1053/j.gastro.2009.06.029] 17 Toennies JL, Tortora G, Simi M, Valdastri P, Webster III R J. Swallowable medical devices for diagnosis and surgery: the state of the art. J Mech E 2010; 224: 1397-1414 [DOI: 10.1243/09544062JMES1879] 18 Wang A, Banerjee S, Barth BA, Bhat YM, Chauhan S, Gottlieb KT, Konda V, Maple JT, Murad F, Pfau PR, Pleskow DK, Siddiqui UD, Tokar JL, Rodriguez SA. Wireless capsule endoscopy. Gastrointest Endosc 2013; 78: 805-815 [PMID: 24119509 DOI: 10.1016/j.gie.2013.06.026] 19 Pezzoli A. Wireless endoscopy capsules should not be released in the environment. Gastrointest Endosc 2014; 80: 191-192 [PMID: 24950652 DOI: 10.1016/j.gie.2014.02.021] 20 Fisher LR, Hasler WL. New vision in video capsule endoscopy: current status and future directions. Nat Rev Gastroenterol Hepatol 2012; 9: 392-405 [PMID: 22565098 DOI: 10.1038/nrgastro.2012.88] 21 Japan Inc Communications, Inc. A Capsule Camera to Save Stomachs. Available from: URL: http: //www.japaninc.com/article.php?articleID=1130 22 Koulaouzidis A, Iakovidis DK, Karargyris A, Plevris JN. Optimizing lesion detection in small-bowel capsule endoscopy: from present problems to future solutions. Expert Rev Gastroenterol Hepatol 2015; 9: 217-235 [PMID: 25169106] 23 Basar MR, Malek F, Juni KM, Shaharom Idris M, Iskandar M. Saleh M. Ingestible Wireless Capsule Technology: A Review of Development and Future Indications. Int J Antennas Propag 2012; Article ID 807165 [DOI: 10.1155/2012/807165] 24 Basar MR, Ahmad MY, Cho J, Ibrahim F. Application of wireless power transmission systems in wireless capsule endoscopy: an overview. Sensors (Basel) 2014; 14: 10929-10951 [PMID: 24949645 DOI: 10.3390/s140610929] 25 RF Co., Ltd. Sayaka. Available from: URL: http: //rfsystemlab.com/en/sayaka/ 26 Filip D, Yadid-Pecht O, Mintchev MP. Progress in self-stabilizing capsules for imaging of the large intestine. In: 17th IEEE International Conference on Electronics, Circuits, and Systems (ICECS); 2010: 231–234 27 Filip D, Yadid-Pecht O, Andrews CN, Mintchev MP. Self-stabilizing colonic capsule endoscopy: pilot study of acute canine models. IEEE Trans Med Imaging 2011; 30: 2115-2125 [PMID: 21803680 DOI: 10.1109/TMI.2011.2163165] 28 Filip D, Yadid-Pecht O, Muench G, Mintchev MP, Andrews CN. Suture marker lesion detection in the colon by self-stabilizing and unmodified capsule endoscopes: pilot study in acute canine models. Gastrointest Endosc 2013; 77: 272-279 [PMID: 23317692 DOI: 10.1016/j.gie.2012.10.016] 29 Bourbakis N, Giakos G, Karargyris A. Design of new-generation robotic capsules for therapeutic and diagnostic endoscopy. Imaging Systems and Techniques (IST), 2010 IEEE International Conference on; 2010 July 1-2; Thessaloniki, Greece. IEEE, 2010: 1-6 [DOI: 10.1109/IST.2010.5548467] 30 Karargyris A, Koulaouzidis A. Capsule-odometer: a concept to improve accurate lesion localisation. World J Gastroenterol 2013; 19: 5943-5946 [PMID: 24124345 DOI: 10.3748/wjg.v19.i35.5943] 31 Karargyris A, Koulaouzidis A. OdoCapsule: next-generation wireless capsule endoscopy with accurate lesion localization and video stabilization capabilities. IEEE Trans Biomed Eng 2015; 62: 352-360 [PMID: 25167544] 32 Wong Kee Song LM, Banerjee S, Barth BA, Bhat Y, Desilets D, Gottlieb KT, Maple JT, Pfau PR, Pleskow DK, Siddiqui UD, Tokar JL, Wang A, Rodriguez SA. Emerging technologies for endoscopic hemostasis. Gastrointest Endosc 2012; 75: 933-937 [PMID: 22445927 DOI: 10.1016/j.gie.2012.01.024] 33 Smith LA, Stanley AJ, Bergman JJ, Kiesslich R, Hoffman A, Tjwa ET, Kuipers EJ, von Holstein CS, Oberg S, Brullet E, Schmidt PN, Iqbal T, Mangiavillano B, Masci E, Prat F, Morris AJ. Hemospray application in nonvariceal upper gastrointestinal bleeding: results of the Survey to Evaluate the Application of Hemospray in the Luminal Tract. J Clin Gastroenterol 2014; 48: e89-e92 [PMID: 24326829 DOI: 10.1097/MCG.0000000000000054] 34 Moglia A, Menciassi A, Schurr MO, Dario P. Wireless capsule endoscopy: from diagnostic devices to multipurpose robotic systems. Biomed Microdevices 2007; 9: 235–243 [PMID: 17160703] 35 Swain P. The future of wireless capsule endoscopy. World J Gastroenterol 2008; 14: 4142-4145 [PMID: 18636658] 36 Pasricha T, Smith BF, Mitchell VR, Fang B, Brooks ER, Gerding JS, Washington MK, Valdastri P, Obstein KL. Controlled colonic insufflation by a remotely triggered capsule for improved mucosal visualization. Endoscopy 2014; 46: 614-618 [PMID: 24845802 DOI: 10.1055/s-0034-1365497] 37 Twomey K, Marchesi JR. Swallowable capsule technology: current perspectives and future directions. Endoscopy 2009; 41: 357-362 [PMID: 19340742 DOI: 10.1055/s-0028-1119640] 38 Munoz F, Alici G, Li W. A review of drug delivery systems for capsule endoscopy. Adv Drug Deliv Rev 2014; 71: 77-85 [PMID: 24384373 DOI: 10.1016/j.addr.2013.12.007] 39 Martin NE, Collison KR, Martin LL, Tardif S, Wilding I, Wray H, Barrett JS. Pharmacoscintigraphic assessment of the regional drug absorption of the dual angiotensin-converting enzyme/neutral endopeptidase inhibitor, M100240, in healthy volunteers. J Clin Pharmacol 2003; 43: 529-538 [PMID: 12751274] 40 Woods SP, Constandinou TG. Towards a micropositioning system for targeted drug delivery in wireless capsule endoscopy. Conf Proc IEEE Eng Med Biol Soc 2011; 2011: 7372-7375 [PMID: 22256042 DOI: 10.1109/IEMBS.2011.6091717] 41 Glass P, Cheung E, Sitti M. A legged anchoring mechanism for capsule endoscopes using micropatterned adhesives. IEEE Trans Biomed Eng 2008; 55: 2759-2767 [PMID: 19126455 DOI: 10.1109/TBME.2008.2002111] 42 Keller J, Fibbe C, Volke F, Gerber J, Mosse AC, Reimann-Zawadzki M, Rabinovitz E, Layer P, Swain P. Remote magnetic control of a wireless capsule endoscope in the esophagus is safe and feasible: results of a randomized, clinical trial in healthy volunteers. Gastrointest Endosc 2010; 72: 941-946 [PMID: 20855064 DOI: 10.1016/j.gie.2010.06.053] 43 Nano based capsule-Endoscopy with Molecular Imaging and Optical biopsy. Available from: URL: ftp: //ftp.cordis.europa.eu/pub/lifescihealth/docs/nemo.pdf 44 Schostek S, Schurr MO. European Research on Wireless Endoscopy–the VECTOR Project. In PHealth 2013: Proceedings of the 10th International Conference on Wearable Micro and Nano Technologies for Personalized Health; 2013 Jun; IOS Press, 2013: 193 45 Quirini M, Menciassi A, Scapellato S, Dario P, Rieber F, Ho CN, Schostek S, Schurr MO. Feasibility proof of a legged locomotion capsule for the GI tract. Gastrointest Endosc 2008; 67: 1153-1158 [PMID: 18513557] 46 Quaglia C, Buselli E, Webster III RJ, Valdastri P, Menciassi A, Dario P. An endoscopic capsule robot: a meso-scale engineering case study. J Micromech Microeng 2009; 19: 105007 [DOI: 10.1088/0960-1317/19/10/105007] 47 Tortora G, Valdastri P, Susilo E, Menciassi A, Dario P, Rieber F, Schurr MO. Propeller-based wireless device for active capsular endoscopy in the gastric district. Minim Invasive Ther Allied Technol 2009; 18: 280-290 [PMID:19707936 DOI: 10.1080/13645700903201167] 48 Ciuti G, Donlin R, Valdastri P, Arezzo A, Menciassi A, Morino M, Dario P. Robotic versus manual control in magnetic steering of an endoscopic capsule. Endoscopy 2010; 42: 148-152 [PMID:20017088 DOI: 10.1055/s-0029-1243808] 49 Valdastri P, Quaglia C, Susilo E, Menciassi A, Dario P, Ho CN, Anhoeck G, Schurr MO. Wireless therapeutic endoscopic capsule: in vivo experiment. Endoscopy 2008; 40: 979-982 [PMID: 19065478 DOI: 10.1055/s-0028-1103424] 50 Supcam. News and events. Available from: URL: http: //www.supcam.eu/news-and-events 51 Lucarini G, Mura M, Ciuti G, Menciassi A. The SUPCAM platform: electromagnetic navigation system for active capsule endoscope. IEEE International Magnetics Conference 2014; Available from: URL: http: //www.supcam.eu/news-and-events/24-electromagnetic-navigation-system-for-active-capsule-endoscope 52 Tozzi A, Ciuti G, Lucarini G, Mura M, Quaglia C, Menciassi A, Battaglia G. Supcam European Project: Preliminary Prototyping and Test of A New Generation Active Endoscopic Colon Capsule. Conference: UEG Week 2014 [DOI: 10.13140/2.1.1935.8722] 53 Morita E, Ohtsuka N, Shindo Y, Nouda S, Kuramoto T, Inoue T, Murano M, Umegaki E, Higuchi K. In vivo trial of a driving system for a self-propelling capsule endoscope using a magnetic field (with video). Gastrointest Endosc 2010; 72: 836-840 [PMID: 20883863 DOI: 10.1016/j.gie.2010.06.016] 54 Gu Y, Xie X, Li G, Sun T, Wang D, Yin Z, Zhang P, Wang Z. Design of Endoscopic Capsule With Multiple Cameras. IEEE Trans Biomed Circuits Syst 2014 Nov 5; Epub ahead of print [PMID: 25376042] 55 Kyoungchul K, Sehyuk Y, Sunhee C, Jeon D. A Robotic Biopsy Device for Capsule Endoscopy. J Med Devices 2012; 6: 031004 [DOI: 10.1115/1.4007100] 56 Hedsund C, Joensson IM, Gregersen T, Fynne L, Schlageter V, Krogh K. Magnet tracking allows assessment of regional gastrointestinal transit times in children. Clin Exp Gastroenterol 2013; 6: 201-208 [PMID: 24399881 DOI: 10.2147/CEG.S51402] 57 Motilis Sàrl. Products: MTS-2. Available from: URL: http: //www.motilis.com/V2/index.php?nav=3&passit= 58 Hasler WL. The use of SmartPill for gastric monitoring. Expert Rev Gastroenterol Hepatol 2014; 8: 587-600 [PMID: 24881810 DOI: 10.1586/17474124.2014.922869] 59 Tran K, Brun R, Kuo B. Evaluation of regional and whole gut motility using the wireless motility capsule: relevance in clinical practice. Therap Adv Gastroenterol 2012; 5: 249-260 [PMID: 22778790 DOI: 10.1177/1756283X12437874] 60 Rao SS, Kuo B, McCallum RW, Chey WD, DiBaise JK, Hasler WL, Koch KL, Lackner JM, Miller C, Saad R, Semler JR, Sitrin MD, Wilding GE, Parkman HP. Investigation of colonic and whole-gut transit with wireless motility capsule and radiopaque markers in constipation. Clin Gastroenterol Hepatol 2009; 7: 537-544 [PMID: 19418602 DOI: 10.1016/j.cgh.2009.01.017] 61 Kuo B, McCallum RW, Koch KL, Sitrin MD, Wo JM, Chey WD, Hasler WL, Lackner JM, Katz LA, Semler JR, Wilding GE, Parkman HP. Comparison of gastric emptying of a nondigestible capsule to a radio-labelled meal in healthy and gastroparetic subjects. Aliment Pharmacol Ther 2008; 27: 186-196 [PMID: 17973643 DOI: 10.1111/j.1365-2036.2007.03564.x] 62 Johannessen EA, Wang L, Wyse C, Cumming DR, Cooper JM. Biocompatibility of a lab-on-a-pill sensor in artificial gastrointestinal environments. IEEE Trans Biomed Eng 2006; 53: 2333-2340 [PMID: 17073339 DOI: 10.1109/TBME.2006.883698] 63 Gora MJ, Sauk JS, Carruth RW, Lu W, Carlton DT, Soomro A, Rosenberg M, Nishioka NS, Tearney GJ. Imaging the upper gastrointestinal tract in unsedated patients using tethered capsule endomicroscopy. Gastroenterology 2013; 145: 723-725 [PMID: 23932950 DOI: 10.1053/j.gastro.2013.07.053] 64 Gora MJ, Sauk JS, Carruth RW, Gallagher KA, Suter MJ, Nishioka NS, Kava LE, Rosenberg M, Bouma BE, Tearney GJ. Tethered capsule endomicroscopy enables less invasive imaging of gastrointestinal tract microstructure. Nat Med 2013; 19: 238-240 [PMID: 23314056 DOI: 10.1038/nm.3052] 65 National Documentation Centre. Health News Repository. Available from: URL: http: //www.nihfw.org/doc/Health News Repository- Jan-Mar 2013.pdf 66 Check-Cap. The first non-invasive, prep-free imaging test for colon cancer. Available from: URL: http: //www.check-cap.com 67 Chatrath H, Rex DK. Potential screening benefit of a colorectal imaging capsule that does not require bowel preparation. J Clin Gastroenterol 2014; 48: 52-54 [PMID: 23632347 DOI: 10.1097/MCG.0b013e318288a2cd] 68 Check-Cap. Available from: URL: http: //www.check-cap.com/Abstract-DDW-2014.pdf 69 Heriot-Watt University. Sonopill. Available from: URL: http: //www.eps.hw.ac.uk/institutes/sensors-signals-systems/grant-sonopill.htm 70 BBC news. Dundee-led project to create an ultrasound pill awarded £5m. Available from: URL: http: //www.bbc.co.uk/news/uk-scotland-tayside-central-21755003 71 Iakovidis DK, Sarmiento R, Silva JS, Histace A, Romain O, Koulaouzidis A, Dehollain C, Pinna A, Granado B, Dray X. Towards Intelligent Capsules for Robust Wireless Endoscopic Imaging of the Gut. IEEE International Conference on Imaging Systems and Techniques; 2014 Oct; Santorini, Greece. IEEE, 2014: 95-100

P-Reviewer: Higaki S, Iakovidis DK, Redondo-Cerezo E

S-Editor: Yu J L-Editor:  E-Editor:

A

B

Figure 1 NORIKA and SAYAKA capsules

A

B

Figure 2 Self-stabilizing capsule endoscope.

A

B

Figure 3 Odocapsule.

Figure 4 Coagulation capsule.

A

B

Figure 5 Enterion capsule.

Figure 6 InteliSite capsule.

A

B

Figure 7 The micro-positioning mechanism used capsule body (A) and holding mechanism.

A

B

Figure 8 MicroBall.

Figure 9 Motility tracking system.