**Name of Journal: *World Journal of Methodology***

**ESPS Manuscript NO: 22397**

**Manuscript Type: Editorial**

**Cation-exchange high-performance liquid chromatography for variant hemoglobins and HbF/A2: What must hematopathologists know about methodology?**

Sharma P *et al.* Hemoglobin HPLC methodology: Essentials for hematopathologists

**Prashant Sharma, Reena Das**

**Prashant Sharma, Reena Das,** Department of Hematology, Postgraduate Institute of Medical Education and Research, Chandigarh 160012, India

**Author contributions:** Sharma P and Das R both made substantial contributions to conception of this article; Sharma P drafted the article and Das R made critical revisions related to important intellectual content. Both authors gave final approval of the version of the article submitted.

**Conflict-of-interest** **statement:** Both authors of the paper declare that they have no conflicting interests (including commercial, personal, political, intellectual or religious interests) that are related to the work submitted for consideration for publication.

**Open-Access:** This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: <http://creativecommons.org/licenses/by-nc/4.0/>

**Correspondence to: Dr. Prashant Sharma, MD, DNB, DM, Associate Professor,** Department of Hematology, Level 5, Research Block A, Postgraduate Institute of Medical Education and Research, Sector 12, Chandigarh 160012, India. [sharma.prashant@pgimer.edu.in](mailto:sharma.prashant@pgimer.edu.in)

**Telephone:** +91-88-72016123

**Received:** August 28, 2015

**Peer-review started:** September 1, 2015

**First decision:** December 7, 2015

**Revised:** January 28, 2016

**Accepted:** February 23, 2016

**Article in press:**

**Published online:**

**Abstract**

Cation-exchange high-performance liquid chromatography (CE-HPLC) is a widely used laboratory test to detect variant hemoglobins as well as quantify hemoglobins F and A2 for the diagnosis of thalassemia syndromes. It’s versatility, speed, reproducibility and convenience have made CE-HPLC the method of choice to initially screen for hemoglobin disorders. Despite its popularity, several methodological aspects of the technology remain obscure to pathologists and this may have consequences in specific situations. This paper discusses the basic principles of the technique, the initial quality control steps and the interpretation of various controls and variables that are available on the instrument output. Subsequent sections are devoted to methodological considerations that arise during reporting of cases. For instance, common problems of misidentified peaks, totals crossing 100%, causes of total area being above or below acceptable limits and the importance of pre-integration region peaks are dealt with. Ultimately, CE-HPLC remains an investigation, the reporting of which combines in-depth knowledge of the biological basics with more than a working knowledge of the technological aspects of the technique.

**Key words:** Anemia; Diagnosis; Hematological disorders; Hematopathology; Hemoglobin; Hemoglobinopathies; High-performance liquid chromatography; Laboratory instrumentation; Red blood cells; Thalassemia

**© The Author(s) 2016.** Published by Baishideng Publishing Group Inc. All rights reserved.

**Core tip:** Interpretation of cation-exchange high-performance liquid chromatography requires in-depth knowledge of the biological basics of the disorders of hemoglobin with knowledge of the technological aspects of the technique. Pathologists may be unaware of the nuances of the technique, the rigorous quality control required and the approach to pitfalls that may be encountered. Here we list the most common of these, and based on literature and our experience, attempt to guide novices in this exciting and useful technology.

Sharma P, Das R. Cation-exchange high-performance liquid chromatography for variant hemoglobins and HbF/A2: What must hematopathologists know about methodology? *World J Methodol* 2016; In press

Cation-exchange high-performance liquid chromatography (CE-HPLC) of a red blood cell lysate is now established as a rapid, accurate and reproducible diagnostic technique to separate various human hemoglobin fractions[1]. It is performed routinely in many laboratories, mostly on fully automated systems. Interpretation of HPLC chromatograms is a complex process requiring inputs from the clinical background of the case as well as complete blood count data[1,2]. Here, we briefly discuss a few technological aspects that influence results, and therefore are concern to reporting pathologists.

**PRINCIPLE OF CE-HPLC**

The basic principle of CE-HPLC involves passing the analyte of interest (a mixture of hemoglobins in solution) at a high pressure (approximately 100-200 kg/cm²) through a cylindrical column packed with small spherical particles (typically 5 μm diameter silica gel, called the stationary phase). Very small sample volumes (usually approximately 5 μL) are applied to the column. Different hemoglobins adsorb onto the silica packing with different intensities based on their ionic interactions. The column is then perfused by a buffer (mobile phase) that constantly varies in pH and ionic strength. Different hemoglobins elute out with the perfusing buffer at different but characteristic time points in response to the continually changing salt gradient[3]. Another variable that affects elution/retention times, the column temperature, is kept fixed throughout the approximately 5.0 to 6.0 min run (Figure 1A).

**HEMOGLOBIN DETECTION AND ASSIGNMENT OF WINDOWS**

The eluted hemoglobin fractions are detected by a flow-cell type photometer that records changes in absorbance at 415 nm (hemoglobin) and 690 nm (background) on an integrating computer system. A chromatogram is generated displaying time on the X-axis and percentages on the Y-axis (Figure 1). The area under the absorption peak approximates the percentage of the fraction detected, and each fraction is assigned a window (*i.e.*, range of retention times). The software controls for overlapping/merging peaks by dropping vertical axes at the troughs[3,4]. The report prepared incorporates numerical as well as graphical data, and their analysis is discussed next. We use the very widely applied Bio-Rad Variant II Turbo output (Bio-Rad Laboratories, Hercules, United States) using the β-Thal Short Programme for illustration (Figure 1), however, the principles remain similar even on other systems[5]. The interpretation of various chromatogram regions and peaks in various windows is summarized in Table 1. A specimen chromatogram with the peaks highlighted is shown in Figure 2.

**CALIBRATORS AND CONTROLS**

All HPLC runs are preceded by priming and then calibration of instrument. Separate calibration factors are obtained for HbA2 and HbF as ratios of expected to obtained values. Since the two values should ideally be equal (*i.e*., a ratio of 1) these are deemed to have passed if they lie between 0.7 and 1.3. These calibration factors are then applied for all subsequent patient samples. The retention time of HbA2 in the calibrator is also a useful indicator of run reliability. Normally it lies between 2.60-2.70 min, and the instrument may need temperature adjustments if wider deviations occur. This is especially common as the column cartridge ages; usual cartridge lifetimes being around 250 injections. Bi-level controls, one normal (HbF 1%-2%, HbA2 1.8%-3.2%) and one elevated (HbF 5%-10%, HbA2 4%-6%), should be analyzed at the beginning as well as the end of each set of patient specimens. The high control in case of Bio-Rad instruments also contains a variant peak that must elute in the S-window. All peaks must be symmetrical, temperature variations being the most common cause again of asymmetry[3,6].

**METHODOLOGICAL CONSIDERATIONS DURING REPORTING**

Once the preliminary checks have passed, reporting of patient samples can proceed. During reporting, attention must be directed to the following areas.

***Total area of analysis***

This must lie between 1 to 3 million μVolt-seconds (Figure 1B). Specimens with lower areas (due to anemia) or increased values (due to polycythemia) must be reanalysed after appropriate manual concentration by removing plasma or dilution by removing red cells respectively[6].

***Cases with total area > 100%***

The total area is the sum of all individual peaks’ areas and can therefore, in the presence of overlapping peaks, can cross 100%. This is especially common in patients with beta-thalassemia major, where the HbF peak overlaps and usually obliterates P2 and P3. In such situations, the percentages can be calculated manually, by taking the area of the peak of interest, and dividing it by total area to get the proportion. For example, in Figure 2, the HbF%, instead of the implausible 100.5% can be calculated as F-peak area × 100 ÷ Total area (*i.e.*, 1537546 × 100 ÷ 1576229 = 97.5%).

***Small unknown peaks***

One or more unknown peaks often occur around the P2/P3 window (retention times 1.3-1.8; Figure 1C). These may be safely ignored if ≤ 1% of total area[6]. Unknown peaks at longer retention times and those > 1% should be paid greater attention as transfusion-transmitted peaks, HbA2’ may present as small peaks[7]. One may also review the sample run previously as carryover peaks are also usually small.

***Misidentified hemoglobin fractions***

In rare cases with very large abnormal peaks, the entire fraction may be misassigned to either another category, or as an unknown peak. This was commoner in older generation analyzers (like the Variant), but can still occur with broad-based HbD and HbF peaks (Figure 2). Alternative techniques are then required to establish the identity of the unknown peak.

***Pre-integration peaks***

These peaks, with retention time < 1 min, are not reflected in the tabular data and need to be looked for on the chromatogram (Figure 1D). The causes of such peaks include HbH, Hb Barts, bilirubin and acetylated HbF (Figure 3). The clinical background and other HPLC findings usually indicate their nature, if found. In addition, bilirubin peaks are usually early, very sharp and thin. HbH peaks are usually dual and of low to moderate height, while post-translationally modified F are usually multiple with their height proportionate to the HbF%[1,3,4]. If required, the software settings may be readjusted manually to include such peaks. This may be especially useful in cases with HbH disease.

***P2 and P3 peaks***

These represent post-translationally modified adult hemoglobin (HbA0) and show normal ranges of 3.8 ± 0.7 and 4.3 ± 0.4 respectively (unpublished data). P2 is comprised of glycated hemoglobin (HbA1c) and levels ≥ 6.5% should be mentioned in the report with the suggestion to exclude diabetes mellitus. Low P2 levels are seen in cases with reticulocytosis. Elevated P3 may indicate HbJ-Meerut (an α-globin chain variant). It is also elevated in cases with the HbE variant (that elutes in the HbA2 window) and in aged specimens[1,3,4]. Incidentally, the P1-peak is virtually always absent in normal specimens.

In conclusion, although CE-HPLC is a rapid, convenient and reliable investigation for hemoglobin disorders, it involves several methodological issues and nuances. Reporting pathologists must be aware of these to extract maximum information from this technology.

**REFERENCES**

1 **Wajcman H**, Moradkhani K. Abnormal haemoglobins: detection & amp; characterization. *Indian J Med Res* 2011; **134**: 538-546 [PMID: 22089618]

2 **Rangan A**, Sharma P, Dadu T, Saxena R, Verma IC, Bhargava M. β-Thalassemia mutations in subjects with borderline HbA₂ values: a pilot study in North India. *Clin Chem Lab Med* 2011; **49**: 2069-2072 [PMID: 21892914 DOI: 10.1515/CCLM.2011.696]

3 Chapter 2: Laboratory techniques for the identification of abnormalities of globin chain synthesis. In Bain BJ. In: Haemoglobinopathy Diagnosis. Bain BJ, editor. 2nd ed. Blackwell Publishing. Oxford, 2006: 26-62 [DOI: 10.1002/9780470988787.ch2]


4 Chapter 3: Haemoglobin pattern analysis (3.1. Chromatographic methods for HbA2 determination). In: Prevention of Thalassaemias and Other Haemoglobin Disorders (Volume 2: Laboratory Protocols). By Old J, Harteveld CL, Traeger-Synodinos J, et al. (eds). 2nd ed. Nicosia, Cyprus: Thalassaemia International Federation, 2012. Available from: URL: http://www.ncbi.nlm.nih.gov/books/NBK190579/

5 **Van Delft P**, Lenters E, Bakker-Verweij M, de Korte M, Baylan U, Harteveld CL, Giordano PC. Evaluating five dedicated automatic devices for haemoglobinopathy diagnostics in multi-ethnic populations. *Int J Lab Hematol* 2009; **31**: 484-495 [PMID: 19486364 DOI: 10.1111/j.1751-553X.2009.01158.x]


6 Instruction Manual: Variant™ II β-thalassemia Short Program Reorder Pack. Ref: 270-2103 and 270-2154 (Release Notes CD-R Version BV231100-A). Bio-Rad Laboratories Inc. Hercules, CA USA, 2015: 10-15

7 **Gupta SK**, Sharma M, Tyagi S, Pati HP. Transfusion-induced hemoglobinopathy in patients of beta-thalassemia major. *Indian J Pathol Microbiol* 2011; **54**: 609-611 [PMID: 21934236 DOI: 10.4103/0377-4929.85112]


**P-Reviewer:** Harn GL, Yao D **S-Editor:** Ji FF **L-Editor: E-Editor:**

****

**Figure 1 Specimen cation-exchange high-performance liquid chromatography output from the Bio-Rad Variant II Turbo instrument (Bio-Rad Laboratories, Hercules, United States) using the β-Thal Short Programme.** Label **A** indicates the total time of analysis (X-axis) is approximately 5 to 6 min. Label **B** indicates that the total area of analysis should lie between 1 and 3 million. Label **C** shows the unknown peaks that may occur, especially in the P2 and P3 regions. Label **D** depicts the preintegration phase (< 1 min) is not reflected in the table above and should be analyzed on the chromatogram. Label **E** shows where the problem of HbF concentration being calculated as > 100% is present in this case. Please see text for details of resolution of this problem.

****

**Figure 2 A specimen chromatogram to show the various peaks that may occur in health and disease.** The patient, with elevated HbA2 and mildly elevated HbF, most likely had β-thalassemia trait with increased HbF (HbF levels between 1%-5% are found in approximately 30% of β-thalassemia trait cases, and occasional ones may have even higher levels).



**Figure 3 A case of thalassemia major with a very prominent and broad-based HbF being reported by the instrument (an older Bio-Rad Variant) as a P2-peak.**

**Table 1 Interpretation of various chromatogram regions and commonly encountered peaks in various high-performance liquid chromatography windows (retention time ranges obtained from Bio-Rad kit inserts)**

|  |  |  |
| --- | --- | --- |
| **Region / window** | **Retention time** | **Interpretation** |
| Pre-integration region | < 1 min | Bilirubin, Hb H, Hb Barts, modified HbF |
| P1 peak | 0.74 | A minuscule peak usually found in specimens with increased HbF |
| F window | 0.98-1.22 | HbF, Hb Okayama |
| P2 window | 1.28-1.50 | Glycated HbA |
| P3 window | 1.50-1.90 | Aged samples, HbJ-Meerut, modified HbE |
| A window | 1.90-3.10 | HbA, Glycated HbS, intact Hb Koln |
| A2 window | 3.30-3.90 | HbE, Hb D-Iran, Hb Lepore, Hb G-Koushatta, Hb Zurich, Hb Korle Bu |
| D window | 3.90-4.30 | Hb D-Punjab, Hb G-Philadelphia |
| S window | 4.30-4.70 | HbS, Hb Q-Thailand, Hb Manitoba |
| C window | 4.90-5.30 | HbC, Hb Constant Spring, Hb Agenogi |
| Further unknown peaks | > 5.30 | Hb Q-India |