**Name of Journal: *World Journal of Clinical Pediatrics***

**ESPS Manuscript NO: 25696**

**Manuscript Type: Minireviews**

**Drug delivery interfaces: A way to optimize inhalation therapy in spontaneously breathing children**

Ari A. Drug delivery interfaces in spontaneously breathing children

**Arzu Ari**

**Arzu Ari,** Department of Respiratory Therapy, Georgia State University, Atlanta, GA 30303-3083, United States

**Author contributions:** Ari A is the sole author of this manuscript.

**Conflict-of-interest statement:** Ari A serves on the advisory board of Bayer Pharmaceuticals.

**Open-Access:** This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: <http://creativecommons.org/licenses/by-nc/4.0/>

**Manuscript source:** Invited manuscript

**Correspondence to: Arzu Ari, PhD, RRT, PT, FAARC,** Department of Respiratory Therapy, Georgia State University, 140 Decatur Street Suite 1228, Atlanta, GA 30303-3083, United States. [arzuari@hotmail.com](mailto:arzuari@hotmail.com)

**Telephone:** +1-404-4131269

**Fax:** +1-404-4131230

**Received:** March 20, 2016

**Peer-review started:** March 22, 2016

**First decision:** April 20, 2016

**Revised:** May 3, 2016

**Accepted:** July 11, 2016

**Article in press:**

**Published online:**

**Abstract**

There are several different types of drug delivery interfaces available on the market. Using the right interface for aerosol drug delivery to children is essential for effective inhalation therapy. However, clinicians usually focus on selecting the right drug-device combination and often overlook the importance of interface selection that lead to suboptimal drug delivery and therapeutic response in neonates and pediatrics. Therefore, it is necessary to critically assess each interface and understand its advantage and disadvantages in aerosol drug delivery to this patient population. The purpose of this paper is to provide a critical assessment of drug delivery interfaces used for the treatment of children with pulmonary diseases by emphasizing advantages and problems associated with their use during inhalation therapy.

**Key words:** Aerosols; Inhalation therapy; Children; masks; Mouthpiece; High flow nasal cannula; Blow-by; Hood; Spacer/valved holding chamber

**© The Author(s) 2016.** Published by Baishideng Publishing Group Inc. All rights reserved.

**Core tip:** Many interfaces exist for aerosol drug delivery to spontaneously breathing children and inhalation therapy with different interfaces has become an important topic of interest among clinicians. However, clinicians usually focus on selecting the right drug-device combination and often overlook the importance of interface selection that lead to suboptimal drug delivery and therapeutic response in neonates and pediatrics. This paper provides a critical assessment of drug delivery interfaces used for the treatment of children with pulmonary diseases by emphasizing advantages and problems associated with their use during inhalation therapy.

Ari A. Drug delivery interfaces: A way to optimize inhalation therapy in spontaneously breathing children. *World J Clin Pediatr* 2016; In press

**INTRODUCTION**

There are several different types of drug delivery interfaces available on the market. Using the right interface for aerosol drug delivery to children is essential for effective inhalation therapy. However, clinicians usually focused on selecting the right drug-device combination and often overlooked the importance of interface selection that lead to suboptimal drug delivery and therapeutic response in neonates and pediatrics[1-6]. Therefore, it is necessary to critically assess each interface and understand its advantage and disadvantages in aerosol drug delivery to neonates and pediatrics. The purpose of this paper is to provide a critical assessment of drug delivery interfaces used for the treatment of children with pulmonary diseases by emphasizing advantages and problems associated with their use during inhalation therapy.

**BLOW-BY**

Blow-by is a technique that is used with a jet nebulizer placed within a distance from the child and directs aerosol plume towards the patient’s face. Historically, aerosolized medications were delivered to neonates and pediatrics using blow-by because it was considered to be an effective technique especially for crying, fussing and uncooperative children. Also, many parents preferred to use blow-by, a mask-free aerosol delivery technique, to avoid struggling with their children during inhalation therapy.

However, there are several disadvantages of this technique. For instance, it cannot be used with pressurized metered-dose inhalers (pMDIs) with valved holding chambers (VHCs) and breath-actuated nebulizers due to poor mask seal that will inhibit valve opening[7]. Also, blow-by cannot be used with mesh nebulizers due to lack of supplemental gas flow[7]. Previous research reported that blow-by is not efficient in aerosol drug delivery to children because it results in 50%-85% lower dose than the facemask[8-11]. Therefore, using blow-by for aerosol therapy is not recommended[7,11-13].

Problems associated with blow-by highlight not only the importance of interface selection in inhalation therapy, but also finding a better alternative for delivering aerosolized medications to neonates and pediatrics. Mouthpiece, facemask, nasal mask, pasifier mask, hood, high flow nasal cannula and VHCs may be viable choices of interface in children and the following sections will describe each interface more in detail.

**MOUTHPIECE**

Through an *in vitro* study, Ari *et al*[14] showed that aerosol delivery via a mouthpiece may provide twice as much drug compared with a facemask and is the most effective interface in spontaneously breathing older pediatrics[15]. Since children less than 3 years of age cannot keep the mouthpiece in their mouth with an adequate seal during inhalation therapy, the mouthpiece is not the right interface for them[16-19]. Therefore, when a mouthpiece cannot be used by a child, choosing another interface such as facemask, high flow nasal cannula or hood is important to improve the efficiency and efficacy of aerosol drug delivery to neonates and pediatrics.

**FACEMASK**

Facemasks are commonly used for aerosol drug delivery to children until they develop sufficient understanding to inhale through the mouthpiece during inhalation therapy. In children who cannot use a mouthpiece until 3 years of age, clinicians should consider using a well-fitting facemask. Therefore, it is essential to select a lightweight and flexible facemask with anatomic contours and small dead space in order to increase tolerability of facemask by children during inhalation therapy[20,21]. Using smaller masks with less dead space in neonates will lead to a greater inhaled dose especially with use of aerosol devices such as mesh nebulizers or pMDIs that do not add gas to the system during treatment.

Facemasks designs can be divided into two categories: (1) Front-loaded facemasks and (2) Bottom-loaded facemasks. Front-loaded facemasks have small entrainment ports on the side of the mask and direct aerosol toward the oronasal area of the patient as opposed to bottom-loaded masks that direct aerosol toward the upper part of the mask. Previous research reported that aerosol deposition with the front-loaded facemask (Bubbles Fish II Mask, PARI, Midlothian, Virginia) was greater than bottom-loaded facemask[8,22-24]. They also have lower deposition in the eye and face compared with bottom-loaded facemask designs[22,23,25].

When a facemask is used for aerosol drug delivery to neonates or pediatrics, clinicians should have a good face-mask seal to maximize the efficiency of treatment and prevent the drug from getting to the eyes and the face of children. However, keeping a good face-mask seal during inhalation therapy is frequently associated with crying and rejection of the facemask. Previous research showed that aerosol drug delivery to children will decrease significantly without an optimum face-mask seal because of leaks, crying or children intolerance of the facemask[2-4,22,25-29]. Janssens *et al*[30] suggested that administration of inhaled medications while children are asleep may be a viable option for inhalation therapy because children have more regular breathing patterns during sleep that may lead to greater lung deposition and better patient outcomes. However, Esposito *et al*[31] reported that 69% of the young children woke up and 75% of them distressed during inhalation therapy with pMDIs using VHCs.

In the past, clinicians believed that crying improves aerosol drug delivery to children because of the large breath at the end of the cry. However, crying results in a very long exhalation followed by fast and short inhalation that leads to deposition of aerosolized medications in the upper respiratory track than in the lower respiratory therapy track. Also, it is difficult to have a good seal with the facemask when a baby cries. Using a facemask with the pMDI - VHC, Tal *et al*[32] found that lung deposition of babies crying was 0.35% as opposed to 2% when they have quite breating. Similarly, Murakami *et al*[33] showed that aerosol deposition in a crying infant using a facemask with a nebulizer was negligible and Iles *et al*[34] reported a 4-fold decrease in lung deposition when infants were crying. According to the findings of the study conducted by Wildhaber *et al*[35] the gastrointestinal deposition in crying children was 50% higher than their non-crying peers.

**PACIFIER MASK**

As a new and innovative development of chidren-oriented drug delivery interface, the pacifier mask (Soother Mask, InspiRx, Somerset, New Jersey) was designed to achieve therapeutic lung deposition in children by eliminating their discomfort, fear and cry with the conventional facemask and keeping them calm through a pacifier. It includes the infant’s own pacifier that is attached to the anterior wall of the mask (Figure 1). The infant keeps the Soother mask sealed to his face by sucking the pacifier during treatment while nasally inhaling aerosolized medications generated by pMDIs/VHCs or nebulizers during inhalation therapy[36,37]. Amirav *et al*[38] compared the Soother mask with a conventional bottom-loaded face mask on bronchodilator delivery in 12 infants less than 1 year of age. Using scintigraphic measurements of aerosol deposition in infants, they reported that lung deposition with the Soother Mask was similar to that with the conventional face mask without a pacifier[38]. Since sucking calms children, the Shooter Mask can be used for prolonged periods of time without rejection by infants and improves compliance to aerosol treatments in infants[18,36-38].

**HIGH FLOW NASAL CANNULA**

Infants and young children are nose breathers. Since previous research showed that nasal delivery of aerosolized medications to the lungs of infants and pediatrics is superior or more effective than oral delivery[39,40], aerosol delivery through high flow nasal cannula (HFNC) has become a popular procedure in the treatment of children with pulmonary diseases. Several *in vitro* studies evaluated aerosol drug delivery through HFNC in infants and pediatrics[41-44]. Using dose quantification with the laser diffraction technique, Bhashyam *et al*[43] determined the efficiency of inhalation therapy through adult and pediatric HFNC with a mesh nebulizer placed downstream of a heated humidifier. They reported that aerosolized medications could be efficiently delivered to pediatrics through HFNC. Ari *et al*[44] compared aerosol drug delivery with helium-oxygen mixture (heliox) and oxygen at 3 L/min and 6 L/min, using a pediatric HFNC with a mesh nebulizer placed on the inspiratory inlet of a heated humidification system. They reported that bronchodilator delivery with heliox at 3 L/min was similar to that with oxygen whereas heliox delivered 2 fold greater aerosol than oxygen at 6 L/min. Sunbul *et al*[42] evaluated bronchodilator delivery using HFNC, bubble continuous positive airway pressure (CPAP) and sigh intermittent mandatory ventilation (SiPAP) with a mesh nebulizer placed proximal to the patient interface and prior to the humidifier. Using spontaneously breathing lung model attached to a low-birth-weight anatomic nasal airway cast, they showed that aerosol delivery with SiPAP was lower than HFNC and the Bubble CPAP. Aerosol deposition through HFNC was less than 2% but higher than drug delivery with the Bubble CPAP. Also, nebulizer placement at the humidifier resulted in greater aerosol deposition in HFNC, SiPAP and Bubble CPAP[42]. According to Perry *et al*[41] HFNC should not be used for bronchodilator delivery to children because the amount of aerosol deposition obtained with different cannula sizes of flows used with HFNC was lower than the amount needed for a clinical response. Also, skin irritation and condensate accumulating in the cannula are potential issues with HFNC. Therefore, clinical studies evaluating the safety and efficacy of aerosol drug delivery with HFNC are warranted.

**NASAL MASK**

Nasal masks were developed in recent years to improve aerosol drug delivery to neonates and pediatrics. The nasal mask is a special type of mask that is placed over the nasal airway during inhalation therapy. A recent *in vitro* study showed that aerosol delivery with the nasal mask was less than that with the facemask in simulated spontaneously breathing infants and young children using a jet nebulizer[24].

**HOOD**

Hood is a good option for aerosol drug delivery to children who cannot use a mouthpiece and tolerate the facemask[18,45-48]. Since there is no attachment to the patient’s face, the likelihood of agitating infants and making them cry with the use of hood for inhalation therapy may be less than facemasks. Aerosol drug delivery via hood is easy to operate and often provided when infants are asleep. Amirav *et al*[49] showed that bronchodilator delivery with the hood and facemask was similar (2.6% and 2.4%, respectively) in 14 wheezing children. Kugelman *et al*[47] reported that both treatment time and discomfort were lower in infants using the hood. In another study, Amirav and colleagues found that respiratory scores of infants with bronchiolitis received aerosol therapy with the hood and facemask were similar, but parents preferred the hood over the masks[48]. It is also important to ensure the optimal position of the child within the hood. Kim *et al*[50] found similar lung deposition in face-up and face down positions during hood nebulization; however, the face-side position has less facial-ocular deposition than face-up position.

**VALVED-HOLDING CHAMBERS (VHCS)**

VHCs are commonly used with pMDIs in order to decrease oropharangeal deposition and minimize hand-breath coordination in children[12,51]. According to previous research, spacers and VHCs should be washed with detergent and air-dry to eliminate static charge and improve aerosol delivery to infants and pediatrics[52-55]. Thus, deposition of drug particles on the inner surface of the spacer or VHC will be eliminated. Another alternative would be to use anti-static spacers/VHCs during inhalation therapy in children[56].

Also, infants and toddlers may not empty aerosolized medication from a large volume spacer of 200-700 mL. Therefore, it is important to use small volume spacers or VHCs so that the concentration of aerosol in the VHC is kept higher and children can inhale all the medication in less time with fewer breaths. Parents need to be educated to actuate one dose at a time into VHC instead of multiple doses and let their children inhale from VHC right after the pMDI has been actuated[12,57].

**EDUCATING PARENTS ABOUT INTERFACES USED IN INHALATION THERAPY**

Typically, inhaled medications are prescribed without demonstrating parents how inhalation therapy should be undertaken with each device and interface. Therefore, parents don’t know how to use each interface and problems that may arise during aerosol drug delivery to children. For instance, when their baby fights with the facemask, some parents may decide to use blow-by without knowing that it will reduce the efficiency of therapy and others force the baby to accept the facemask by holding it tightly on the baby’s face and believing that crying improves aerosol drug delivery to their children. As a result, parents report poor response to inhalation therapy to their physicians who usually decide to increase the dose or change the inhaled agent as they assume parents’ technique in aerosol drug delivery is adequate[18]. Therefore, parental awareness and training on proper technique with each interface during inhalation therapy is essential. Table 1 includes descriptions, advantages and disadvantages of each interface used for aerosol drug delivery to spontaneously breathing neonates and pediatrics. After careful instructions on how to use and handle an aerosol device, clinicians should reinforce instructions on a regular basis and the choice of drug delivery interface should be re-assessed[58].

In conclusion, delivering aerosolized drugs through different interfaces to children poses a number of challenges. Clearly, there is a need to develop more acceptable and child-friendly interfaces in order to improve aerosol drug delivery to this patient population. New interfaces should take into account the special needs and respiratory characteristics of children. Meanwhile, educating parents and healthcare professionals about drug delivery interfaces used in inhalation therapy is essential for the well-being of neonates and pediatrics.

**REFERENCES**

1 **Nikander K**, Berg E, Smaldone GC. Jet nebulizers versus pressurized metered dose inhalers with valved holding chambers: effects of the facemask on aerosol delivery. *J Aerosol Med* 2007; **20** Suppl 1: S46-55; discussion S55-8 [PMID: 17411405 DOI: 10.1089/jam.2007.0588]

2 **Janssens HM**, Tiddens HA. Facemasks and aerosol delivery by metered dose inhaler-valved holding chamber in young children: a tight seal makes the difference. *J Aerosol Med* 2007; **20** Suppl 1: S59-63; discussion S63-5 [PMID: 17411407 DOI: 10.1089/jam.2007.0578]

3 **Esposito-Festen J**, Ates B, van Vliet F, Hop W, Tiddens H. Aerosol delivery to young children by pMDI-spacer: is facemask design important? *Pediatr Allergy Immunol* 2005; **16**: 348-353 [PMID: 15943599 DOI: 10.1111/j.1399-3038.2005.00285]

4 **Erzinger S**, Schueepp KG, Brooks-Wildhaber J, Devadason SG, Wildhaber JH. Facemasks and aerosol delivery in vivo. *J Aerosol Med* 2007; **20** Suppl 1: S78-83; discussion S83-4 [PMID: 17411409 DOI: 10.1089/jam.2007.0572]

5 **Ari A**, Fink JB. Effective bronchodilator resuscitation of children in the emergency room: device or interface? *Respir Care* 2011; **56**: 882-885 [PMID: 21679497 DOI: 10.4187/respcare.01375]

6 **Ari A**, Hess D, Myers TR, Rau JL. A Guide to Aerosol Delivery Devices for Respiratory Therapists. Dallas, Texas: American Association for Respiratory Care, 2009

7 **DiBlasi RM**. Clinical Controversies in Aerosol Therapy for Infants and Children. *Respir Care* 2015; **60**: 894-914; discussion 914-6 [PMID: 26070582 DOI: 10.4187/respcare.04137]

8 **Lin HL**, Restrepo RD, Gardenhire DS, Rau JL. Effect of face mask design on inhaled mass of nebulized albuterol, using a pediatric breathing model. *Respir Care* 2007; **52**: 1021-1026 [PMID: 17650358]

9 **Restrepo RD**, Dickson SK, Rau JL, Gardenhire DS. An investigation of nebulized bronchodilator delivery using a pediatric lung model of spontaneous breathing. *Respir Care* 2006; **51**: 56-61 [PMID: 16381619]

10 **Rubin BK**. Bye-bye, blow-by. *Respir Care* 2007; **52**: 981 [PMID: 17650350]

11 **Ari A**, Restrepo RD. Aerosol delivery device selection for spontaneously breathing patients: 2012. *Respir Care* 2012; **57**: 613-626 [PMID: 22472501 DOI: 10.4187/respcare.01756]

12 **Ari A**, Fink JB. Aerosol therapy in children: challenges and solutions. *Expert Rev Respir Med* 2013; **7**: 665-672 [PMID: 24224509 DOI: 10.1586/17476348.2013.847369]

13 **Ari A**, Fink JB. Guidelines for aerosol devices in infants, children and adults: which to choose, why and how to achieve effective aerosol therapy. *Expert Rev Respir Med* 2011; **5**: 561-572 [PMID: 21859275 DOI: 10.1586/ers.11.49]

14 **Ari A**, de Andrade AD, Sheard M, AlHamad B, Fink JB. Performance Comparisons of Jet and Mesh Nebulizers Using Different Interfaces in Simulated Spontaneously Breathing Adults and Children. *J Aerosol Med Pulm Drug Deliv* 2015; **28**: 281-289 [PMID: 25493535 DOI: 10.1089/jamp.2014]

15 **Ditcham W**, Murdzoska J, Zhang G, Roller C, von Hollen D, Nikander K, Devadason SG. Lung deposition of 99mTc-radiolabeled albuterol delivered through a pressurized metered dose inhaler and spacer with facemask or mouthpiece in children with asthma. *J Aerosol Med Pulm Drug Deliv* 2014; **27** Suppl 1: S63-S75 [PMID: 25054483 DOI: 10.1089/jamp.2014.1139]

16 **Everard ML**. Aerosol delivery to children. *Pediatr Ann* 2006; **35**: 630-636 [PMID: 16999296 DOI: 10.3928/0090-4481-20060901-06]

17 **Everard ML**. Inhalation therapy for infants. *Adv Drug Deliv Rev* 2003; **55**: 869-878 [PMID: 12842605 DOI: 10.1016/S0169-409X(03)00082-6]

18 **Amirav I**, Newhouse MT. Aerosol therapy in infants and toddlers: past, present and future. *Expert Rev Respir Med* 2008; **2**: 597-605 [PMID: 20477295 DOI: 10.1586/17476348.2.5.597]

19 **Devadason SG**. Recent advances in aerosol therapy for children with asthma. *J Aerosol Med* 2006; **19**: 61-66 [PMID: 16551216 DOI: 10.1089/jam.2006.19.61]

20 **Amirav I**, Mandelberg A. Face masks for aerosols-there is more science... *Pediatr Pulmonol* 2010; **45**: 221-223 [PMID: 20146372 DOI: 10.1002/ppul.21163]

21 **Amirav I**, Newhouse MT. Review of optimal characteristics of face-masks for valved-holding chambers (VHCs). *Pediatr Pulmonol* 2008; **43**: 268-274 [PMID: 18219694 DOI: 10.1002/ppul.20767]

22 **Smaldone GC**, Berg E, Nikander K. Variation in pediatric aerosol delivery: importance of facemask. *J Aerosol Med* 2005; **18**: 354-363 [PMID: 16181009 DOI: 10.1089/jam.2005.18.354]

23 **Sangwan S**, Gurses BK, Smaldone GC. Facemasks and facial deposition of aerosols. *Pediatr Pulmonol* 2004; **37**: 447-452 [PMID: 15095329 DOI: 10.1002/ppul.10454]

24 **El Taoum KK**, Xi J, Kim J, Berlinski A. In Vitro Evaluation of Aerosols Delivered via the Nasal Route. *Respir Care* 2015; **60**: 1015-1025 [PMID: 25587167 DOI: 10.4187/respcare.03606]

25 **Smaldone GC**, Sangwan S, Shah A. Facemask design, facial deposition, and delivered dose of nebulized aerosols. *J Aerosol Med* 2007; **20** Suppl 1: S66-75; discussion S75-7 [PMID: 17411408 DOI: 10.1089/jam.2007.0579]

26 **Amirav I**, Newhouse MT. Aerosol therapy with valved holding chambers in young children: importance of the facemask seal. *Pediatrics* 2001; **108**: 389-394 [PMID: 11483804 DOI: 10.1542/peds.108.2.389]

27 **Esposito-Festen JE**, Ates B, van Vliet FJ, Verbraak AF, de Jongste JC, Tiddens HA. Effect of a facemask leak on aerosol delivery from a pMDI-spacer system. *J Aerosol Med* 2004; **17**: 1-6 [PMID: 15120007 DOI: 10.1089/089426804322994406]

28 **Smaldone GC**. Assessing new technologies: patient-device interactions and deposition. *Respir Care* 2005; **50**: 1151-1160 [PMID: 16122399]

29 **Hayden JT**, Smith N, Woolf DA, Barry PW, O'Callaghan C. A randomised crossover trial of facemask efficacy. *Arch Dis Child* 2004; **89**: 72-73 [PMID: 14709514]

30 **Janssens HM**, van der Wiel EC, Verbraak AF, de Jongste JC, Merkus PJ, Tiddens HA. Aerosol therapy and the fighting toddler: is administration during sleep an alternative? *J Aerosol Med* 2003; **16**: 395-400 [PMID: 14977430 DOI: 10.1089/089426803772455659]

31 **Esposito-Festen J**, Ijsselstijn H, Hop W, van Vliet F, de Jongste J, Tiddens H. Aerosol therapy by pressured metered-dose inhaler-spacer in sleeping young children: to do or not to do? *Chest* 2006; **130**: 487-492 [PMID: 16899849 DOI: 10.1378/chest.130.2.487]

32 **Tal A**, Golan H, Grauer N, Aviram M, Albin D, Quastel MR. Deposition pattern of radiolabeled salbutamol inhaled from a metered-dose inhaler by means of a spacer with mask in young children with airway obstruction. *J Pediatr* 1996; **128**: 479-484 [PMID: 8618180 DOI: 10.1016/S0022-3476(96)70357-8]

33 **Murakami G**, Igarashi T, Adachi Y, Matsuno M, Adachi Y, Sawai M, Yoshizumi A, Okada T. Measurement of bronchial hyperreactivity in infants and preschool children using a new method. *Ann Allergy* 1990; **64**: 383-387 [PMID: 2321816]

34 **Iles R**, Lister P, Edmunds AT. Crying significantly reduces absorption of aerosolised drug in infants. *Arch Dis Child* 1999; **81**: 163-165 [PMID: 10490528 DOI: 10.1136/adc.81.2.163]

35 **Wildhaber JH**, Dore ND, Wilson JM, Devadason SG, LeSouëf PN. Inhalation therapy in asthma: nebulizer or pressurized metered-dose inhaler with holding chamber? In vivo comparison of lung deposition in children. *J Pediatr* 1999; **135**: 28-33 [PMID: 10393600 DOI: 10.1016/S0022-3476(99)70323-9]

36 **Amirav I**, Newhouse MT, Luder A, Halamish A, Omar H, Gorenberg M. Feasibility of aerosol drug delivery to sleeping infants: a prospective observational study. *BMJ Open* 2014; **4**: e004124 [PMID: 24670428 DOI: 10.1136/bmjopen-2013-004124]

37 **Amirav I**, Luder AS, Halamish A, Raviv D, Kimmel R, Waisman D, Newhouse MT. Design of aerosol face masks for children using computerized 3D face analysis. *J Aerosol Med Pulm Drug Deliv* 2014; **27**: 272-278 [PMID: 24074142 DOI: 10.1089/jamp.2013.1069]

38 **Amirav I**, Luder A, Chleechel A, Newhouse MT, Gorenberg M. Lung aerosol deposition in suckling infants. *Arch Dis Child* 2012; **97**: 497-501 [PMID: 22362720 DOI: 10.1136/archdischild-2011-301236]

39 **Chua HL**, Collis GG, Newbury AM, Chan K, Bower GD, Sly PD, Le Souef PN. The influence of age on aerosol deposition in children with cystic fibrosis. *Eur Respir J* 1994; **7**: 2185-2191 [PMID: 7713202 DOI: 10.1183/09031936.94.07122185]

40 **Amirav I**, Borojeni AA, Halamish A, Newhouse MT, Golshahi L. Nasal versus oral aerosol delivery to the "lungs" in infants and toddlers. *Pediatr Pulmonol* 2014; Epub ahead of print [PMID: 24482309 DOI: 10.1002/ppul.22999]

41 **Perry SA**, Kesser KC, Geller DE, Selhorst DM, Rendle JK, Hertzog JH. Influences of cannula size and flow rate on aerosol drug delivery through the Vapotherm humidified high-flow nasal cannula system. *Pediatr Crit Care Med* 2013; **14**: e250-e256 [PMID: 23628834 DOI: 10.1097/PCC.0b013e31828a7f79]

42 **Sunbul FS**, Fink JB, Harwood R, Sheard MM, Zimmerman RD, Ari A. Comparison of HFNC, bubble CPAP and SiPAP on aerosol delivery in neonates: An in-vitro study. *Pediatr Pulmonol* 2015; **50**: 1099-1106 [PMID: 25491434 DOI: 10.1002/ppul.23123]

43 **Bhashyam AR**, Wolf MT, Marcinkowski AL, Saville A, Thomas K, Carcillo JA, Corcoran TE. Aerosol delivery through nasal cannulas: an in vitro study. *J Aerosol Med Pulm Drug Deliv* 2008; **21**: 181-188 [PMID: 18518794 DOI: 10.1089/jamp.2007.0662]

44 **Ari A**, Harwood R, Sheard M, Dailey P, Fink JB. In vitro comparison of heliox and oxygen in aerosol delivery using pediatric high flow nasal cannula. *Pediatr Pulmonol* 2011; **46**: 795-801 [PMID: 21438178 DOI: 10.1002/ppul.21421]

45 **Amirav I**, Shakked T, Broday DM, Katoshevski D. Numerical investigation of aerosol deposition at the eyes when using a hood inhaler for infants--a 3D simulation. *J Aerosol Med Pulm Drug Deliv* 2008; **21**: 207-214 [PMID: 18518796 DOI: 10.1089/jamp.2007.0619]

46 **Shakked T**, Broday DM, Katoshevski D, Amirav I. Administration of aerosolized drugs to infants by a hood: a three-dimensional numerical study. *J Aerosol Med* 2006; **19**: 533-542 [PMID: 17196081 DOI: 10.1089/jam.2006.19.533]

47 **Kugelman A**, Amirav I, Mor F, Riskin A, Bader D. Hood versus mask nebulization in infants with evolving bronchopulmonary dysplasia in the neonatal intensive care unit. *J Perinatol* 2006; **26**: 31-36 [PMID: 16341026 DOI: 10.1038/sj.jp.7211434]

48 **Amirav I**, Oron A, Tal G, Cesar K, Ballin A, Houri S, Naugolny L, Mandelberg A. Aerosol delivery in respiratory syncytial virus bronchiolitis: hood or face mask? *J Pediatr* 2005; **147**: 627-631 [PMID: 16291353 DOI: 10.1016/j.jpeds.2005.05.035]

49 **Amirav I**, Balanov I, Gorenberg M, Groshar D, Luder AS. Nebuliser hood compared to mask in wheezy infants: aerosol therapy without tears! *Arch Dis Child* 2003; **88**: 719-723 [PMID: 12876173 DOI: 10.1136/adc.88.8.719]

50 **Kim J**, Xi J, Si X, Berlinski A, Su WC. Hood nebulization: effects of head direction and breathing mode on particle inhalability and deposition in a 7-month-old infant model. *J Aerosol Med Pulm Drug Deliv* 2014; **27**: 209-218 [PMID: 23808762 DOI: 10.1089/jamp.2013.1051]

51 **Muchão FP**, Perín SL, Rodrigues JC, Leone C, Silva Filho LV. [Evaluation of the knowledge of health professionals at a pediatric hospital regarding the use of metered-dose inhalers]. *J Bras Pneumol* 2008; **34**: 4-12 [PMID: 18278370 DOI: S1806-37132008000100003]

52 **Wildhaber JH**, Janssens HM, Piérart F, Dore ND, Devadason SG, LeSouëf PN. High-percentage lung delivery in children from detergent-treated spacers. *Pediatr Pulmonol* 2000; **29**: 389-393 [PMID: 10790251 DOI: 10.1002/(SICI)1099-0496]

53 **Piérart F**, Wildhaber JH, Vrancken I, Devadason SG, Le Souëf PN. Washing plastic spacers in household detergent reduces electrostatic charge and greatly improves delivery. *Eur Respir J* 1999; **13**: 673-678 [PMID: 10232445 DOI: 10.1183/09031936.99.13367399]

54 **Dompeling E**, Oudesluys-Murphy AM, Janssens HM, Hop W, Brinkman JG, Sukhai RN, de Jongste JC. Randomised controlled study of clinical efficacy of spacer therapy in asthma with regard to electrostatic charge. *Arch Dis Child* 2001; **84**: 178-182 [PMID: 11159302 DOI: 10.1136/adc.84.2.178]

55 **Anhøj J**, Bisgaard H, Lipworth BJ. Effect of electrostatic charge in plastic spacers on the lung delivery of HFA-salbutamol in children. *Br J Clin Pharmacol* 1999; **47**: 333-336 [PMID: 10215759 DOI: 10.1046/j.1365-2125.1999.00893]

56 **Bisgaard H**, Anhøj J, Klug B, Berg E. A non-electrostatic spacer for aerosol delivery. *Arch Dis Child* 1995; **73**: 226-230 [PMID: 7492160 DOI: 10.1136/adc.73.3.226]

57 **Wildhaber JH**, Devadason SG, Eber E, Hayden MJ, Everard ML, Summers QA, LeSouëf PN. Effect of electrostatic charge, flow, delay and multiple actuations on the in vitro delivery of salbutamol from different small volume spacers for infants. *Thorax* 1996; **51**: 985-988 [PMID: 8977597 DOI: 10.1136/thx.51.10.985]

58 **Lannefors L**. Inhalation therapy: Practical considerations for nebulisation therapy. *Physical Therapy Review* 2006; **11**: 21-27 [DOI: 10.1179/108331906X98976]

**P-Reviewer:** Sandblom G, Shelat VG **S-Editor:** Qiu S **L-Editor: E-Editor:**

**Table 1 Descriptions, advantages and disadvantages of each interface used for aerosol drug delivery to spontaneously breathing neonates and pediatrics**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Interface** | **Description** | **Advantages** | **Disadvantages** | **Suggestions for the best practice** |
| Blow-by | A technique that directs aerosol plume towards the patient’s face by placing a jet nebulizer within a distance from the child that ranges from 1 to 30 cm | 1Easy to use  1Comfortable and easy to tolerate by the patient  1A mask-free aerosol delivery technique  1Used with fussing, crying and uncooperative children | 1Inefficient aerosol drug delivery to children  1Drug delivery with blow-by is 50%-85% less than the facemask.  1Cannot be used with pMDIs, breath- actuated nebulizers and mesh nebulizers | 1Inhalation therapy with blow-by is not efficient; therefore, it should not be used for aerosol drug delivery to neonates and pediatrics |
| Mouthpiece | A cylindrical tube that extends between the lips so that aerosol can pass through the oropharynx to reach lower respiratory tract | 1Efficient inhalation therapy in children  1Aerosol drug delivery with a mouthpiece is two-fold more than that with a face mask | 1Children less than 3 years of age cannot use a mouthpiece.  1An adequate consistent seal is needed during inhalation therapy | 1The mouthpiece should not be used for children who are less than 3 years old.  1When using a mouthpiece child should be encouraged to keep it in their mouth during therapy  1If a child cannot keep the mouthpiece in his mouth with an adequate seal during aerosol drug delivery, another interface should be used for inhalation therapy |
| Facemask | An interface that covers the nose and mouth. It is kept in place through an elastic band that extends beyond the back of the head or neck | 1Can be used in children all years of age.  1Can be used with nebulizers and pMDIs to deliver aerosolized medications to neonates and pediatrics | 1A good face-mask seal is needed for optimum aerosol drug delivery.  1Is frequently associated with crying, intolerance and rejection of the mask.  1Crying and leaks between face and mask decrease aerosol drug delivery to children | 1Select a lightweight and flexible facemask with anatomic contours to increase tolerability of face mask by children during therapy  1Choose a facemask with small dead space and have a good face-mask seal to increase delivery efficiency of inhalation therapy.  1Use another interface if the patient starts to fuss, and cry during aerosol drug delivery with a facemask |
| Pacifier Mask | A face mask with the attachment of the infant’s own pacifier | 1A new and innovative facemask design that eliminates fear, discomfort and cry with the standard facemask  1A children-oriented drug delivery interface designed to achieve therapeutic lung deposition in children  1Improves compliance to inhalation therapy in infants |  | 1May be a good option for children who fuss, cry and does not tolerate other interfaces used for aerosol drug delivery in neonates and pediatrics |
| Nasal Mask | An interface that covers the nose to allow aerosol to pass through the nasopharynx to reach the lower respiratory tract | 1Easy to use  1Better tolerance than the facemask | 1Aerosol delivery with the nasal mask is less than that with the standard facemask |  |
| High Flow Nasal Cannula | A tubing with two small prongs that are inserted into the nares to allow aerosol pass through the nasopharynx and reach the lower respiratory tract | 1Efficient delivery of aerosolized medications to neonates and pediatrics.  1Children may tolerate HFNC better than the facemask | 1More information about the safety and efficacy of aerosol drug delivery though HFNC is needed.  1Cannot be used with pMDIs | 1When using mesh nebulizers for aerosol drug delivery to neonates and pediatrics, place the nebulizer prior to the heated humidifier |
| Hood | An enclosure that covers the head and neck of a neonate or small children to deliver aerosol to the lungs while isolating it from ambient air | 1A good option for aerosol delivery to children who cannot use a mouthpiece and tolerate the facemask.  1Likelihood of agitating infants and making them cry is low.  1Aerosol delivery with the hood is the same as the facemask.  1Parents prefer the hood over the mask | 1User may need additional training and practice to provide proper inhalation therapy with the hood.  1More time and parts may be needed for the set-up | 1 Use the hood for aerosol drug delivery to children who cannot use a mouthpiece and tolerate the facemask.  1Put the infant in the face-side position when using the hood for inhalation therapy because it has less facial-ocular deposition than face-up position |
| Valved Holding Chamber | A chamber shaped interface with a one-way valve that allows aerosols to be contained in the chamber during aerosol therapy | 1Reduces oropharangeal deposition  1Minimize hand-breath coordination during inhalation therapy  1Improves efficiency of aerosol therapy | 1Electrostatic charge and large volume VHCs result in a decrease in aerosol drug delivery to children | 1Wash the VHC with detergent and air dry before inhalation therapy in order to eliminate static charge and improve aerosol delivery to neonates and pediatrics.  1Choose small volume VHCs for aerosol therapy.  1Actuate one-dose at a time into VHC instead of multiple doses |

VHC: Valved holding chambers; pMDIs: Pressurized metered-dose inhalers.



**Figure 1 Soother Mask (Reproduced with permission from the InspiRx, Somerset, New Jersey).**