**Name of Journal: *World Journal of Gastroenterology***

**ESPS Manuscript NO: 32652**

**Manuscript Type: MINIREVIEWS**

**Magnetic anchor guidance for endoscopic submucosal dissection and other endoscopic procedures**

Bhatt A *et al.* Magnetic anchor guidance and ESD

Mohamed Mortagy, Neal Mehta, Mansour A Parsi, Seiichiro Abe, Tyler Stevens, John J Vargo, Yutaka Saito, Amit Bhatt

**Mohamed Mortagy,** Department of Internal Medicine, Homer Stryker MD School of Medicine, Western Michigan University, Kalamazoo, MI 49008, United States

**Neal Mehta,** Department of Internal Medicine, Medicine Institute, Cleveland Clinic, Cleveland, OH 44195, United States

**Mansour A Parsi, Tyler Stevens, John J Vargo, Amit Bhatt,** Department of Gastroenterology and Hepatology, Digestive Disease Institute, Cleveland Clinic, Cleveland, OH 44195, United States

**Seiichiro Abe, Yutaka Saito,** Endoscopy Division, National Cancer Center Hospital, Tokyo 104-0045, Japan

**Author contributions:** All authors contributed equally to this work. All authors played a role in concept and design; acquisition of data; interpretation of data; drafting of the manuscript and critical revision of the manuscript for important intellectual content.

**Conflict-of-interest statement:** All authors have no conflicts-of-interest

**Open-Access:** This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: <http://creativecommons.org/licenses/by-nc/4.0/>

**Manuscript source:** Unsolicited manuscript

**Correspondence to:**  **Amit Bhatt, MD,** Department of Gastroenterology and Hepatology, Digestive Disease Institute, A30, The Cleveland Clinic, 9500 Euclid Ave, Cleveland, OH 44195, United States. [bhatta3@ccf.org](mailto:bhatta3@ccf.org)

**Telephone:** +1-216-4444880

**Fax:** +1-216-4446305

**Received:** January 18, 2017

**Peer-Review Started:** January 19, 2017

**First Decision:** February 10, 2017

**Revised:** February 24, 2017

**Accepted:** March 15, 2017

**Article in Press:**

**Published Online:**

**Abstract**

Endoscopic submucosal dissection (ESD) is a well-established, minimally invasive treatment for superficial neoplasms of the gastrointestinal tract. The universal adoption of ESD has been limited by its slow learning curve, long procedure times, and high risk of complications. One technical challenge is the lack of a second hand that can provide traction, as in conventional surgery. Reliable tissue retraction that exposes the submucosal plane of dissection would allow for safer and more efficient dissection. Magnetic anchor guided endoscopic submucosal dissection (MAG-ESD) has potential benefits compared to other current traction methods. MAG-ESD offers *dynamic* tissue retraction independent of the endoscope mimicking a surgeon’s “second hand”. Two types of magnets can be used: electromagnets and permanent magnets. In this article we review the MAG-ESD technology, published work and studies of magnets in ESD. We also review the use of magnetic anchor guidance systems (MAGS) in natural orifice transluminal endoscopic surgery and the idea of magnetic non-contact retraction using surface ferromagentization. We discuss the current limitations, the future potential of MAG-ESD and the developments needed for adoption of this technology.

**Key words:** Endoscopic submucosal dissection; Magnetic anchor guidance; Gastric cancer; Esophageal cancer; Magnets; Traction; Natural orifice transluminal endoscopic surgery

© **The Author(s) 2017.** Published by Baishideng Publishing Group Inc. All rights reserved

**Core tip:** Magnetic anchor guided endoscopic submucosal dissection (MAG-ESD) is a promising technology that can offer tissue retraction independent of the endoscope which acts like an endoscopist’s “second hand”. The objective of this article is to present a scientific review of the technology, published work, limitations, and future potential of MAG-ESD.

Mortagy M, Mehta N, Parsi MA, Abe S, Stevens T, Vargo JJ, Saito Y, Bhatt A. Magnetic anchor guidance for endoscopic submucosal dissection and other endoscopic procedures.*World J Gastroenterol* 2017; In press

**INTRODUCTION**

Endoscopic submucosal dissection (ESD) is an advanced endoscopic technique that allows for *en-bloc* resection of early gastrointestinal tract cancers[1-3]. By achieving *en-bloc* resection, ESD has shown higher curative resection rates and lower recurrence rates than endoscopic mucosal resection (EMR)[4].

ESD was developed in Japan and has become standard of care for managing early cancers and tumors of the esophagus, stomach and colon. Its adoption in the Western world has been slow due to lower volume of gastric cancer, longer procedure times, and slow learning curve[5-7]. ESD also carries a significant risk of complications. The risk of immediate and delayed bleeding in gastric ESD is 7% and 5.5% respectively[8]. The risk of bleeding in colorectal ESD is 1.5%[1] and is rare in esophageal ESD[9].

During conventional surgery, dissection is performed with one hand or device cutting and a second hand or device retracting the tissue so that the appropriate plane of dissection is easily seen. In ESD, the entire procedure is done through a single endoscope, and there is no second hand to help as in conventional surgery[10-12]. This limits visualization of the dissection plane increasing the technical challenge of the procedure (Figure 1). Adequate tissue tension and clear visibility of the dissection plane are important for effective and safe dissection[12-14]. Inability to visualize the dissection plane can lead to inadvertent cutting of blood vessels resulting in bleeding, and difficulty in stopping bleeding as the bleeding point is poorly visualized. It has been estimated that approximately 70% of procedure time during ESD is spent handling bleeding by inexperienced operators[15]. If the submucosal plane of dissection is not well visualized, inadvertent cutting of the muscle layer may lead to perforation. The risk of perforation in gastric, esophageal and colonic ESD is 1.2-5.2%, 0-6%and 2.4% respectively[16,8,1].

Several methods have been developed to help provide traction during ESD procedures[17]. Examples are the clip-with-line method[18], percutaneous traction method[19], sinker-assisted method[20], magnetic anchor method[12], external forceps method[21], internal-traction methods (clip band method[22] and medical ring method[23]), double-channel-scope method[24], outer-route method[25], double-scope method[26], endoscopic surgical platform method[27], robot-assisted method[28] and the Sakamoto and Osada clip (SO-Clip) method[29].

Magnetic anchor guided - endoscopic submucosal dissection (MAG-ESD) is an attractive method of traction as it has potential benefits compared to other current traction methods. It delivers traction independent of the endoscope by creating an invisible second hand for the operator. Unlike the double channel and outer-route methods, MAG-ESD does not interfere with the complex endoscopic movements needed to perform ESD. Unlike the clip-with-line, percutaneous traction, and sinker-assisted methods, MAG-ESD provides *dynamic* traction in a sense that the external magnet can be moved to change the direction of retraction. Permanent and electromagnets are the only two types of magnets available[30].

**ELECTROMAGNETS**

Electromagnets are made of a coil of an electric conductor which acts as a magnet only when an electric current passes through. The main advantage of electromagnets is that the magnetic field can be quickly changed by controlling the amount of electric current used. Their main disadvantage is being large and cumbersome. Another significant disadvantage of electric magnets is being electric current dependent. Without a current, no force is produced, making anchoring impossible. The only current which can be used is direct current, necessitating the need for an external power source and wired control system, because voltage and capacity of modern batteries are too low to supply the necessary energy. Moreover, every current pass generates thermal energy, generating heat close to the endoscopist.

**USE OF ELECTROMAGNETS IN MAG-ESD**

***Initial concept of MAG-ESD***

The concept of MAG-ESD was first postulated and experimented by Kobayashi *et al[*12]. They proposed that if a magnetic field is properly controlled and powered, it could generate sufficient force to use a micro-forceps for traction of the mucosa during EMR and thereby developed a magnetic anchor system to prove their theory. The magnetic anchor consisted of three parts: a handmade magnetic weight composed of magnetic stainless steel (SUS420F), a micro-forceps and a connecting thread. This system was tested in an ex-vivo porcine model. First, an incision was made by standardEMR technique in the mucosa. Then, a magnetic anchor was attached to the edge of the incised mucosa. The tip of mucosa was lifted by increasing the electric current through the electromagnet of the magnetic control system. Direction of traction for magnetic anchor could be simply manipulated by changing the position of the electromagnet over the animal or by changing the position of the animal itself. Hemorrhage was rare because blood vessels could be clearly visualized endoscopically and hemostatic procedures were conducted before cutting blood vessels. However, separation of the magnetic weight from the forceps and slipping of the micro-forceps from the mucosa were observed. In these cases, new magnetic anchors could be inserted without any problems and the malfunctioned anchors could be easily retrieved. One main drawback of this model was that the electromagnets used were large and cumbersome and not practical for most endoscopy suites.

***First human experience***

The feasibility of MAG-ESD in the resection of early gastric cancer (EGC) in humans was first looked at by Gotoda *et al*[31]. Twenty-five patients with EGC > 20 millimeters (mm) in diameter located in the gastric body were enrolled. The structure of the magnetic anchor was the same as that used in Kobayashi’s study and the magnet used was electromagnet as well. All tumors were resected *en-bloc* without any perforations or significant bleeding. All magnetic anchors were safely retrieved. The magnetic anchor system was reported to be supportive in 23 out of 25 operations by the performing endoscopists. All ESD scars were completely healed after 8 weeks. Neither recurrent cancer nor distant metastases were observed in any of the patients during follow up. The authors concluded that MAG-ESD is a feasible and safe method that enables excellent visualization of the tissue and facilitates gastric ESD in patients with early gastric cancer.

**PERMANENT MAGNETS**

Permanent magnets are made from a material that stays magnetized (ferromagnetic). Permanent magnets have the advantage of being small in size. Neodymium-iron-boron (NeFeB) rare earth magnets are the strongest permanent magnets available. They are highly resistant to demagnetization because of their atomic structure[32]. They are utilized in a variety of commercial applications including MRI machines. However, Neodymium magnets have a low resistance to corrosion, necessitating a protective coating to prevent potential release of harmful products. Thus, all neodymium magnets should be isolated from contact with tissues and fluids inside the body. The shielding material must be inert to the human body and non-obstructive to the magnetic field (*e.g.,* gold or epoxy)[30].

**USE OF PERMANENT MAGNETS IN MAG-ESD**

***Live animal trial***

The Feasibility of counter-traction with a simplified MAG-ESD technique was evaluated in a live animal experiment in two beagle dogs[33]. The external magnet was a hand-held permanent magnet that can be locked by a flexible arm, so there was no need for an assistant to keep holding it in position during the procedure (Figure 2). Ten MAG-ESDs using neodymium permanent magnets were successfully performed in all cases without perforation in the stomach. Adequate counter-traction with good visualization was obtained using the external magnet except in two cases. None of the magnetic anchors became displaced during the procedure and all were retrieved after endoscopic resection.

**STANDARD ESD *VS* MAG ESD**

The first comparative study between standard ESD and MAG-ESD was done on simulated gastric lesions in ex-vivo animal model[34]. Ten ESDs were performed using a MAG-ESD device and ten were performed by the standard ESD technique. The MAG-ESD device consisted of two neodymium magnets (one internal and one external), with sutures, and standard endoscopic clips. The magnet could be freely manipulated using a three dimensional hemisphere allowing the assistant to adjust the direction of counter-traction in accordance with the operator’s preference and providing the endoscopist with the broadest view of the dissection line. The MAG-ESD device provided dynamic multidirectional counter-traction during ESD. Endoscopists reported increased visualization in MAG-ESD compared to standard ESD. The total procedure time and the lesion dissection time for the MAG-ESD were significantly lower than the control group. No perforations or other adverse events occurred in either group. The authors concluded that this device reduced procedure time for an operator who was experienced in ESD and enhanced exposure of dissection line compared to standard ESD.

**MAGNETIC ANCHOR GUIDANCE IN NATURAL ORIFICE TRANSLUMINAL ENDOSCOPIC SURGERY**

Magnetic anchor guidance (MAG) has also been used in natural orifice transluminal endoscopic surgery (NOTES). In NOTES, operative instruments are passed through an opening created in a natural orifice such as the stomach, colon or vagina leaving the patient with no visible scars. NOTES faces similar challenges to ESD. The entire procedure is performed through a single endoscope and there is loss of triangulation and the second hand that is normally present in surgical procedures. Moreover, the instruments working area is reduced resulting in frequent collision of instruments known as “Sword fighting”.

Triangulation is the ability of the surgeon to manipulate tissue with traction and counter-traction permitting appropriate dissection and suturing. Flexible endoscopic surgery need to apply the same principles of triangulation to be an effective surgical approach[35,36]. Multiple endoscopic prototypes have been experimented in NOTES, but they failed for multiple reasons[37]. First, additional angulations at the tip of a flexible endoscope sacrificed stability which is needed for safe dissection and suturing. Second, angulation at the tip of flexible endoscope makes tactile feedback less reliable. Third, adding more functions to flexible endoscopes increases the complexity of the procedure leading to its impracticality by the operator. A single operator is only able to control a limited number of buttons and wheels. Adding more functions to the flexible endoscope allowing more angulations, more control buttons, wheels and levers will lead to practical limitations due to increased complexity. It has been concluded that current instruments are not suitable for effective surgery in the NOTES environment[38].

The magnetic anchor guidance system (MAGS) concept was tested in a laparoscopic prototype with a system of external neodymium magnetic anchors, tissue retractors and an internal camera[39]. This system aims at controlling an intra-abdominal laparoscope and multiple instruments introduced through a single trocar into the abdominal cavity. All instruments had neodymium permanent magnets. Surgeons were able to position the retractors under the liver or the spleen to elevate them by manipulating external magnetic anchors in a porcine model. The MAGS platform was sufficient to securely anchor the tissue retractors and the camera. These instruments could be moved within the abdomen without loss of magnetic coupling. Two non-survival porcine laparoscopic nephrectomies were successfully completed without complications using this system.

In another study with the same MAG platform, four non-survival porcine transvaginal NOTES cholecystectomy using magnetically anchored instruments were successfully performed[40,41]. This platform was also used successfully in the manipulation and stabilization of mesh for transcolonic endoscopic ventral hernia repair[42].

MAG has also been used in 40 patients undergoing single port laparoscopic cholecystectomy[43]. This MAG system consisted of two magnetic forceps and two external magnets. By moving the first external magnet, the first forceps grasping the gallbladder fundus was pulled upwards lifting the liver and exposing the gallbladder. By moving the second magnet, the second forceps grasping the Hartmann’s pouch was pulled towards the right lower quadrant or towards the epigastrium yielding the best exposure of Calot’s triangle. Then, the gallbladder was dissected laparoscopically. However, in 23 of the 40 patients, assistance was required using a 1-mm needle inserted through the abdominal wall to delicately separate the liver or help in the Calot triangle’s dissection.

**MAGNETIC NON-CONTACT RETRACTION BY SURFACE FERROMAGNETIZATION**

Wang *et al[*44] postulated that magnetic non-contact retraction has potential advantages over the existing direct physical retraction means such as forceps. Non-contact retraction provides complete atraumatic retraction and avoids tumor cell exfoliation. They conducted a pilot study using surface magnetic retraction of the gastric mucosa to facilitate mucosal resection in sixteen porcine stomach specimens. They prepared magnetic media of ferromagnetic micro-particles (Stainless steel 410) dispersed in cyanoacrylate liquid. The media was surface-glued to the target mucosa. A 3-mm neodymium permanent magnet was placed in contact with the media, and the mucosa was magnetically retracted by 10 mm. An in-vitro surgical resection of the retracted mucosa was then performed with the use of electrosurgical snares. The magnet was capable of maintaining retraction of the mucosa, via the medium, throughout the procedure. The data obtained showed that the force required to retract and lift the mucosa by 10 mm was 0.135 Newtons (N). There is a linear relationship between the retraction force and distance. The estimated total retracting force during cutting with Snares was 0.324 ± 0.154 N. They concluded that non-contact magnetic grasping is feasible with the development of the necessary sensor and control technology. The main limitation of this idea is that the magnetic force decreases sharply as the air gap increases. Hence, more powerful magnets would be needed for non-contact retraction at a safe operating distance.

**LIMITATIONS OF MAG-ESD**

In its current form, the most significant limitation of MAG-ESD is the coupling strength of magnets. Magnetic force decays over distance exponentially. So, the strength of magnetic force varies based on the patient’s abdominal thickness which depends mainly on the amount of abdominal fat. Experimental studies have proven that the maximal thickness of the abdominal wall where MAG-ESD is practical varies from 1.5 centimeters (cm) up to 4 cm[45,46]. So, MAG-ESD in its current form can’t be adopted for obese patients and is limited to thin and pediatric patients only[44]. More studies need to be done to be able to determine the appropriate relation between coupling force of magnets, weight of anchored instruments and abdominal wall thickness.

In non-obese women, the abdominal wall measured 1.7 ± 0.7 cm at the umbilicus and 2.0 ± 0.7 cm in the left upper quadrant. For obese women (BMI ≥ 30), thickness was 3.0 ± 1.2 cm at the umbilicus and 4.4 ± 1.0 in the left upper quadrant[47].

If we want to use more than one external magnet to control more than one intraabdominal/intraluminal tool as what is done in MAGS that is used in NOTES, we will be faced by another limitation. The maximum number of external magnets in the operating area is unknown. The minimum separation distance between external magnets in current external platforms is 3 cm before magnet attraction could be problematic. Such problems could be magnet-magnet interference and also operator hand-magnet collisions[39].

Some authors point out to the possible effects of magnetic fields on human tissues[30]. An example is the possible effect on blood flow which might cause clotting. Another problem is the interaction between magnets and other instruments used during procedures. Moreover, a ferromagnetic foreign body in human tissue can interact with the magnetic field. This technology is contraindicated with pacemakers[48], metal foreign body or recently implanted metal orthopedic prosthesis.

To test possible damages that might be caused by magnets on tissues, a study was conducted on three pigs by placing magnetic devices laparoscopically in the four abdominal quadrants[49]. The devices were left in place for 2 to 4 hours. Full-thickness abdominal wall sections were harvested from these quadrants plus control at 0, 2, or 14 days after surgery. Histologic assessment was then performed. No gross or microscopic tissue damage was seen. Moreover, magnets are used in other applications in medicine, such as MRI, that are considered to be safe.

In summary, the main limitations of MAG-ESD technique are the coupling strength of magnets, abdominal wall thickness and possible harmful effects of magnets on human tissues. Moreover, the large size of the magnets and the high cost are significant problems that need to be overcome.

**FUTURE**

MAG-ESD has the potential to act as a surgeon’s “second hand” to make complex endoscopic procedures like ESD easier to perform. However, to achieve this, it has to overcome a number of challenges and limitations. Techniques like the clip-with-line method are able to offer basic retraction simply and affordably with equipment already available in an endoscopy suite. For the more complex MAG-ESD to be appealing, it has to offer a significant advantage during complex endoscopic procedures such as ESD, over its competitors. More recently, a trial by Rivas *et al*[50]used the Magnetic Surgical System developed by Levita Magnetics Corp. (San Mateo, CA) to successfully perform fifty reduced-port laparoscopic cholecystectomies with no serious adverse events and excellent exposure of the surgical site. If the safety, efficacy, and ease of use of this system could be translated to ESD, magnetic endoscopy could be more appealing.

In summary, several major limitations that need to be overcome include the sheer size of the magnets and its high costs. As magnetic attraction decays over distance, its effect may vary in patients with different abdominal wall thickness. Moreover, air insufflated during endoscopy will increase the distance between magnets. A magnetic control system that could adapt magnetic strength to distance would reduce the variability of MAG-ESD. The more precise the control of the internal magnet, the more it might act like the desired second hand. With increased precision, magnets may be used for indications other than retraction, such as tissue dissection, improve visualization, and ultimately enhance the field of endoscopic surgery.

**ACKNOWLEDGEMENTS**

We would like to thank Mohamed Elmassry, MD for his significant contribution to this article by providing the artwork.

**REFERENCES**

1 **Saito Y**, Kawano H, Takeuchi Y, Ohata K, Oka S, Hotta K, Okamoto K, Homma K, Uraoka T, Hisabe T, Chang DK, Zhou PH. Current status of colorectal endoscopic submucosal dissection in Japan and other Asian countries: progressing towards technical standardization. *Dig Endosc* 2012; **24** Suppl 1: 67-72 [PMID: 22533756 DOI: 10.1111/j.1443-1661.2012.01282.x]

2 **Draganov PV**, Gotoda T, Chavalitdhamrong D, Wallace MB. Techniques of endoscopic submucosal dissection: application for the Western endoscopist? *Gastrointest Endosc* 2013; **78**: 677-688 [PMID: 24021491 DOI: 10.1016/j.gie.2013.07.033]

3 **Bhatt A**, Abe S, Kumaravel A, Vargo J, Saito Y. Indications and Techniques for Endoscopic Submucosal Dissection. *Am J Gastroenterol* 2015; **110**: 784-791 [PMID: 25623656 DOI: 10.1038/ajg.2014.425]

4 **Cao Y**, Liao C, Tan A, Gao Y, Mo Z, Gao F. Meta-analysis of endoscopic submucosal dissection versus endoscopic mucosal resection for tumors of the gastrointestinal tract. *Endoscopy* 2009; **41**: 751-757 [PMID: 19693750 DOI: 10.1055/s-0029-1215053]

5 **Rösch T**, Sarbia M, Schumacher B, Deinert K, Frimberger E, Toermer T, Stolte M, Neuhaus H. Attempted endoscopic en bloc resection of mucosal and submucosal tumors using insulated-tip knives: a pilot series. *Endoscopy* 2004; **36**: 788-801 [PMID: 15326574 DOI: 10.1055/s-2004-825838]

6 **Choi IJ**, Kim CG, Chang HJ, Kim SG, Kook MC, Bae JM. The learning curve for EMR with circumferential mucosal incision in treating intramucosal gastric neoplasm. *Gastrointest Endosc* 2005; **62**: 860-865 [PMID: 16301026 DOI: 10.1016/j.gie.2005.04.033]

7 **Gotoda T**, Friedland S, Hamanaka H, Soetikno R. A learning curve for advanced endoscopic resection. *Gastrointest Endosc* 2005; **62**: 866-867 [PMID: 16301027 DOI: 10.1016/j.gie.2005.07.055]

8 **Oda I**, Suzuki H, Nonaka S, Yoshinaga S. Complications of gastric endoscopic submucosal dissection. *Dig Endosc* 2013; **25** Suppl 1: 71-78 [PMID: 23368986 DOI: 10.1111/j.1443-1661.2012.01376.x]

9 **Gotoda T**, Ho KY, Soetikno R, Kaltenbach T, Draganov P. Gastric ESD: current status and future directions of devices and training. *Gastrointest Endosc Clin N Am* 2014; **24**: 213-233 [PMID: 24679233 DOI: 10.1016/j.giec.2013.11.009]

10 **Repici A**. Endoscopic submucosal dissection: established, or still needs improving? *Gastrointest Endosc* 2009; **69**: 16-18 [PMID: 19111684 DOI: 10.1016/j.gie.2008.07.030]

11 **Fukami N**. What we want for ESD is a second hand! Traction method. *Gastrointest Endosc* 2013; **78**: 274-276 [PMID: 23867374 DOI: 10.1016/j.gie.2013.04.192]

12 **Kobayashi T**, Gotohda T, Tamakawa K, Ueda H, Kakizoe T. Magnetic anchor for more effective endoscopic mucosal resection. *Jpn J Clin Oncol* 2004; **34**: 118-123 [PMID: 15078906 DOI: 10.1093/jjco/hyh025]

14 **Lee BI.** Debates on colorectal endoscopic submucosal dissection - traction for effective dissection: Gravity is enough. *Clin Endosc* 2013; **46**: 467-471 [PMID: 24143304 DOI: 10.5946/ce.2013.46.5.467]

13 **Oyama T**. Counter traction makes endoscopic submucosal dissection easier. *Clin Endosc* 2012; **45**: 375-378 [PMID: 23251884 DOI: 10.5946/ce.2012.45.4.375]

15 **Yamamoto S**, Uedo N, Ishihara R, Kajimoto N, Ogiyama H, Fukushima Y, Yamamoto S, Takeuchi Y, Higashino K, Iishi H, Tatsuta M. Endoscopic submucosal dissection for early gastric cancer performed by supervised residents: assessment of feasibility and learning curve. *Endoscopy* 2009; **41**: 923-928 [PMID: 19802773 DOI: 10.1055/s-0029-1215129]

16 **Ono S**, Fujishiro M, Koike K. Endoscopic submucosal dissection for superficial esophageal neoplasms. *World J Gastrointest Endosc* 2012; **4**: 162-166 [PMID: 22624067 DOI: 10.4253/wjge.v4.i5.162]

17 **Imaeda H**, Hosoe N, Kashiwagi K, Ohmori T, Yahagi N, Kanai T, Ogata H. Advanced endoscopic submucosal dissection with traction. *World J Gastrointest Endosc* 2014; **6**: 286-295 [PMID: 25031787 DOI: 10.4253/wjge.v6.i7.286]

18 **Lee BI**. Debates on colorectal endoscopic submucosal dissection - traction for effective dissection: gravity is enough. *Clin Endosc* 2013; **46**: 467-471 [PMID: 24143304]

19 **Kondo H**, Gotoda T, Ono H, Oda I, Kozu T, Fujishiro M, Saito D, Yoshida S. Percutaneous traction-assisted EMR by using an insulation-tipped electrosurgical knife for early stage gastric cancer. *Gastrointest Endosc* 2004; **59**: 284-288 [PMID: 14745409]

20 **Saito Y**, Emura F, Matsuda T, Uraoka T, Nakajima T, Ikematsu H, Gotoda T, Saito D, Fujii T. A new sinker-assisted endoscopic submucosal dissection for colorectal cancer. *Gastrointest Endosc* 2005; **62**: 297-301 [PMID: 16046999]

21 **Imaeda H**, Iwao Y, Ogata H, Ichikawa H, Mori M, Hosoe N, Masaoka T, Nakashita M, Suzuki H, Inoue N, Aiura K, Nagata H, Kumai K, Hibi T. A new technique for endoscopic submucosal dissection for early gastric cancer using an external grasping forceps. *Endoscopy* 2006; **38**: 1007-1010 [PMID: 16673308]

22 **Parra-Blanco A**, Nicolas D, Arnau MR, Gimeno-Garcia AZ, Rodrigo L, Quintero E. Gastric endoscopic submucosal dissection assisted by a new traction method: the clip-band technique. A feasibility study in a porcine model (with video). *Gastrointest Endosc* 2011; **74**: 1137-1141 [PMID: 22032320 DOI: 10.1016/j.gie.2011.07.037]

23 **Matsumoto K**, Nagahara A, Sakamoto N, Suyama M, Konuma H, Morimoto T, Sagawa E, Ueyama H, Takahashi T, Beppu K, Shibuya T, Osada T, Yoshizawa T, Ogihara T, Watanabe S. A new traction device for facilitating endoscopic submucosal dissection (ESD) for early gastric cancer: the "medical ring". *Endoscopy* 2011; **43** Suppl 2 UCTN: E67-E68 [PMID: 21341187]

24 **Ishigooka M,** Uchizawa M, Kusama K, Takahashi b, Takagi H, Morizono R, Koyama J. Endoscopic resection with local injection of HSE solution by direct incision of submucosa for early gastric cancer (S-ERHSE). *Endoscopia Digestiva* 2002; **11**: 1753-1757

25 **Motohashi O**, Nishimura K, Nakayama N, Takagi S, Yanagida N. Endoscopic submucosal dissection (two-point fixed ESD) for early esophageal cancer. *Dig Endosc* 2009; **21**: 176-179 [PMID: 19691765 DOI: 10.1111/j.1443-]

26 **Hirao M**, Masuda K, Asanuma T, Naka H, Noda K, Matsuura K, Yamaguchi O, Ueda N. Endoscopic resection of early gastric cancer and other tumors with local injection of hypertonic saline-epinephrine. *Gastrointest Endosc* 1988; **34**: 264-269 [PMID: 3391382]

27 **Diana M**, Chung H, Liu KH, Dallemagne B, Demartines N, Mutter D, Marescaux J. Endoluminal surgical triangulation: overcoming challenges of colonic endoscopic submucosal dissections using a novel flexible endoscopic surgical platform: feasibility study in a porcine model. *Surg Endosc* 2013; **27**: 4130-4135 [PMID: 23793807 DOI: 10.1007/]

28 **Ho KY**, Phee SJ, Shabbir A, Low SC, Huynh VA, Kencana AP, Yang K, Lomanto D, So BY, Wong YY, Chung SC. Endoscopic submucosal dissection of gastric lesions by using a Master and Slave Transluminal Endoscopic Robot (MASTER). *Gastrointest Endosc* 2010; **72**: 593-599 [PMID: 20646698 DOI: 10.1016/j.gie.2010]

29 **Ritsuno H**, Sakamoto N, Osada T, Goto SP, Murakami T, Ueyama H, Mori H, Matsumoto K, Beppu K, Shibuya T, Nagahara A, Ogihara T, Watanabe S. Prospective clinical trial of traction device-assisted endoscopic submucosal dissection of large superficial colorectal tumors using the S-O clip. *Surg Endosc* 2014; **28**: 3143-3149 [PMID: 24879138 DOI: 10.1007/s00464-014-3572-0]

30 **Kobiela J**, Grymek S, Wojanowska M, Łubniewski M, Makarewicz W, Dobrowolski S, Łachiński A, J, Śledziński Z. Magnetic instrumentation and other applications of magnets in NOTES. *Wideochir Inne Tech Maloinwazyjne* 2012; **7**: 67-73 [PMID: 23256005 DOI: 10.5114/wiitm.2011.25665]

31 **Gotoda T**, Oda I, Tamakawa K, Ueda H, Kobayashi T, Kakizoe T. Prospective clinical trial of magnetic-anchor-guided endoscopic submucosal dissection for large early gastric cancer (with videos). *Gastrointest Endosc* 2009; **69**: 10-15 [PMID: 18599053 DOI: 10.1016/j.gie.2008.03.1127]

32 **Sagawa M**, Fujimura S, Yamamoto H, Matsuura Y, Hiraga K. Permanent magnet materials based on the rare earth-iron-boron tetragonal compounds. *IEEE Trans Magn* 1984; **20**: 1584-1589 [DOI: 10.1109/TMAG.1984.1063214]

33 **Matsuzaki I**, Miyahara R, Hirooka Y, Funasaka K, Furukawa K, Ohno E, Nakamura M, Kawashima H, Maeda O, Watanabe O, Ando T, Kobayashi M, Goto H. Simplified magnetic anchor-guided endoscopic submucosal dissection in dogs (with videos). *Gastrointest Endosc* 2014; **80**: 712-716 [PMID: 25085334 DOI: 10.1016/j.gie.2014.05.334]

34 **Aihara H**, Ryou M, Kumar N, Ryan MB, Thompson CC. A novel magnetic countertraction device for endoscopic submucosal dissection significantly reduces procedure time and minimizes technical difficulty. *Endoscopy* 2014; **46**: 422-425 [PMID: 24573770 DOI: 10.1055/s-0034-1364940]

35 **Spaun GO**, Zheng B, Swanström LL. A multitasking platform for natural orifice translumenal endoscopic surgery (NOTES): a benchtop comparison of a new device for flexible endoscopic surgery and a standard dual-channel endoscope. *Surg Endosc* 2009; **23**: 2720-2727 [PMID: 19360364 DOI: 10.1007/s00464-009-0476-5]

36 **Thompson CC**, Ryou M, Soper NJ, Hungess ES, Rothstein RI, Swanstrom LL. Evaluation of a manually driven, multitasking platform for complex endoluminal and natural orifice transluminal endoscopic surgery applications (with video). *Gastrointest Endosc* 2009; **70**: 121-125 [PMID: 19394008 DOI: 10.1016/j.gie.2008.11.007]

37 **von Renteln D**, Vassiliou MC, Rösch T, Rothstein RI. Triangulation: the holy grail of endoscopic surgery? *Surg Endosc* 2011; **25**: 1355-1357 [PMID: 21424191 DOI: 10.1007/s00464-011-1650-0]

38 **Karimyan V**, Sodergren M, Clark J, Yang GZ, Darzi A. Navigation systems and platforms in natural orifice translumenal endoscopic surgery (NOTES). *Int J Surg* 2009; **7**: 297-304 [PMID: 19481186 DOI: 10.1016/j.ijsu.2009.05.007]

39 **Park S**, Bergs RA, Eberhart R, Baker L, Fernandez R, Cadeddu JA. Trocar-less instrumentation for laparoscopy: magnetic positioning of intra-abdominal camera and retractor. *Ann Surg* 2007; **245**: 379-384 [PMID: 17435544 DOI: 10.1097/01.sla.0000232518.01447.c7]

40 **Scott DJ**, Tang SJ, Fernandez R, Bergs R, Goova MT, Zeltser I, Kehdy FJ, Cadeddu JA. Completely transvaginal NOTES cholecystectomy using magnetically anchored instruments. *Surg Endosc* 2007; **21**: 2308-2316 [PMID: 17704871 DOI: 10.1007/s00464-007-9498-z]

41 **Scott DJ,** Tang SJ, Bergs R, Fernandez R. Magnetically- Anchored Instruments for Transgastric Endoscopic Surgery. SAGES 2006; Apr 29, 2006; Dallax, Texas, USA.

42 **Ryou M**, Thompson CC. Magnetic retraction in natural-orifice transluminal endoscopic surgery (NOTES): addressing the problem of traction and countertraction. *Endoscopy* 2009; **41**: 143-148 [PMID: 19214894 DOI: 10.1055/s-0028-1119454]

43 **Dominguez G**, Durand L, De Rosa J, Danguise E, Arozamena C, Ferraina PA. Retraction and triangulation with neodymium magnetic forceps for single-port laparoscopic cholecystectomy. *Surg Endosc* 2009; **23**: 1660-1666 [PMID: 19415381 DOI: 10.1007/s00464-009-0504-5]

44 **Wang Z**, Wang L, Tang B, Frank T, Brown S, Cuschieri A. Retraction by surface ferromagnetisation of target tissues: preliminary studies on feasibility of magnetic retraction for endoscopic surgery. *Surg Endosc* 2008; **22**: 1838-1844 [PMID: 18071794 DOI: 10.1007/s00464-007-9716-8]

45 **Zeltser IS**, Bergs R, Fernandez R, Baker L, Eberhart R, Cadeddu JA. Single trocar laparoscopic nephrectomy using magnetic anchoring and guidance system in the porcine model. *J Urol* 2007; **178**: 288-291 [PMID: 17499803 DOI: 10.1016/j.juro.2007.03.001]

46 **Best SL**, Bergs R, Gedeon M, Paramo J, Fernandez R, Cadeddu JA, Scott DJ. Maximizing coupling strength of magnetically anchored surgical instruments: how thick can we go? *Surg Endosc* 2011; **25**: 153-159 [PMID: 20533063 DOI: 10.1007/s00464-010-1149-0]

47 **Milad MP**, Terkildsen MF. The spinal needle test effectively measures abdominal wall thickness before cannula placement at laparoscopy. *J Am Assoc Gynecol Laparosc* 2002; **9**: 514-518 [PMID: 12386366 DOI: 10.1016/S1074-3804(05)60529-4]

48 **Wolber T**, Ryf S, Binggeli C, Holzmeister J, Brunckhorst C, Luechinger R, Duru F. Potential interference of small neodymium magnets with cardiac pacemakers and implantable cardioverter-defibrillators. *Heart Rhythm* 2007; **4**: 1-4 [PMID: 17198980 DOI: 10.1016/j.hrthm.2006.09.012]

49 **Best SL**, Kabbani W, Scott DJ, Bergs R, Beardsley H, Fernandez R, Mashaud LB, Cadeddu JA. Magnetic anchoring and guidance system instrumentation for laparo-endoscopic single-site surgery/natural orifice transluminal endoscopic surgery: lack of histologic damage after prolonged magnetic coupling across the abdominal wall. *Urology* 2011; **77**: 243-247 [PMID: 20947148 DOI: 10.1016/j.urology.2010.05.041]

50 **Rivas H**, Robles I, Riquelme F, Vivanco M, Jiménez J, Marinkovic B, Uribe M. Magnetic Surgery: Results from First Prospective Clinical Trial in 50 Patients. *Ann Surg* 2016; Epub ahead of print [PMID: 27759614]

**P-Reviewer:** Amornyotin S, Fiori E, Tepes B **S-Editor:** Qi Y **L-Editor: E-Editor:**

**Specialty type:** Gastroenterology and hepatology

**Country of origin:** United States

**Peer-review report classification**

Grade A (Excellent): A



Grade B (Very good): B

Grade C (Good): C

Grade D (Fair): 0

Grade E (Poor): 0

Figure 1: Endoscopic Submucosal Dissection With and Withou



**Figure 1** **Endoscopic submucosal dissection with and without retraction.** Endoscopic submucosal dissection (ESD) without retraction on the left demonstrates the mucosa impeding the visual field. ESD with retraction using the clip-line method on the right allows for better visualization of the surgical field with the mucosa pulled away.




Figure 2 Magnetic anchor system using neodymium permanent magnet. ESD: Endoscopic submucosal dissection.