**Name of Journal:** *World Journal of Orthopedics*

**Manuscript NO:** 49248

**Manuscript Type:** CASE REPORT

**Pseudotumor recurrence in a post-revision total hip arthroplasty with stem neck modularity: A case report**

Desai BR *et al*. Post-revision pseudotumor recurrence

Bhumit R Desai, Gonzalo E Sumarriva, George F Chimento

**Bhumit Desai**, The University of Queensland School of Medicine, Ochsner Clinical School, New Orleans, LA 70121, United States

**Gonzalo Sumarriva, George Chimento**, Department of Orthopaedics, Ochsner Clinic Foundation, New Orleans, LA 70121, United States

**ORCID number:** Bhumit R Desai (0000-0002-5456-3939); Gonzalo E Sumarriva (0000-0001-6399-168X); George F Chimento (0000-0001-7507-5662).

**Author contributions:** Chimento GF was the patient’s primary joint surgeon; Sumarriva GE analyzed and interpreted the imaging findings; Desai BR reviewed the literature and drafted the manuscript; Desai BR, Sumarriva GE, and Chimento GF were responsible for the revision of the manuscript for important intellectual content; all authors issued final approval for the version to be submitted.

**Informed consent statement:** Informed written consent was obtained from the patient for publication of this report and any accompanying images.

**Conflict-of-interest statement:** The authors declare that they have no conflict of interest.

**CARE Checklist (2016) statement:** The authors have read the CARE Checklist (2016), and the manuscript was prepared and revised according to the CARE Checklist (2016).

**Open-Access:** This article is an open-access article which was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution Non Commercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: <http://creativecommons.org/licenses/by-nc/4.0/>

**Manuscript source:** Unsolicited manuscript

**Corresponding author: George F Chimento, FACS, MD, Chairman, Surgeon,** Department of Orthopedics, Ochsner Clinic Foundation, 1514 Jefferson Highway**,** New Orleans, LA 70121, United States. gchimento@ochsner.org

**Telephone:** +1-866-6247637

**Fax:** +1-504-8422988

**Received:** May 21, 2019

**Peer-review started:** May 23, 2019

**First decision:** July 31, 2019

**Revised:** September 19, 2019

**Accepted:** November 28, 2019

**Article in press:**

**Published online:**

**Abstract**

***BACKGROUND***

This unique presentation of hip swelling is only the seventh reported case of pseudotumor recurrence in a non-metal-on-metal total hip arthroplasty (THA) construct. The constellation of patient symptoms and laboratory findings contradict the expected elevated serum metal ion levels associated with the formation of pseudotumor. The presentation, lab trends, and imaging findings contribute to the growing base of knowledge surrounding the effects of corrosion in arthroplasty constructs with stem-neck modularity.

***CASE SUMMARY***

A 74-year-old man status post primary THA presented with left hip swelling and elevated serum metal ions five years after implantation of a modular stem-neck prosthesis. The swelling was diagnosed as pseudotumor based on laboratory trends and imaging findings and was treated with revision arthroplasty that completely resolved the initial hip swelling. The patient presented with recurrent hip swelling and recurrent pseudotumor findings on imaging in the same hip four months later. Non-operative management with ultrasound-guided hip aspiration resulted in symptom relief and resolution of the recurrent swelling. After one year of follow-up, the patient had no further recurrences of hip swelling.

***CONCLUSION***

This case of post-revision pseudotumor recurrence elucidates attributable patient, surgical, and implant factors with a discussion of diagnostics, management, and prognosis for patients with pseudotumor in non-metal-on-metal arthroplasty constructs.

**Key words:** Revision arthroplasty; Pseudotumor; Stem-neck modularity; Pseudotumor recurrence; Metal-on-metal arthroplasty; Modular stem; Case report

**© The Author(s) 2019.** Published by Baishideng Publishing Group Inc. All rights reserved.

**Core tip:** This unique presentation of hip swelling is only the seventh reported case of pseudotumor recurrence in a non-metal-on-metal total hip arthroplasty construct. These patients require a low threshold for clinical evaluation, using multiple modalities such as history, physical exam, magnetic resonance imaging, and serum metal ion level testing. The presentation, lab trends, and imaging findings of this case contribute to the growing base of knowledge surrounding the effects of corrosion in arthroplasty constructs with stem-neck modularity.

Desai BR, Sumarriva GE, Chimento GF.Pseudotumor recurrence in a post-revision total hip arthroplasty with stem neck modularity: A case report. *World J Orthop* 2019; In press

**INTRODUCTION**

Total hip arthroplasty (THA) yields excellent patient satisfaction and offers improvement in mobility and quality of life[1].Advancements in technology have shaped prosthesis design to accommodate patient factors, level of activity, and surgical approaches. THA constructs can be composed of three components-the femoral component, acetabular component, and bearing surface. One variation of prosthesis design is modular neck systems, which allow a wide range of intraoperative adjustments of the femoral head, adjustment of leg length, and horizontal femoral offset [2]. Modularity aims to reproduce the natural biomechanics of the hip through the adjustment of individual prosthetic components. Proximal stem-neck modularity includes a junction that can have a double taper subjective to both axial and bending stresses. This stress leads to an uncommon complication known as fretting corrosion[2]. Destruction at contact sites from relative micromotion between two components of the prosthesis can lead to mid-stem failure in modular stem prostheses[2].

Metal-on-metal (MoM) THA is an alternative prosthetic design that was commonly used before unexpected complications led to its disuse. The concept was originally designed with the advantage of increased joint stability and fewer dislocations given a larger surface area for contact. An unforeseen complication of MoM arthroplasty was pseudotumor formation–at rates as high as 59% in studied cohorts[3].This association has been seen consistently in metal-on metal THA[4].A study by Ming showed the elapsed time between MoM THA and revision arthroplasty for pseudotumor formation to be 61 mo[5]. In the same study cohort of 97 patients, 3% had post-revision pseudotumor recurrence[5]. Recently, pseudotumor formation has been linked to non-metal-on-metal THA, especially in patients with modular constructs. Based on the literature review conducted by the authors of this report, only seven cases of pseudotumor recurrence have been associated in THA with non-MoM modular constructs[5,6].

The authors of this paper present a case of post-revision pseudotumor recurrence in a THA with a modular femoral component. The patient was informed that data concerning the case would be submitted for publication, and patient consent was obtained. Table 1 outlines a timeline of the patient presentation.

**CASE PRESENTATION**

***Chief complaints***

A 74-year-old African-American male was referred to the senior author on the orthopedic joint service with left hip swelling and pain more than five years after THA (Figures 1 and 2). The patient presented with a one-month history of left hip swelling. Associated symptoms included discomfort with movement, decreased range of motion, and left lower extremity swelling. There was no recent history of trauma and the left limb was neurovascular-intact.

***History of presenting illness***

The patient presented with left hip swelling and was evaluated with physical exam and imaging. Magnetic resonance imaging (MRI) showed a well-defined complex lesion anterior to the left hip prosthesis (Figure 3) and a septated simple fluid collection posterior to the left hip prosthesis; further evaluation revealed elevated serum metal ion levels-serum cobalt 12 ng/mL (ref. range 0.0-0.9 ng/mL) and serum chromium 2.8 ng/mL (ref. range < 0.3 ng/mL) The lesion was first unsuccessfully managed with aspiration by interventional radiology. After continued discomfort and signs of component loosening on imaging, the patient opted for with revision arthroplasty. During revision, the initial incision was used to resect necrotic soft tissue and drain a large blood-tinged fluid collection in the articular space. The collected tissue was sent for culture and pathology, which showed fibrin and fibrous tissue without the presence of neutrophils. On extraction of the head and neck, there was a significant degree of corrosion along the mid-portion and body of the neck. The acetabular component was well-fixed and in good condition to accommodate a new constrained liner. The revision was completed by extracting the proximal segment of the femoral stem and replacing it with an 18 mm x 155 mm modular stem, a 25-mm neutral proximal body, and 28 mm neutral femoral head. The new components were firmly seated, demonstrating stability and no impingement on range of motion. At two-week follow-up patient was ambulating well with a cane with no acute post-operative complications (Figure 4).

The patient presented again four months post-revision with a two-week history of recurrent left hip swelling. A new lesion was noted on MRI in a location similar to the initial lesion (Figure 5). Serum metal ion levels at this time had returned to baseline (serum cobalt = 1.5 ng/mL; serum chromium = 1.4 ng/mL). Ultrasound-guided fine needle aspiration of the lesion yielded 15cc of clotted blood, resembling the aspiration findings of the initial lesion suggesting adverse local tissue reaction (ALTR) recurrence.

***History of past illness***

The patient’s only notable past medical history was hypertension well-controlled on medication.

***Laboratory examinations***

New pseudotumor formation in a location similar to the initial lesion was noted on MRI. Serum metal ion levels were elevated above the normal range, but significantly lower than on initial presentation for pseudotumor formation (serum cobalt = 1.5 ng/mL; serum chromium = 1.4 ng/mL).

***Imaging examinations***

MRI at time of initial presentation showed a well-defined complex lesion anterior to the left hip prosthesis and a septate simple fluid collection posterior to the left hip prosthesis. Repeat MRI at recurrence confirmed a similar lesion larger in size at the same location.

**FINAL DIAGNOSIS**

The final diagnosis of the presented case is post-revision ALTR recurrence*.*

**TREATMENT**

An ultrasound-guided hip aspiration provided symptomatic relief and resolution of hip swelling.

**OUTCOME AND FOLLOW-UP**

At clinical follow-up 18 mo post-revision there were no radiographic signs of implant loosening and the patient reported no further episodes of hip swelling with satisfactory return to baseline mobility.

**DISCUSSION**

The goal of prosthetic modularity is to offer intraoperative adjustments to manage leg length, anteversion, and offset to better replicate natural hip biomechanics[7].Dual modularity refers to arthroplasty construct components (head, neck, and stem) that may be interchanged, specifically at the proximal portion where the head-neck and neck-stem junctions are located. Modularity at the head-neck junction was originally designed to allow movements in different planes and to reduce friction and wear[2]. However, the additional junctions pose complications such as fracture, corrosion, fretting, dissociation, and cold-welding[8]. The specific concern of fretting and crevice corrosion is a known association established in modular constructs[9]. Micromotion at the neck-stem junction from natural bending movements can predispose components to fretting corrosion and metal debris deposition[9].

Pseudotumor is a term historically associated with metal-on-metal arthroplasty implants. However, with the growing recognition that cystic masses causing pain, pressure effects, bone and soft-tissue destruction can be associated with implants of different bearings and degrees of modularity, the literature has adapted new terms such as asymptomatic lymphocytic vasculitis-associated lesions based on histological findings and the more encompassing term ALTR. A recent systematic review of pseudotumors suggested an incidence ranging from 1% up to 6% of all hip THAs[4]. These pseudotumors are neither infective nor neoplastic and have been proposed to occur secondary to excessive wear debris, foreign body reaction, metal hypersensitivity, and certain patient/surgical-related factors[6].

One commonly suggested etiology of ALTRs is metal ion deposition generated by mechanical wear and surface corrosion at modular junctions[10]. Fretting is a relative oscillatory tangential movement of small amplitude which can occur between contacting surfaces of two immobile and fixed object[11]. In presence of body fluids, fretting can be accompanied by corrosive effects. The presence of water within human tissues initiates a repassivation reaction, where the exposed metal surface forms an oxide and releases hydrogen ions. Negatively charged chloride ions migrate to balance these positively charged hydrogen ions, producing hydrochloric acid, which dissolves metal alloys. The resulting corrosion debris generated from the modular junction appears responsible for ALTR[11].

Looking further at instances of pseudotumor recurrence, a study by Kwon looked at 40 revisions after the presence of symptomatic ALTR diagnosed on metal artifact reduction sequence MRI with elevated metal ion levels[6]. These patients all had metal-on-polyethylene bearings on titanium alloy femoral stems so there was only one site of modularity at the head-neck junction. They used the Goldberg criteria to assess fretting and corrosion and found that three-quarters of the patients had Grade 3 tissue damage and metal debris at the femoral component’s trunnion neck taper[6]. They were revised to a ceramic femoral head with a titanium adaptor sleeve over the femoral stem trunnion. In the patient that experienced recurrence, there was no linear correlation to post-revision serum metal ion levels. Rather they suggested that pre-revision MRI findings of lesions with abductor deficiency and intra-operative tissue damageshould be stronger predictive factors for post-revision complications such as pseudotumor recurrence[6].

Thus, it would lead us to believe that with an increase in the number of modular junctions, there exists an increased propensity for fretting corrosion and an increased incidence of ALTRs. Several recent studies looking at revision outcomes in THA with dual modularity suggest exactly this[5,8-11]. The benefits of dual modularity are clear, it provides greater flexibility and facilitates the adjustment of leg length, femoral anteversion, and offset in order to accurately reproduce the center of rotation of the hip. However, both axial and bending stresses at the stem-neck junction strongly influence fretting corrosion eventually leading to mechanical failure of the modular neck and the subsequent need for revision.

Kwon looked at outcomes after revision dual taper THA in about 200 patients with symptomatic ALTR due to taper corrosion[6]. A staggering 20% of the cohort experienced post-revision complications. 3% of those patients experienced pseudotumor recurrence. 5 of the patients were revised to a cemented monoblock femoral stem with a ceramic head, so recurrence may be less likely attributable to stem-neck modularity or head-neck taper corrosion. The authors attributed recurrence to surgical, patient, and implant factors. They hypothesized that incomplete surgical debridement and inadequate removal of pseudotumor can contribute. However, extensive debridement needs to be performed with preservation of nearby neurovascular structures.

The uniqueness of this case report is highlighted by the post-revision ALTR recurrence. Previously reported cases of stem-neck modularity leading to pseudotumor formation exhibited implant damage between two metal components[12]. This lends to the notion that micromotion between adjoining metal-on-metal components are susceptible to the same corrosion and metal ion deposition seen in MoM arthroplasty. In order for this mechanism to be true however, one would expect more cases of pseudotumor formation given the number of THAs performed yearly using stem-neck modularity.

Current recommendations in symptomatic patients with ALTRs suggest avoiding overreliance on any single investigative tool in the clinical decision-making process for revision surgery[13]. There should be a low threshold for clinical evaluation, especially in patients dual modularity THA. A combination approach has been recommended by the consensus statement between the American Association of Hip and Knee Surgeons, American Academy of Orthopaedic Surgeons, and Hip Society in regards to managing symptomatic patients with dual modular THA. The recommendation states “And in patients where revision in indicated, knowledge of the pre-revision risk factors we talked about earlier can mitigate the incidence of complications and need for re-revision.”

**CONCLUSION**

This case highlights a post-revision pseudotumor recurrence in a modular femoral stem THA not utilizing metal-on-metal prostheses. This article aims to elucidate patient, surgical, and implant factors that may contribute to pseudotumor recurrence with a discussion of diagnostics, management, and prognosis.

**REFERENCES**

1 **Ritter MA**, Albohm MJ. Overview: maintaining outcomes for total hip arthroplasty. The past, present, and future. *Clin Orthop Relat Res* 1997; : 81-87 [PMID: 9372760 DOI: 10.1097/00003086-199711000-00009]

2 **Krishnan H**, Krishnan SP, Blunn G, Skinner JA, Hart AJ. Modular neck femoral stems. *Bone Joint J* 2013; **95-B**: 1011-1021 [PMID: 23908413 DOI: 10.1302/0301-620X.95B8.31525]

3 **Hart AJ**, Satchithananda K, Liddle AD, Sabah SA, McRobbie D, Henckel J, Cobb JP, Skinner JA, Mitchell AW. Pseudotumors in association with well-functioning metal-on-metal hip prostheses: a case-control study using three-dimensional computed tomography and magnetic resonance imaging. *J Bone Joint Surg Am* 2012; **94**: 317-325 [PMID: 22336970 DOI: 10.2106/JBJS.J.01508]

4 **Daniel J**, Holland J, Quigley L, Sprague S, Bhandari M. Pseudotumors associated with total hip arthroplasty. *J Bone Joint Surg Am* 2012; **94**: 86-93 [PMID: 22218386 DOI: 10.2106/JBJS.J.01612]

5 **Dimitriou D**, Liow MH, Tsai TY, Leone WA, Li G, Kwon YM. Early Outcomes of Revision Surgery for Taper Corrosion of Dual Taper Total Hip Arthroplasty in 187 Patients. *J Arthroplasty* 2016; **31**: 1549-1554 [PMID: 26895822 DOI: 10.1016/j.arth.2016.01.015]

6 **Kwon YM**, Rossi D, MacAuliffe J, Peng Y, Arauz P. Risk Factors Associated With Early Complications of Revision Surgery for Head-Neck Taper Corrosion in Metal-on-Polyethylene Total Hip Arthroplasty. *J Arthroplasty* 2018; **33**: 3231-3237 [PMID: 29958755 DOI: 10.1016/j.arth.2018.05.046]

7 **Hsu AR**, Gross CE, Levine BR. Pseudotumor from modular neck corrosion after ceramic-on-polyethylene total hip arthroplasty. *Am J Orthop (Belle Mead NJ)* 2012; **41**: 422-426 [PMID: 23365811]

8 **Kop AM**, Swarts E. Corrosion of a hip stem with a modular neck taper junction: a retrieval study of 16 cases. *J Arthroplasty* 2009; **24**: 1019-1023 [PMID: 18835686 DOI: 10.1016/j.arth.2008.09.009]

9 **Canham CD**, Muradov PI, Simpson JB, Incavo SJ. Corrosion and adverse local tissue reaction after total hip arthroplasty with a modular titanium alloy femoral neck. *Arthroplast Today* 2017; **3**: 211-214 [PMID: 29204482 DOI: 10.1016/j.artd.2017.03.003.]

10 **De Martino I**, Assini JB, Elpers ME, Wright TM, Westrich GH. Corrosion and Fretting of a Modular Hip System: A Retrieval Analysis of 60 Rejuvenate Stems. *J Arthroplasty* 2015; **30**: 1470-1475 [PMID: 25817187 DOI: 10.1016/j.arth.2015.03.010]


11 **Baxmann M**, Jauch SY, Schilling C, Blömer W, Grupp TM, Morlock MM. The influence of contact conditions and micromotions on the fretting behavior of modular titanium alloy taper connections. *Med Eng Phys* 2013; **35**: 676-83; discussion 676 [PMID: 22940445 DOI: 10.1016/j.medengphy.2012.07.013]

12 **Leung P**, Kudrna JC. Growth of an intrapelvic pseudotumor associated with a metal-on-metal total hip arthroplasty after revision arthroplasty causing a femoral nerve neuropathy. *Arthroplast Today* 2016; **2**: 105-109 [PMID: 28326410 DOI: 10.1016/j.artd.2016.07.001]


13 **Kwon YM**, Lombardi AV, Jacobs JJ, Fehring TK, Lewis CG, Cabanela ME. Risk stratification algorithm for management of patients with metal-on-metal hip arthroplasty: consensus statement of the American Association of Hip and Knee Surgeons, the American Academy of Orthopaedic Surgeons, and the Hip Society. *J Bone Joint Surg Am* 2014; **96**: e4 [PMID: 24382732 DOI: 10.2106/JBJS.M.00160]

**P-Reviewer:** Papachristou G **S-Editor:** Wang J **L-Editor:** **E-Editor:**


**Specialty type:** Orthopedics   
**Country of origin:** United States   
**Peer-review report classification**  
**Grade A (Excellent):** 0  
**Grade B (Very good):** B  
**Grade C (Good):** 0  
**Grade D (Fair):** 0 **Grade E (Poor):** 0




**Figure 1 Pre-operative anterior-posterior pelvic radiograph showing degenerative changes of left hip with severe joint space narrowing, subchondral sclerosis, and cyst formation.**




**Figure 2 Initial post-operative anterior-posterior pelvic radiograph showing the components of the prosthesis well-aligned and well seated.**



**Figure 3 Pre-revision T2-weighted coronal magnetic resonance image showing a well-defined 7.8 cm complex lesion anterior to the left hip prosthesis.**



**Figure 4 Post-revision anterior-posterior hip radiograph showing an aligned and well-fixed left hip prosthesis.**



**Figure 5 Post-revision short-T1 inversion recovery coronal magnetic resonance image demonstrating pseudotumor recurrence.**

**Table 1 Patient** **timeline describing presentation, symptom progression, and clinical course**

|  |  |
| --- | --- |
| **Timeline** | **Patient situation** |
| February 2012 | Past medical history: Degenerative osteoarthritis;  Underwent left total hip arthroplasty |
| November 2017 | Initial presentation for left hip swelling;  Fluid collection seen on magnetic resonance imaging |
| December 2017 | Treated with revision total hip arthroplasty;  Immediate resolution of pseudotumor post-revision |
| March 2018 | Recurrent left hip swelling;  Resolution after repeat magnetic resonance imaging and joint aspiration |
| March 2019 | After 1-yr of follow-up no further hip swelling recurrences or pain symptoms |