**Name of Journal:** *World Journal of Gastroenterology*

**Manuscript NO:** 58874

**Manuscript Type:** ORIGINAL ARTICLE

***Basic Study***

**Fedora-type magnetic compression anastomosis device for intestinal anastomosis**

Chen H *et al.* Fedora-type magnetic compression anastomosis device

Huan Chen, Tao Ma, Yue Wang, Hao-Yang Zhu, Zhe Feng, Rong-Qian Wu, Yi Lv, Ding-Hui Dong

**Huan Chen, Tao Ma, Yue Wang, Hao-Yang Zhu, Zhe Feng, Rong-Qian Wu, Yi Lv, Ding-Hui Dong,** National Local Joint Engineering Research Center for Precision Surgery & Regenerative Medicine, The First Affiliated Hospital of Xi'an Jiaotong University, Xi’an 710061, Shaanxi Province, China

**Huan Chen, Tao Ma, Yue Wang, Zhe Feng, Rong-Qian Wu, Yi Lv, Ding-Hui Dong,** Department of Hepatobiliary Surgery, The First Affiliated Hospital of Xi’an Jiaotong University, Xi’an 710061, Shaanxi Province, China

**Author contributions:** All authors helped to perform the research; Dong DH and Lv Y conceived and designed the experiments; Chen H and Dong DH contributed to performing the animal experiment; Chen H, Feng Z, and Wang Y collected and analyzed the data; Chen H, Zhu HY, Ma T, and Wu RQ contributed to manuscript writing; Dong DH and Lv Y contributed to critical revision of the manuscript; all authors read and approved the final manuscript.

**Supported by** the National Natural Science Foundation of China (to Lv Y), No. 81470896.

**Corresponding author: Ding-Hui Dong, MD, PhD, Surgeon, Teacher,** National Local Joint Engineering Research Center for Precision Surgery & Regenerative Medicine, The First Affiliated Hospital of Xi'an Jiaotong University, No. 76 West Yanta Road, Xi’an 710061, Shaanxi Province, China. dongdinghui0510@163.com

**Received:** August 12, 2020

**Revised:** August 29, 2020

**Accepted:** September 10, 2020

**Published online:** November 14, 2020

**Abstract**

BACKGROUND

Although previous studies have confirmed the feasibility of magnetic compression anastomosis (MCA), there is still a risk of long-term anastomotic stenosis. For traditional MCA devices, a large device is associated with great pressure, and eventually increased leakage.

AIM

To develop a novel MCA device to simultaneously meet the requirements of pressure and size.

METHODS

Traditional nummular MCA devices of all possible sizes were used to conduct ileac anastomosis in rats. The mean (± SD) circumference of the ileum was 13.34 ± 0.12 mm. Based on short- and long-term follow-up results, we determined the appropriate pressure range and minimum size. Thereafter, we introduced a novel “fedora-type” MCA device, which entailed the use of a nummular magnet with a larger sheet metal.

RESULTS

With traditional MCA devices, the anastomoses experienced stenosis and even closure during the long-term follow-up when the anastomat was smaller than Φ5 mm. However, the risk of leakage increased when it was larger than Φ4 mm. On comparison of the different designs, it was found that the “fedora-type” MCA device should be composed of a Φ4-mm nummular magnet with a Φ6-mm sheet metal.

CONCLUSION

The diameter of the MCA device should be greater than 120% of the enteric diameter. The novel “fedora-type” MCA device controls the pressure and optimizes the size.

**Key Words:** Magnetic compression anastomosis; Anastomotic stenosis; Size of anastomat; Compression pressure; Fedora-type magnetic compression anastomosis device

**Citation:** Chen H, Ma T, Wang Y, Zhu HY, Feng Z, Wu RQ, Lv Y, Dong DH. Fedora-type magnetic compression anastomosis device for intestinal anastomosis. *World J Gastroenterol* 2020; 26(42): 6614-6625

**URL:** https://www.wjgnet.com/1007-9327/full/v26/i42/6614.htm

**DOI:** https://dx.doi.org/10.3748/wjg.v26.i42.6614

**Core Tip:** To address some of the deficiencies in the current magnetic compression anastomosis (MCA) model, we explored the optimal size and pressure of the MCA device for intestinal anastomosis in rats. We found that the suggested diameter of the MCA device should be larger than 120% of the enteric diameter to avoid stenosis. Further, we developed a novel “fedora-type” MCA device for the current model, using a Φ4-mm nummular magnet with a Φ6-mm sheet metal. This model safely formed anastomosis and ensured long-term anastomosis. This novel anastomat controlled pressure and optimized the size, thus meeting our stipulated requirements.

**INTRODUCTION**

Since Obora *et al*[1] used magnetic compression anastomosis (MCA) to successfully reconstruct vessels for the first time in 1978, MCA has been proven to be capable of compressing and penetrating various tissues[2]. Thus, MCA has been applied in many scenarios, especially for conditions in the digestive tract, such as esophageal[3-5], intestinal[6-8], gastrointestinal[9-11], biliary-intestinal[12-14], and pancreas-intestinal anastomoses[15]. However, research has shown that there is a risk of long-term anastomotic stenosis and even closure after MCA[15-20]; this eventually restricted further clinical application of MCA.

Therefore, effective and reliable MCA must satisfy all of the following criteria: Appropriate pressure, safe formation of anastomosis without leakage in the short-term follow-up, adequate size, and avoidance of anastomotic stenosis or closure in the long-term follow-up. Unfortunately, previous studies mostly focused on the formation of anastomosis[10,21,22], and thus long-term outcomes were neglected. Conversely, for traditional MCA devices, the compression force was positively correlated with the size. Thus, larger anastomosis was associated with a higher risk of leakage[22].

Thus, for MCA, there are three uncertainties that require clarification. First, the minimum initial size of anastomosis needs to be determined for reconstruction of the digestive tract of a certain size. Second, the suitable compression pressure range to form anastomosis without leakage needs to be determined for the particular tissue to be anastomosed. Third, clarity is required to determine the most effective design of a novel MCA device to simultaneously meet the compression pressure and size requirements.

To address these gaps, we designed the following two experiments. First, based on the anatomical characteristics of the rat intestine, we used traditional nummular MCA devices of all possible sizes to conduct ileac side-to-side anastomosis. Based on the short-term follow-up results, we determined the appropriate pressure range required for MCA. According to the long-term follow-up results, we confirmed the minimum size required to avoid anastomotic stenosis or closure. Second, based on the results of the former experiment, we introduced a novel design concept, known as the “fedora-type,” to the MCA device to simultaneously meet the requirements of both pressure and size, so that stable anastomosis could be formed.

**MATERIALS AND METHODS**

***Study design and ethical considerations***

All experimental protocols were approved by the Committee on the Ethics of Animal Experiments of Xi'an Jiaotong University (No. XJTULAC2020-1281). This research was conducted based on the guidelines for the Care and Use of Laboratory Animals from Xi'an Jiaotong University Health Science Center. A total of 105 male Sprague-Dawley rats weighing 240-260 g were obtained from the Experimental Animal Center, Xi'an Jiaotong University, Xi'an, China. The circumference of the intestine was measured for each rat during the operation, and the mean (± standard deviation, SD) was 13.34 ± 0.12 mm. All rats were anesthetized by isoflurane inhalation and were commonly treated pre- and post-operation. Postoperative complications and survival rates were observed.

***Experiment 1****:* ***Comparison of traditional nummular MCA devices***

Sixty rats were divided into four groups (groups 1.1-1.4), with 15 rats in each group. Traditional nummular MCA devices with different sizes were used in each group. As shown in Figure 1A, the MCA device involved a pair of nummular magnets (parent and daughter parts, NdFeB and N45). The diameters of the MCA devices in groups 1.1-1.4 were 3, 4, 5, and 6 mm, respectively, and the corresponding mean (± SD) compression pressures were 54.56 ± 1.40, 126.07 ± 1.38, 147.56 ± 3.42, and 152.60 ± 2.67 kPa, respectively.

After anesthesia, a 3-cm midline incision was made, and the small intestine was removed and covered with sterile gauze in normal warm saline. Then, a 6-mm incision was made 12 cm distal to the cecum. Afterwards, the parent and daughter parts of the MCA device were inserted into the intestine from the incision, reaching 6 cm proximal and distal to the incision, respectively. After adjusting the locations of the magnets, they were gently coupled to compress the ileum wall. The incisions made in the intestine and abdominal wall were sutured (Figure 2).

***Experiment 2****:* ***Development of a fedora-type MCA device***

Forty-five rats were randomly divided into three groups (groups 2.1-2.3) with 15 rats in each group. Based on experiment 1, a self-made “fedora-type” MCA device with different designs was adopted in each group. This device also consisted of parent and daughter parts. Each part involved a nummular magnet (NdFeB, N45) and a larger sheet metal (Ti6Al4V), just like a fedora cap, as shown in Figure 1B and C. The nummular magnets for all the groups were Φ4 mm, and the sheet metals for groups 2.1-2.3 were Φ4, Φ5, and Φ6 mm, respectively. Additionally, the mean (± SD) compression pressures for the different groups were 126.07 ± 1.38, 80.69 ± 0.88, and 56.03 ± 0.61 kPa, respectively.

The surgical procedure used was the same as that described in experiment 1.

***X-ray examination***

X-ray fluoroscopy was conducted to confirm the accurate coupling of daughter and parent parts immediately after the operation (Figure 2E2 and F2). Routine X-rays were performed every day to verify the device’s movement and stable coupling in the digestive tract until the devices were discharged.

***Tissue harvest and analysis***

On postoperative days 30, 90, and 180, five rats in each group were euthanized to collect the anastomotic tissue specimens. The gross appearance of specimens was assessed based on a widely accepted scale, as shown in supplementary Table 1[23]. The sizes of the anastomosis were measured and analyzed using ImageJ\_v1.8.0. The mechanical properties were evaluated based on bursting pressure using a self-made manometer. The histological morphology of ileac stomas was evaluated using Masson’s trichrome staining and hematoxylin and eosin (HE) staining.

***Statistical analysis***

SPSS Statistics Software version 23.0 (IBM Corporation, Armonk, NY, United States) was used for all analyses. Categorical variables are reported as numbers and proportions, and were compared using Chi-squared or nonparametric tests as appropriate. Normal continuous variables are reported as the mean ± SD and were compared using analysis of variance tests. Abnormal variables are reported as medians [interquartile range (IQR)] and were compared using nonparametric tests. All hypothesis tests were two-sided, and *P* values < 0.05 were considered statistically significant. The significance levels (α) for *post hoc* tests were adjusted accordingly.

**RESULTS**

***Experiment 1****:* ***Comparison of traditional nummular MCA devices***

**Survival rate**, **expulsion time, and bursting pressure:** No notable difficulties were encountered, and blood loss during the surgical procedure was minimal. There were no significant differences in the survival rates between the groups (groups 1.1-1.4, 93.3%, 100%, 73.3%, and 73.3%; *P* = 0.083) (Table 1). However, the combined survival rate for groups 1.1 and 1.2 was significantly higher than that of groups 1.3 and 1.4 (96.7% *vs* 73.3%, *P* = 0.026).

Routine X-ray fluoroscopy showed that all traditional nummular MCA devices coupled tightly after operation. The larger devices appeared to require shorter expulsion time. The median expulsion times were 3 (IQR 3-4), 3 (IQR 3-4), 2 (IQR 1-3), and 2.5 (IQR 2-3) d for groups 1.1-1.4, respectively (*P* = 0.002) (Table 1).

The bursting pressure for group 1.4 was lower than that in the other groups on the 30th postoperative day (*P* = 0.032) (Table 1). There was no significant difference in the bursting pressure between any of the groups on postoperative days 90 and 180 (Table 1).

**Size of anastomosis:** On postoperative days 30 and 90, it was observed that as the size of the MCA device increased, the circumference of the anastomosis increased (*P* < 0.008, adjusted α = 0.008). On the 180th postoperative day, the circumference also increased with size, with the exception of that in group 1.1 when compared to group 1.2 (group 1.1 *vs* group 1.2, *P* = 0.044; *P* < 0.008 for other comparisons; adjusted α = 0.008) (Table 1).

For the smaller groups (groups 1.1 and 1.2), the anastomosis circumferences decreased as time progressed (group 1.1: 2.47 ± 0.18, 1.20 ± 0.18, and 0.35 ± 0.19 mm for postoperative days 30, 90, and 180, respectively, *P* < 0.001; group 1.2: 8.84 ± 0.31, 5.90 ± 0.27, and 2.07 ± 0.37 mm for postoperative days 30, 90, and 180, respectively, *P* < 0.017, adjusted α = 0.017) (Figure 3A1-A3 and B1-B3). In group 1.1, the anastomoses were nearly closed by the 90th postoperative day. In group 1.2, closure of anastomoses occurred by the 180th postoperative day. As for the larger groups (groups 1.3 and 1.4), no significant differences in the circumference were found between the different time points (group 1.3, *P* = 0.811; group 1.4, *P* = 0.830) (Figure 3C1-C3 and D1-D3).

**Morphological analysis:** On the 30th postoperative day, the gross appearance of the anastomoses in the smaller groups was better than that in the larger groups. In groups 1.1 and 1.2, the anastomoses were clean and intact, and the mucosa was smooth and flat without any ulcers or erosions (Figure 3A4, A5, B4, and B5). However, the adhesion around the anastomoses was severe in groups 1.3 and 1.4, and the mucosa was not smooth and flat (Figure 3C4, C5, D4, and D5). As shown in Table 1, the adhesion scores for groups 1.3 and 1.4 were significantly higher than those in groups 1.1 and 1.2, respectively (*P* < 0.008 for both, adjusted α = 0.008).

The histological morphology showed that the serosal, submucosal, and mucosal layers were interrupted by scar tissue in the larger groups (Figure 4A1 and A2). However, it was continuous in the smaller groups (Figure 4B1 and B2).

***Experiment 2: Development of a fedora-type MCA device***

***S*urvival rate**, **expulsion time, and bursting pressure:** The surgical procedures went well for all of the different fedora-type MCA devices used. After the operation, X-ray fluoroscopy showed that the daughter and parent parts for all the fedora-type MCA devices were tightly coupled. There was no significant difference in the survival rates (groups 2.1-2.3: 93.33%, 100%, and 93.33%, respectively, *P* = 0.434) or expulsion time (groups 2.1-2.3: 3 (IQR 3-3.25), 4 (IQR 2-5), and 4 (IQR 3-5) d, respectively, *P* = 0.175) between different fedora-type MCA devices. Additionally, there was no significant difference in the bursting pressure based on the different fedora-type MCA devices used (Table 2).

**Size of anastomosis:** On the 30th, 90th, and 180th postoperative days, the larger fedora-type MCA devices had a larger anastomosis circumference (*P* < 0.017 for all, adjusted α = 0.017) (Table 2). Based on the findings from the former experiment, the circumferences of the anastomoses in the smaller fedora-type MCA device (group 2.1) decreased as time progressed (8.04 ± 0.62 mm, 5.36 ± 0.32 mm, and 2.45 ± 0.67 mm for postoperative days 30, 90, and 180, respectively; *P* < 0.017 for all, adjusted α = 0.017), and the stomas were nearly closed by the 180th postoperative day (Figure 5A1-A3). There were no significant differences in the circumference at the different postoperative time points for the large fedora-type MCA devices (group 2.2: *P* = 0.749; group 2.3: *P* = 0.712) (Figure 5B1-B3 and C1-C3).

**Morphological analysis:** On the 30th postoperative day, the gross appearance of anastomoses in all groups did not significantly differ. The anastomoses were clean and intact for all designs of the fedora-type MCA devices on the 30th postoperative day, and all the mucosae were smooth and flat, without any ulcers or erosions (Figure 5A4, A5, B4, B5, C4, and C5). As shown in Table 2, the difference in the adhesion score between the groups was not significant (*P* = 0.985). The HE and Masson’s trichrome staining in all groups showed that the serosal, submucosal, and mucosal layers were continuous (Figure 4C1 and C2).

**DISCUSSION**

Although previous studies have confirmed the feasibility of MCA in animal experiments[24-26] and clinical practice[27-29], there is still a risk of anastomotic stenosis or even closure in the long run after MCA[15-17]. One interesting finding regarding MCA is the correlation between the size of anastomosis and the MCA device. Therefore, in the current study, traditional nummular MCA devices with different sizes were used to explore the suitable size and pressure for MCA. However, for traditional MCA devices, as the pressure increased, the size also increased[22]. Larger MCA devices increased the risk of leakage; therefore, we developed a novel “fedora-type” MCA device to allow for a large size but low pressure. Each part of the fedora-type MCA device had a nummular magnet with a larger sheet metal. After comparison, the optimal design for the fedora-type MCA device was that with a Φ4-mm nummular magnet and a Φ6- mm sheet metal.

The anastomat influenced the outcome of MCA in terms of pressure and size. The pressure affects the ischemic necrosis speed of the compressed tissue. If this speed surpasses the healing of anastomotic tissue, leakage could occur[21,22]. However, if the pressure is too low, dissociation of the MCA device might occur[10]. Furthermore, the importance of size is embodied in the following two aspects. First, if the size is too small, the anastomosis would narrow or even close with time; this is perhaps due to the insufficient shunt. Conversely, if the size is too large, placement and discharge of the anastomat will be difficult[30]. Thus, pressure influenced the short-term outcome of anastomotic formation for MCA, while size influenced the long-term outcome of anastomotic stenosis or closure for MCA. The existing limited basic work regarding MCA devices has mostly been focused on the effect of pressure, with a relatively short-term follow-up period (no more than 3 mo)[10,21,22]. These previous studies have ignored the importance of the size, which required subgroups and long-term follow-up. However, anastomotic stenosis or closure was identified as the real challenge for MCA devices in the gut[15,16,20].

To our knowledge, this is the first study to simultaneously explore the optimal size and pressure of traditional nummular MCA devices for intestinal anastomosis in the rat model, with a 6-mo follow-up period. The rat model simplified the subgroups. Thus, all sizes of traditional MCA devices were explored; this was crucial to investigate the relationship between anastomat, gut sizes, and anastomotic stenosis. This study showed that 5-6 mm was the optimal size range for ileac side-to-side anastomosis in the rat model. When the size was smaller than 5 mm, the anastomosis formed was small, and anastomotic stenosis or closure occurred in the long-term follow-up. While the size reached up to 7 mm, it was difficult to insert it into the intestine. In the current model, the mean (± SD) circumference of the intestine was 13.34 ± 0.12 mm, meaning that the diameter was approximately 4.2 mm. Thus, we speculated that the size of the MCA device should be larger than 120% of the enteric diameter, otherwise the anastomosis stoma would not receive sufficient shunt. This would result in stenosis or closure in the long-term follow-up. Unfortunately, the size was only approximately 58%-66% of the enteric diameter in a previously published study[6-8]. This study also demonstrated that 54.56 ± 1.40 kPa to 126.07 ± 1.38 kPa was the optimal compression pressure range, in accordance with previously published studies.

Although we determined the optimal size and pressure, they were almost impossibly achieved by traditional MCA devices, which were either of large or small size and achieved high or low pressure, respectively. The high pressure increased the risk of leakage, while the small size caused anastomotic stenosis or even closure. Devices that were large in size and led to a low amount of pressure were the ideal design for MCA devices in the gut. Therefore, we developed a novel MCA device to meet these parameters, which we called a “fedora-type” MCA device. Both parts of the novel anastomat consisted of a nummular magnet and a larger sheet metal. This allowed for control of the compression pressure by adjustment of the magnet, and for optimal size by allowing for the sheet metal to be changed. The novel design broke the internal connection between size and compression pressure in MCA devices and allowed for a large size and low pressure. Of all the different designs for the fedora-type MCA device used, the Φ4-mm nummular magnet with a Φ6-mm sheet metal could safely form anastomosis after operation and ensure long-term stability. It should be noted that the pressure produced by this design was almost the same as that of the Φ3-mm traditional nummular MCA device, which was the smallest one used in the first experiment in this study (54.56 ± 1.40 kPa *vs* 56.03 ± 0.61 kPa). However, the circumference of anastomosis at 6 mo was comparable to that of the Φ6- mm traditional MCA device (16.33 ± 0.37 mm *vs* 16.42 ± 0.31 mm, *P* = 0.893). This confirmed that the anastomotic stenosis was associated with the size of the MCA device, instead of the pressure.

This study was subject to several limitations that merit consideration. These results are only applicable to rats; models in larger animals and further clinical trials are needed to test this hypothesis and guide clinical application. Although some results of the current work cannot be directly translated into clinical practice, such as the size of MCA device, other results would provide important guidance for further clinical application. For example, with an adequate number of animals, we demonstrated that the diameter of MCA device should be greater than 120% of the enteric diameter to ensure the stability of intestinal anastomosis. In this study, the anastomotic specimens at postoperative days 30, 90, and 180 were analyzed. The anastomotic specimens from a longer follow-up duration might be more convincing. However, we suspect that if the anastomosis remained stable for 6 mo, stenosis would rarely occur.

**CONCLUSION**

To address some of the deficiencies in the current MCA model, we explored the optimal size and pressure of the MCA device for intestinal anastomosis in rats. We found that the suggested diameter of the MCA device should be larger than 120% of the enteric diameter to avoid stenosis. Then, we developed a novel “fedora-type” MCA device for the current model, using a Φ4-mm nummular magnet with a Φ6-mm sheet metal. This model safely formed anastomosis and ensured long-term anastomosis. This novel anastomat controlled pressure and optimized the size, thus meeting our stipulated requirements for a large size and small force device.

**ARTICLE HIGHLIGHTS**

***Research background***

The feasibility of magnetic compression anastomosis (MCA) has been confirmed by previous studies; however, there is still a risk of long-term anastomotic stenosis. In fact, anastomat influences the outcome of MCA in terms of pressure and size. High pressure increases the risk of leakage, while small size causes anastomotic stenosis or even closure. One defect of traditional MCA lies in the correlation between the size of anastomosis and the MCA device. For traditional MCA devices, a large size has represented large pressure, eventually leading to increased leakage, meaning “large size & large force”.

***Research motivation***

Studies have shown that there is a risk of long-term anastomotic stenosis and even closure after MCA; this has restricted further clinical application of MCA.

***Research objectives***

This study aimed to explore the optimal size and pressure of the MCA device for intestinal anastomosis in rats. Thereafter, a novel MCA device (“fedora-type” MCA device) was developed to simultaneously meet the requirements of pressure and size.

***Research methods***

We designed the following two experiments. First, based on the anatomical characteristics of rat intestines, we used traditional nummular MCA devices with all possible sizes to conduct ileac side-to-side anastomosis. Based on the short-term results, we determined the appropriate pressure range required for MCA. According to the long-term results, we confirmed the minimum size required to avoid anastomotic stenosis or closure. Second, based on the results of the former experiment, we introduced a novel design concept, referred to as the “fedora-type” MCA device, to simultaneously meet the requirements of both pressure and size, so that stable anastomosis could be formed.

***Research results***

The optimal size range was 5-6 mm for ileac side-to-side anastomosis in the rat model (the diameter of the MCA device should be within 120%-140% of the enteric diameter). When the size was smaller than 5 mm, anastomotic stenosis or closure occurred. This study also demonstrated that 54.56 ± 1.40 kPa to 126.07 ± 1.38 kPa was the optimal compression pressure range. Traditional MCA cannot meet both of these requirements. This newly developed “fedora-type” MCA device consisted of a nummular magnet and a larger sheet metal. This allowed for control of the compression pressure by adjustment of the magnet, and for optimal size by allowing for the sheet metal to be changed. The novel design broke the internal connection between size and compression pressure in MCA devices and allowed for a large size and low pressure. Of all the different designs for the fedora-type MCA device used, the Φ4 mm nummular magnet with a Φ6 mm sheet metal could safely form anastomosis after operation and ensure long-term stability.

***Research conclusions***

The diameter of the MCA device should be larger than 120% of the enteric diameter to avoid stenosis. This novel anastomat controlled pressure and optimized the size respectively, thus meeting our stipulated requirements for a large size and small force device. The “fedora-type” MCA device for this model, using a Φ4 mm nummular magnet with a Φ6 mm sheet metal, safely formed anastomosis and ensured long-term anastomosis.

***Research perspectives***

Models in larger animals and further clinical trials are needed to test this hypothesis and guide clinical application.

**ACKNOWLEDGEMENTS**

The authors thank all the staff at the National Local Joint Engineering Research Center for Precision Surgery & Regenerative Medicine for their help and provision of facilities to conduct this study.

**REFERENCES**

1 **Obora Y**, Tamaki N, Matsumoto S. [Nonsuture microvascular anastomosis using magnet rings (author's transl)]. *Neurol Med Chir (Tokyo)* 1980; **20**: 497-505 [PMID: 6157124 DOI: 10.2176/nmc.20.497]

2 **Lv Y**, Shi Y; Scientific Committee of the First International Conference of Magnetic Surgery. Xi'an consensus on magnetic surgery. *Hepatobiliary Surg Nutr* 2019; **8**: 177-178 [PMID: 31098373 DOI: 10.21037/hbsn.2019.03.01]

3 **Noh M**, Mooney DP, Trumper DL. Magnet-Assisted Hydraulic Bougienage for Correction of Long-Gap Esophageal Atresia. *IEEE Trans Biomed Eng* 2018; **65**: 2178-2189 [PMID: 29989954 DOI: 10.1109/TBME.2017.2786733]

4 **Oetzmann von Sochaczewski C**, Lindner A, Heimann A, Balus A, Patel VH, Harrison MR, Muensterer OJ. Beyond Magnamosis: A Method to Test Sutureless Esophageal Anastomotic Devices in Living Swine by Creating an Esophageal Bypass Loop for Natural Oral Nutrition. *J Laparoendosc Adv Surg Tech A* 2019; **29**: 852-855 [PMID: 30882275 DOI: 10.1089/lap.2018.0778]

5 **Muensterer OJ**, Sterlin A, Oetzmann von Sochaczewski C, Lindner A, Heimann A, Balus A, Dickmann J, Nuber M, Patel VH, Manfredi MA, Jennings RW, Smithers CJ, Fauza DO, Harrison MR. An experimental study on magnetic esophageal compression anastomosis in piglets. *J Pediatr Surg* 2020; **55**: 425-432 [PMID: 31128845 DOI: 10.1016/j.jpedsurg.2019.04.029]

6 **Lebares CC**, Graves CE, Lin MY, Fidelman N, Cello J, Harrison MR, Rogers S. Endoscopic Magnetic Compression Anastomosis For Small Bowel Bypass in a High Operative Risk Setting. *Surg Laparosc Endosc Percutan Tech* 2019; **29**: e84-e87 [PMID: 31107851 DOI: 10.1097/SLE.0000000000000669]

7 **Toselli L**, Martinez-Ferro M, Cervio G, Kwiat D, Imamura-Ching J, Graves CE, Gaston B, Harrison M. Magnetic Compression Anastomosis (Magnamosis) for Functional Undiversion of Ileostomy in Pediatric Patients. *J Laparoendosc Adv Surg Tech A* 2017; **27**: 1314-1317 [PMID: 28976806 DOI: 10.1089/lap.2017.0300]

8 **Machytka E**, Bužga M, Zonca P, Lautz DB, Ryou M, Simonson DC, Thompson CC. Partial jejunal diversion using an incisionless magnetic anastomosis system: 1-year interim results in patients with obesity and diabetes. *Gastrointest Endosc* 2017; **86**: 904-912 [PMID: 28716404 DOI: 10.1016/j.gie.2017.07.009]

9 **An Y**, Zhang Y, Liu H, Ma S, Fu S, Lv Y, Yan X. Gastrojejunal anastomosis in rats using the magnetic compression technique. *Sci Rep* 2018; **8**: 11620 [PMID: 30072707 DOI: 10.1038/s41598-018-30075-8]

10 **Lambe T**, Ríordáin MG, Cahill RA, Cantillon-Murphy P. Magnetic compression in gastrointestinal and bilioenteric anastomosis: how much force? *Surg Innov* 2014; **21**: 65-73 [PMID: 23592733 DOI: 10.1177/1553350613484824]

11 **Diaz R**, Davalos G, Welsh LK, Portenier D, Guerron AD. Use of magnets in gastrointestinal surgery. *Surg Endosc* 2019; **33**: 1721-1730 [PMID: 30805789 DOI: 10.1007/s00464-019-06718-w]

12 **Kawabata H**, Hitomi M, Inoue N, Kawakatsu Y, Okazaki Y, Miyata M. Intraductal Ultrasonography as a Local Assessment Before Magnetic Compression Anastomosis for Obstructed Choledochojejunostomy. *Gastroenterology Res* 2017; **10**: 255-258 [PMID: 28912914 DOI: 10.14740/gr842w]

13 **Saito R**, Tahara H, Shimizu S, Ohira M, Ide K, Ishiyama K, Kobayashi T, Ohdan H. Biliary-duodenal anastomosis using magnetic compression following massive resection of small intestine due to strangulated ileus after living donor liver transplantation: a case report. *Surg Case Rep* 2017; **3**: 73 [PMID: 28547740 DOI: 10.1186/s40792-017-0349-4]

14 **Matsuura R**, Ueno T, Tazuke Y, Tanaka N, Yamanaka H, Takama Y, Nakahata K, Yamamichi T, Maeda N, Osuga K, Yamanouchi E, Okuyama H. Magnetic compression anastomosis for postoperative biliary atresia. *Pediatr Int* 2017; **59**: 737-739 [PMID: 28626977 DOI: 10.1111/ped.13295]

15 **Liu XM**, Li Y, Xiang JX, Ma F, Lu Q, Guo YG, Yan XP, Wang B, Zhang XF, Lv Y. Magnetic compression anastomosis for biliojejunostomy and pancreaticojejunostomy in Whipple's procedure: An initial clinical study. *J Gastroenterol Hepatol* 2019; **34**: 589-594 [PMID: 30278106 DOI: 10.1111/jgh.14500]

16 **Qiao W**, Shi A, Ma F, Yan X, Duan J, Wu R, Li D, Lv Y. Further Development of Magnetic Compression for Gastrojejunostomy in Rabbits. *J Surg Res* 2020; **245**: 249-256 [PMID: 31421370 DOI: 10.1016/j.jss.2019.07.078]

17 **Ellebaek MBB**, Qvist N, Rasmussen L. Magnetic Compression Anastomosis in Long-Gap Esophageal Atresia Gross Type A: A Case Report. *European J Pediatr Surg Rep* 2018; **6**: e37-e39 [PMID: 29796381 DOI: 10.1055/s-0038-1649489]

18 **Woo R**, Wong CM, Trimble Z, Puapong D, Koehler S, Miller S, Johnson S. Magnetic Compression Stricturoplasty For Treatment of Refractory Esophageal Strictures in Children: Technique and Lessons Learned. *Surg Innov* 2017; **24**: 432-439 [PMID: 28745145 DOI: 10.1177/1553350617720994]

19 **Fan C**, Zhang H, Yan X, Ma J, Wang C, Lv Y. Advanced Roux-en-Y hepaticojejunostomy with magnetic compressive anastomats in obstructive jaundice dog models. *Surg Endosc* 2018; **32**: 779-789 [PMID: 28779259 DOI: 10.1007/s00464-017-5740-5]

20 **Liu XM**, Yan XP, Zhang HK, Ma F, Guo YG, Fan C, Wang SP, Shi AH, Wang B, Wang HH, Li JH, Zhang XG, Wu R, Zhang XF, Lv Y. Magnetic Anastomosis for Biliojejunostomy: First Prospective Clinical Trial. *World J Surg* 2018; **42**: 4039-4045 [PMID: 29947988 DOI: 10.1007/s00268-018-4710-y]

21 **Zhao G**, Ma J, Yan X, Li J, Ma F, Wang H, Liu Y, Lv Y. Optimized force range of magnetic compression anastomosis in dog intestinal tissue. *J Pediatr Surg* 2019; **54**: 2166-2171 [PMID: 30929946 DOI: 10.1016/j.jpedsurg.2019.03.005]

22 **Xue F**, Guo HC, Li JP, Lu JW, Wang HH, Ma F, Liu YX, Lv Y. Choledochojejunostomy with an innovative magnetic compressive anastomosis: How to determine optimal pressure? *World J Gastroenterol* 2016; **22**: 2326-2335 [PMID: 26900294 DOI: 10.3748/wjg.v22.i7.2326]

23 **Dziki AJ**, Duncan MD, Harmon JW, Saini N, Malthaner RA, Trad KS, Fernicola MT, Hakki F, Ugarte RM. Advantages of handsewn over stapled bowel anastomosis. *Dis Colon Rectum* 1991; **34**: 442-448 [PMID: 1953849 DOI: 10.1007/BF02049926]

24 **Bai J**, Huo X, Ma J, Lv Y, Yan X. Magnetic compression technique for colonic anastomosis in rats. *J Surg Res* 2018; **231**: 24-29 [PMID: 30278935 DOI: 10.1016/j.jss.2018.05.006]

25 **Gao Y**, Wu RQ, Lv Y, Yan XP. Novel magnetic compression technique for establishment of a canine model of tracheoesophageal fistula. *World J Gastroenterol* 2019; **25**: 4213-4221 [PMID: 31435174 DOI: 10.3748/wjg.v25.i30.4213]

26 **Bruns NE**, Glenn IC, Craner DR, Schomisch SJ, Harrison MR, Ponsky TA. Magnetic compression anastomosis (magnamosis) in a porcine esophagus: Proof of concept for potential application in esophageal atresia. *J Pediatr Surg* 2019; **54**: 429-433 [PMID: 30309731 DOI: 10.1016/j.jpedsurg.2018.09.014]

27 **Li Y**, Sun H, Yan X, Wang S, Dong D, Liu X, Wang B, Su M, Lv Y. Magnetic compression anastomosis for the treatment of benign biliary strictures: a clinical study from China. *Surg Endosc* 2020; **34**: 2541-2550 [PMID: 31399950 DOI: 10.1007/s00464-019-07063-8]

28 **Kubo M**, Wada H, Eguchi H, Gotoh K, Iwagami Y, Yamada D, Akita H, Asaoka T, Noda T, Kobayashi S, Nakamura M, Ono Y, Osuga K, Yamanouchi E, Doki Y, Mori M. Magnetic compression anastomosis for the complete dehiscence of hepaticojejunostomy in a patient after living-donor liver transplantation. *Surg Case Rep* 2018; **4**: 95 [PMID: 30112678 DOI: 10.1186/s40792-018-0504-6]

29 **Parlak E**, Eminler AT, Koksal AS, Toka B, Uslan MI, Sokmensuer C, Guven M. A new method for lumen restoration in a patient with aphagia: Oro-oesophageal through-the-scope magnetic compression anastomosis. *Clin Otolaryngol* 2019; **44**: 1214-1217 [PMID: 30968566 DOI: 10.1111/coa.13337]

30 **Zhao G**, Yan X, Ma L, Liu W, Zhang J, Guo H, Liu Y, Lv Y. Biomechanical and Performance Evaluation of Magnetic Elliptical-Ring Compressive Anastomoses. *J Surg Res* 2019; **239**: 52-59 [PMID: 30802705 DOI: 10.1016/j.jss.2019.01.063]

**Footnotes**

**Institutional review board statement:** The study was reviewed and approved by the Ethics Committee of the First Affiliated Hospital of Xi'an Jiaotong University (Permit Number: XJTU1AF2015LSL-046).

**Institutional animal care and use committee statement:** All experimental protocols were approved by the Committee on the Ethics of Animal Experiments of Xi'an Jiaotong University (Permit Number: XJTULAC2020-1281).

**Conflict-of-interest statement:** All authors declare no conflicts of interest related to this article.

**Data sharing statement:** No additional data are available.

**ARRIVE guidelines statement:** The authors have read the ARRIVE guidelines, and the manuscript was prepared and revised according to the ARRIVE guidelines.

**Open-Access:** This article is an open-access article that was selected by an in-house editor and fully peer-reviewed by external reviewers. It is distributed in accordance with the Creative Commons Attribution NonCommercial (CC BY-NC 4.0) license, which permits others to distribute, remix, adapt, build upon this work non-commercially, and license their derivative works on different terms, provided the original work is properly cited and the use is non-commercial. See: http://creativecommons.org/Licenses/by-nc/4.0/

**Manuscript source:** Unsolicited manuscript

**Peer-review started:** August 12, 2020

**First decision:** August 22, 2020

**Article in press:** September 10, 2020

**Specialty type:** Gastroenterology and hepatology

**Country/Territory of origin:** China

**Peer-review report’s scientific quality classification**

Grade A (Excellent): 0

Grade B (Very good): B


Grade C (Good): 0

Grade D (Fair): 0


Grade E (Poor): 0

**P-Reviewer:** Harrison MR **S-Editor:** Yan JP **L-Editor:** Wang TQ **P-Editor:** Li JH


**Figure Legends**



**Figure 1 Magnetic compression anastomosis devices.** A: Traditional nummular magnetic compression anastomosis (MCA) devices of different sizes used in experiment 1; B: Fedora-type MCA devices with different design used in experiment 2; C: Schematic diagram of the fedora-type MCA device.




**Figure 2 Surgical procedure and X-ray fluoroscopy.** A: The small intestine was removed; B: A 6 mm incision was made 12 cm distal to the cecum (black arrow); C: The daughter part (yellow arrow) was inserted; D: The parent part (blue arrow) was inserted; E1: Two magnets of the traditional nummular magnetic compression anastomosis (MCA) device were coupled (purple arrow) to compress the ileac wall; E2: Accurate coupling of the daughter and parent magnets in experiment 1 was confirmed using X-ray; F1: Two parts of the fedora-type MCA device were coupled (gray arrow); F2: Accurate coupling of the daughter and parent parts in experiment 2 was confirmed using X-ray.




**Figure 3 Gross appearance of anastomoses using traditional nummular magnetic compression anastomosis devices.** A: Group 1.1 (Φ3 mm): The size of anastomosis 30 d after magnetic compression anastomosis (MCA) (A1), the size of anastomosis 180 d after MCA (A2), the change in anastomosis circumferences after MCA (A3), serosa side of anastomosis (A4), and mucosa side of anastomosis (A5); B: Group 1.2 (Φ4 mm): The size of anastomosis 30 d after MCA (B1), the size of anastomosis 180 d after MCA (B2), the change in anastomosis circumferences after MCA (B3), serosa side of anastomosis (B4), and mucosa side of anastomosis (B5); C: Group 1.3 (Φ5 mm): The size of anastomosis 30 d after MCA (C1), the size of anastomosis 180 d after MCA (C2), the change in anastomosis circumferences after MCA (C3), serosa side of anastomosis (C4), and mucosa side of anastomosis (C5); D: Group 1.4 (Φ6 mm): The size of anastomosis 30 d after MCA (D1), the size of anastomosis 180 d after MCA (D2), the change in anastomosis circumferences after MCA (D3), serosa side of anastomosis (D4), and mucosa side of anastomosis (D5). Yellow arrows: Anastomosis; blue arrows: Anastomotic line.

**A B C**



**Figure 4 Microscopic appearance of anastomosis**. A: Larger size groups of traditional nummular magnetic compression anastomosis (MCA) devices (Group 1.3 and 1.4): Hematoxylin and eosin staining (A1) and Masson’s trichrome staining (A2); B: Smaller size groups of traditional nummular MCA devices (Group 1.1 and 1.2): Hematoxylin and eosin staining (B1) and Masson’s trichrome staining (B2); C: Fedora-type MCA device: Hematoxylin and eosin staining (C1) and Masson’s trichrome staining (C2).



**Figure 5 Gross appearance of anastomoses using the fedora-type magnetic compression anastomosis devices.** A: Group 2.1 (with a Φ4 mm sheet metal): The size of anastomosis 30 d after magnetic compression anastomosis (MCA) (A1), the size of anastomosis 180 d after MCA (A2), the change in anastomosis circumferences after MCA (A3), serosa side of anastomosis (A4), and mucosa side of anastomosis (A5); B: Group 2.2 (with a Φ5 mm sheet metal): The size of anastomosis 30 d after MCA (B1), the size of anastomosis 180 d after MCA (B2), the change in anastomosis circumferences after MCA (B3), serosa side of anastomosis (B4), and mucosa side of anastomosis (B5); C: Group 2.3 (with a Φ6 mm sheet metal): The size of anastomosis 30 d after MCA (C1), the size of anastomosis 180 d after MCA (C2), the change in anastomosis circumferences after MCA (C3), serosa side of anastomosis (C4), and mucosa side of anastomosis (C5). Yellow arrows: Anastomosis; blue arrows: Anastomotic line.

**Table 1 Results of traditional nummular magnetic compression anastomosis devices with different sizes**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | **Group 1.1** | **Group 1.2** | **Group 1.3** | **Group 1.4** | ***P*** **value** |
| Survival rate | 93.3% (14/15) | 100% (15/15) | 73.3% (11/15) | 73.3 (11/15) | 0.083 |
| Discharge time (d) | 3 (IQR 3-4) | 3 (IQR 3-4) | 2 (IQR 1-4) | 2.5 (IQR 2-3) | 0.002 |
| Adhesion score | | | | | |
| 0 | 92.9% (13/14) | 86.7% (13/15) | 27.3% (3/11) | 27.3% (3/11) | < 0.001 |
| 1 | 7.1% (1/14) | 13.3% (2/15) | 27.3% (3/11) | 9.1 (1/11) |
| 2 | 0 (0/14) | 0 (0/15) | 9.1% (1/11) | 27.3% (3/11) |
| 3 | 0 (0/14) | 0 (0/15) | 18.2% (2/11) | 18.2% (2/11) |
| 4 | 0 (0/14) | 0 (0/15) | 18.2% (2/11) | 18.2% (2/11) |
| Circumference of anastomotic stomas (mm) | | | | |  |
| 30 d | 2.47 ± 0.18 | 8.84 ± 0.31 | 13.54 ± 0.31 | 15.98 ± 0.73 | < 0.001 |
| 90 d | 1.20 ± 0.18 | 5.90 ± 0.27 | 13.73 ± 0.49 | 16.43 ± 0.30 | < 0.001 |
| 180 d | 0.35 ± 0.19 | 2.07 ± 0.37 | 13.24 ± 0.68 | 16.33 ± 0.37 | < 0.001 |
| Bursting pressure (mmHg) | | | | |  |
| 30 d | 247.64 ± 10.78 | 245.18 ± 7.77 | 242.90 ± 11.56 | 205.725 ± 8.06 | 0.032 |
| 90 d | 264.55 ± 7.87 | 269.46 ± 9.30 | 261.47 ± 9.72 | 256.03 ± 15.63 | 0.830 |
| 180 d | 263.32 ± 10.85 | 258.62 ± 10.19 | 261.08 ± 12.06 | 265.05 ± 11.26 | 0.978 |

**Table 2 Results of fedora-type magnetic compression anastomosis devices with different designs**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | **Group 2.1** | **Group 2.2** | **Group 2.3** | ***P* value** |
| Survival rate | 93.3% (14/15) | 100% (15/15) | 93.3% (14/15) | 0.434 |
| Discharge time (d) | 3 (IQR 3-3.25) | 4 (IQR 2-5) | 4 (IQR 3-5) | 0.175 |
| Adhesion score | | | | |
| 0 | 85.7% (12/14) | 86.7% (13/15) | 85.7% (12/14) | 0.985 |
| 1 | 7.1% (1/14) | 13.3% (2/15) | 7.1% (1/14) |
| 2 | 7.1% (1/14) | 0 (0/15) | 0 (0/14) |
| 3 | 0 (0/14) | 0 (0/15) | 7.1% (1/14) |
| 4 | 0 (0/14) | 0 (0/15) | 0 (0/14) |
| Circumference of anastomotic stomas (mm) | | | |  |
| 30 d | 8.04 ± 0.62 | 13.10 ± 0.43 | 15.85 ± 0.47 | < 0.001 |
| 90 d | 5.36 ± 0.32 | 13.56 ± 0.58 | 16.20 ± 0.52 | < 0.001 |
| 180 d | 2.45 ± 0.67 | 13.57 ± 0.47 | 16.42 ± 0.31 | < 0.001 |
| Bursting pressure (mmHg) | | | |  |
| 30 d | 242.80 ± 8.90 | 239.32 ± 9.18 | 250.88 ± 7.71 | 0.634 |
| 90 d | 259.14 ± 7.42 | 267.00 ± 9.38 | 261.14 ± 12.01 | 0.842 |
| 180 d | 258.35 ± 14.46 | 260.82 ± 11.78 | 265.85 ± 14.07 | 0.972 |