**Name of journal: World Journal of Cardiology**

**ESPS Manuscript NO: 8296**

**Columns: Review**

Newer methods of cardiac output monitoring

Mehta Y *et al*. Cardiac output monitoring

Yatin Mehta, Dheeraj Arora

**Yatin Mehta, Dheeraj Arora,** Institute of Critical Care and Anesthesiology, Medanta The Medicity, Haryana 122001, India

**Author contributions:** Both authors mutually contributed to.

**Correspondence to: Yatin Mehta, MD, MNAMS, FRCA, Chairman, I**nstitute of Critical Care and Anesthesiology, Medanta The Medicity, Sector 38, Gurgaon, Haryana 122001, India. [yatinmehta@hotmail.com](mailto:yatinmehta@hotmail.com)

**Telephone:** +91-124-4141414

**Received:** December 23, 2013 **Revised:** May 7, 2014

**Accepted:** July 15, 2014

**Published online:**

**Abstract**

Cardiac output (CO) is the volume of blood ejected by each ventricle per minute and is the product of stroke volume and heart rate. CO can thus be manipulated by alteration in heart rate or rhythm, preload, contractility and afterload. Moreover it gives important information about tissue perfusion and oxygen delivery. CO can be measured by various methods and thermodilution method using pulmonary artery catheter (PAC) is till date considered as gold standard method. Complications associated with PAC led to development of newer methods which are minimally or non-invasive. Newer methods fulfil other properties like continuous and reproducible reading, cost effective, reliable during various physiological states and have fast response time. These methods are validated against the gold standard with good level agreement. In this review we have discussed various newer methods of CO monitoring and their effectiveness in clinical use.

© 2014 Baishideng Publishing Group Inc. All rights reserved.

**Key words:** Cardiac Output; Pulse contour analysis; Pulse power analysis; Bioimpedance; Doppler; Echocardiography

**Core tip:** This is review of newer methods of cardiac output monitoring which are minimally invasive and have lesser complications as compared to gold standard methods.

Mehta Y, Arora D. Newer methods of cardiac output monitoring. *World J Cardiol* 2014; In press

**INTRODUCTION**

Cardiac output (CO) monitoring is an important tool in high risk critically ill surgical patients in whom large fluid shifts are expected along with bleeding and hemodynamic instability.It is an important component of goal directed therapy (GDT) *i.e.,* when a monitor is used in conjunction with administration of fluids and vasopressors to achieve set therapeutic endpoints thereby improving patient care and outcome. CO cannot be measured reliably by clinical examination and routine assessment. There are various methods of CO monitoring based on Ficks principle, thermodilution, Doppler, pulse contour analysisand bioimpedance. Each method has its own merits and demerits (Table 1). An ideal CO monitor should be minimally or non-invasive, continuous, cost effective, reproducible, reliable during various physiological states and have fast response time[1]. Advances in the computer software and hardware have led to development of newer methods of CO monitoring with minimal or no vascular access.

Methods of CO monitoring are broadly classified as follows: (1) Invasive- Intermittent bolus pulmonary artery thermodilution, Continuous pulmonary artery thermodilution; (2) Minimally invasive-Lithium dilution (LiDCO), Pulse contour analysis (PiCCO and FloTrac), Esophgeal Doppler (ED), transesophgeal echocardiography (TEE); and (3) Non-invasive-Partial gas rebreathing, Thoracic bioimpedance and bioreactance, ECOM (endotracheal cardiac output monitor), Doppler method and Photoelectric plethysmography.

**INVASIVE METHODS**

***Cardiac output measurement by pulmonary artery catheter***

Pulmonary artery catheter (PAC) as a monitor to measure flow and pressure was developed by Dexter[2] and modified later on by Swan[3] to measure CO and central filling pressures. It is still considered as gold standard monitor to measure CO since 1970’s[4].It has been used as a monitoring tool in high risk surgeries and critical care units.

However, its use has been associated with various complications like pneumothorax, arrhythmia, infection, pulmonary artery rupture, valve injury, knotting and thrombosis leading to embolism[5,6].Also, various technical errors may lead to false readings like loss of injectate, variability of temperature, thermistor malfunction, clot over catheter tip, coiling of catheter or timing of injectate > 4 s. Moreover, intracardiac shunts, mechanical ventilation or valvular dysfunction may lead to incorrect readings. These errors and adverse effects led to the development of less invasive methods of CO monitoring[7,8].Thus the main objective of present review article is to focus on the newer methods of CO monitoring that are validated with the gold standard method and have ease of use and lesser complications.

**CONTINUOUS CO MEASUREMENT BY PAC**

Continuous CO (CCO, Edwards Lifesciences, Irvine, California, United States) is a modification of PAC with copper filament in the catheter that remains in the right ventricle. There is intermittent heating of blood in the right heart by the filament and the resultant signal is captured by thermistor near the tip of the catheter. Average value of CO measured over time is displayed on the monitor. Main advantages of CCO over conventional PAC are avoidance of repeated boluses thus reducing the infection risk and operator errors[5].Moreover, continuous monitoring of stroke volume (SV), systemic vascular resistance (SVR) and mixed venous saturation (Svo2) can also be performed with this catheter. We found CCO to be comparable to conventional intermittent thermodilution CO in patients undergoing off pump coronary artery bypass grafting surgery (OPCAB) at various time points[9].

Literature review regarding use of PAC in operating room and intensive care units (ICU) revealed both benefits and risks. Gore *et al*[10] showed that PAC use increased mortality after myocardial infarctionand SUPPORT trial also showed increased mortalityat 30 d[11].Complications have led some authors to call for complete moratorium on PAC use[12]. Various randomized controlled trials (RCT) also demonstrated increased incidence of adverse events in comparison to central venous pressure (CVP) 1.5% *vs* 0.7% with no significant difference in mortality and length of stay in hospital[13].Later on PAC-MAN trial failed to show any benefit or harm with the use of PAC[14].Its use in patients undergoing OPCAB also showed no difference in mortality and final outcome[15].ESCAPE trial demonstrated functional improvement with PAC guided therapy used in patients with congestive heart failure[16].

Inspite of various arguments PAC is still considered as the “Gold Standard” for monitoring of CO. However, due to inherent risk associated with its use investigators are trying to develop a minimally or non-invasive monitor for CO which has all the characteristics of an ideal monitor. Various methods based on arterial pulse contour analysis, plethysmography, Fick’s principle or bioimpedance have been developed. Its values should be within limits of agreement (Bland Altman analysis)[17] of the “gold standard”. We will discuss these methods in the present review.

**MINIMALLY INVASIVE METHODS**

***Pulse power analysis***

This method is based on the principle that change of the blood pressure about the mean is directly related to the SV. Various factors affect its accuracy like compliance of the arterial tree, wave reflection, damping of the transducer and aortic systolic outflow[18].

Lithium dilution CO (LiDCO, Cambridge, United Kingdom) system combines pulse contour analysis with lithium indicator dilution for continuous monitoring of SV and stroke volume variation (SVV). Root mean square method is applied to the arterial pressure signal and called “nominal SV” and using a patient specific calibration factor is further scaled to an “actual SV”[19]. It is a minimally invasive technique first described in 1993[18] and requires a venous (central or peripheral) line and an arterial catheter. A bolus of lithium chloride is injected into venous line and arterial concentration is measured by withdrawing blood across disposable lithium sensitive sensor containing an ionophor selectively permeable to Li. CO is calculated based on Li dose and area according to the concentration time circulation[20].

It requires calibration every 8 h and during major hemodynamic changes. It is contraindicated in patients on Li therapy and calibration is also affected by neuromuscular blockers as quaternary ammonium residue causes electrode to drift[20].Its accuracy is affected by aortic regurgitation, intraaortic balloon pump (IABP), damped arterial line, postaortic surgery, arrhythmia and intra or extracardiac shunts[5,20].

This device has been studied in relation with PAC. Linton *et al*[18] found good correlation with PAC.Good correlation with PAC has also been found in patients undergoing liver transplantation[21].Pearse *et al*[22] studied it for early goal directed therapy (EGDT) and revealed fewer complications and shorter length of hospital stay.

***Pulse contour analysis***

It is based on the principle that area under the systolic part of the arterial pressure waveform is proportional to the SV[23].It was first described by Erlanger and Hooker in 1904 and suggested that CO was proportional to arterial pulse pressure[24].In this method the area is measured post diastole to end of ejection phase divided by aortic impedance that measures SV. It also measures SVV and pulse pressure variation (PVV) which is useful in predicting fluid responsiveness. SVV is the difference between maximum and minimum SV over the respiratory cycle and is caused by changes in preload with alteration in intrathoracic pressure. In addition to that shape of the arterial waveform (dP/dt), arterial compliance, SVR and patient specific calibration factors are also required for calibration[24].In 1970’s first algorithm was developed to continuously analyse the pressure waveform from arterial line[25].

**PICCO system:** The PiCCO system (PULSION medical system, Munich, Germany) was the first pulse contour device introduced and was replaced with PiCCO2 in 2007[26].It requires both central venous (femoral or internal jugular) and arterial cannulation (femoral/radial). Indicator solution injected via central venous cannula and blood temperature changes are detected by a thermistor tip catheter placed in the artery. Thus, it combines pulse contour analysis with the transpulmonary thermodilution CO to determine hemodynamic variables. It requires manual calibration every 8 h and hourly during hemodynamic instability[27].

In addition, thermodilution curve can be used to measure intrathoracic blood volume (ITBV), global end diastolic volume (GEDV) and extravascular lung water (EVLW). GEDV and ITBV are a measure of cardiac preload and EVLW (interstitial, intracellular or intra alveolar) is a mean to quantify pulmonary edema. It also measures SVV/PVV which is marker of fluid responsiveness[28].

PiCCO is a relatively invasive method as it requires both arterial and venous cannulation. Its accuracy may be affected be vascular compliance, aortic impedence and peripheral arterial resistance. Moreover, air bubble, clots and inadequate indicator may also affect the accuracy. Valvular regurgitation, aortic aneurysm, significant arrhythmia and rapidly changing temperature may also affect its accuracy[29].

Various validation studies have found good correlation with PAC during CABG[30].However, that is not the case in patients undergoing OPCAB[31].Innon-cardiac and critically ill patients good correlation has been observed[32].Significant errors have been reported during hemodynamic instability requiring recalibration[33].

**FloTrac system:** FloTrac (Edwards LifeSciences. Irvine, United States) is a pulse contour device introduced in 2005 and is a minimally invasive method as it requires only an arterial line (femoral or radial).The systemdoes not need any external calibration, isoperator independent and easy to use. It is based on the principle that there is a linear relationship between the pulse pressure and SV[19,34].

The algorithm used in this system uses SD of 2000 arterial waveform points which is calculated by arterial pressure waveform sampled each 20 seconds at 100Hz. It incorporates characteristics of the arterial waveform with patient specific demographics. The SV is estimated by following equation-

SV= SDAPX µ

SDAP= Standard deviation of data points that reflects pulse pressure

µ= Conversion factor depends on arterial compliance, mean arterial pressure, waveform characteristics

Vascular compliance is patient’s biometric values (sex, age, height and sex)[35] and waveform characteristics assessed by skewness (degree of asymmetry) and kurtosis (degree of peakedness) of the individual arterial pressure waveform. A change in vascular tone is represented by skewness and kurtosis.The conversion factor µ enables calculation of SV without external calibration. Second generation devices also developed that calibrate every minute leading to improved CO measurement[36].A third generation device with Dynamo tone technology that has automatic adjustment for change in the vascular tone has also been made[37].Good arterial waveform quality is a prerequisite for accurate reading of CO. Accuracy is affected in patients with significant arrhythmias, IABP or morbid obesity[38].

Various studies have validated the efficacy of Flo Trac with PAC and find good correlation.We have studied Flo Trac with PAC in patients undergoing OPCAB and found good agreement. The mean bias and limits of agreement (2 standard deviations) expressed in liters per minute at respective points of measurement were -0.54 ± 1.12,-0.37 ± 1.0, -0.42 ± 1.50, -0.25 ± 1.18, -0.31 ± 1.28, 0.41 ± 1.0, 0.06 ± 1.50, and 0.09 ± 1.40[39].However, in patients with low SVR undergoing liver transplantation or septicemia it is not found as accurate as PAC[40-42].It is found to be useful in patients undergoing major abdominal surgery who receivedGDT[43].Moreover, the site of the arterial cannulation is also an important determinant of accuracy. In severe vasoconstriction radial artery reading will underestimate the CO while in volume responsive patient volume redistribution to cerebral circulation will also impair the pulse contour analysis through radial artery[3].

**Pressure recording analytic method:** PRAM- MostCare (Vytech, Padova, Italy) measures the area under the curve of arterial waveform. Major advantage is that it does not require external calibration and internal calibration is done by morphology of the arterial waveform. Pressure recording analytic method (PRAM) technology analyses whole cardiac cycle and area under the pressure wave (P/t) is determined[44].The P/t is divided into diastolic and systolic phase with 2 impedances based on different characteristics. However the accuracy of this method is still not proven.

**EV1000/Volume view:** A new calibrated pulse wave analysis method (VolumeView™/EV1000™, Edwards Lifesciences, Irvine, CA, United States) has been developed.It is based on pulse pressure analysis, which is calibrated by transpulmonarythermodilution and is currently under trial. Its comparison with PICCO2 system in critically ill patients found comparable results[45].However; very few studies are available for its validation. We have just finished a study on its use for GDT in OPCAB and found it to be very useful.

***Esophageal doppler***

Esophageal Doppler uses a flexible probe with transducer at the tip. It is of the size of anorogastric tube and can be placed for longer period in intubated patients. At the midthoracic level itmeasures flow as it is presumed to be parallel to the descending aorta. Since aorta is considered as a cylinder,the flow can be measured by multiplying cross-sectional area (CSA) and velocity. Doppler ultrasound is used to measure the SV. Once an optimal flow profile has been obtained, the blood flow velocity is determined from the shift in frequency of red blood cells. This is done by the ultrasound processor using the Doppler equation:

V =fd × c/2 × f0 × cos Ɵ

V = velocity of blood, fd = Doppler shift in frequency, c = speed of ultrasound in tissue (1540 m s21), f0 = initial ultrasound frequency, and Ɵ = the angle of ultrasound beam in relation to the blood flow.

The velocity–time integral (VTI) is calculated from the area under the velocity–time curve and used as the stroke distance.The area can be calculated bynomogram or direct measurement. Thus SV is calculated as CSA × VTI and CO is calculated as SV × HR[24].FTc *i.e.,* corrected time flow can also be determined which is used as measure of cardiac preload[46].

Major limiting factor is that it measures flow only in descending thoracic aorta which is 70% of total flow. A correction factor needs to be added to compensate aortic arch flow. Moreover discrepancies in flow may be seen in aortic coarctation, aneurysm or crossclamp, IABP and various metabolic states. Various factors like changes in pulse pressure, vascular compliance, volume status or inotropes may affect the CSA. In circulatory failure, it has been shown that CSA should be measured directly to prevent any inaccuracy in readings. Unchanged CSA may lead to underestimation of CO[24].Accurate velocity can only be determined by proper positioning of the probe which must be within 20° of the axial flow.

Various studies have compared ED with PAC and found good agreement with low bias. A meta analysis revealed it as a reliable method with low bias with limited efficacy[47]. ED has also been used in GDT and shown greater improvement in SV and CO with faster recovery and shorter LOS[48].In cardiac surgery, decreased hospital and ICU stay with decreased incidence of gut mucosal perfusion, without major complications has been shown with ED[49].We also studied this device in patients undergoing OPCAB and found that in comparison with PAC it cannot be used as a sole method for CO monitoring[50].

***Transesophgeal echocardiography***

Transesophgeal echocardiography (TEE) has now been a widely used monitor in perioperative setting. It is an important tool for the assessment of cardiac structures, filling status and cardiac contractility[51]. Moreover; aortic pathology can also be detected by TEE. Doppler technique is used to measure CO by Simpson’s rule measuring SV multiplied by HR. Flow is measured by area under the Doppler velocity waveform that gives VTI and CSA is calculated by planimetry. Measurement can be done at the level of pulmonary artery, mitral or aortic valve. TEE views used for measurement are midesophageal aortic long axis view and deep transgastric long axis view with pulsed and continuous wave Doppler respectively. The ultrasound beam is parallel to the blood flow in transgastric view.

TEE has been validated with PAC with good limits of agreement[52].It is a useful tool in hemodynamically unstable patient under mechanical ventilation[53].However, a skilled operator is required, limited availability and cost factor are major limitations for its use.Standard TEE probe cannot be kept in the patient for too long. Hemodynamic TEE (hTEE) is a disposable thinner TEE probe which can be left in situ for several days.

**NON INVASIVE METHODS**

***Partial gas rebreathing***

It is also known as the NICO system (Novametrix Medical Systems, Wallingford, Conn, United States) or partial gas re-breathing monitor and uses indirect Fick’s principle to calculate CO. It is used in intubated patients under mechanical ventilation. At steady state, the amount of CO2 entering the lungs via the pulmonary artery is proportional to the CO and equals the amount exiting the lungs via expiration and pulmonary veins.

During 30 s of re-breathing, the amount entering does not change, but the amount eliminated by expiration decreases and endtidal CO2increases in proportion to the CO[24].CO is calculated according to following formula-

CO = VCO2/CvCO2− CaCO2

Here VCO2 is CO2 consumption, CaCO2 and CvCO2 is arterial and venous CO2 content respectively. The diffusion rate of carbon dioxide is 22 times more rapid than that of oxygen, it is assumed that no difference in venous CO2 (CvCO2) will occur, whether under normal or rebreathing conditions. A disposable circuit is connected to the ventilator circuit along with infrared CO2 sensor, pneumotachometer and a rebreathing valve. Partial rebreathing is initiated every three minutes by opening the valve and pulmonary blood flow is calculated by difference between normal and rebreathing ratio[54].

Major limitation is that tracheal intubation with fixed ventilator setting is required. It is also not very accurate in patients with severe chest trauma, significant intrapulmonary shunt, high CO states and low minute ventilation[24].Validation studies have not found accuracy of this device with PAC. Studies have shown underestimation preoperatively and overestimation postoperatively after cardiac surgery[55].Thus it has limited clinical applicability in comparison to PAC.

***Thoracic bioimpedance***

Thoracic bioimpedance (TEB) is a non-invasive method of CO monitoring. Initially it was used by astronauts in 1960s[56]. it is based on the hypothesis by considering thorax as a cylinder perfused with fluid with specific resistivity. It measures the electrical resistance of the thorax to a high frequency, low amplitude current[24].

Electrodes six in number are placed (two on either side of neck and four in lower thorax) on the patient and the resistance to current flowing from the outermost to innermost electrodes is measured. The bioimpedance is indirectly proportional to the content of thoracic fluid. Tissue fluid volume, pulmonary and venous blood, and the aortic blood volume all contribute to the TEB measurement.Changes inCO will change the amount of aortic blood and will be reflected in a change TEB[5].SV is calculated using the formula[24]:

SV= VEPTxVETxEPCI,

VEPT = volume of electrically participating tissue (gender, height, and weight)

VET= ventricular ejection time taken from the R–R interval

EPCI=ejection phase contractility index which is indirectly proportional to TEB

Major limitations like interference with electrocautery, proper electrode placement, patient’s movements and arrhythmia may affect its accuracy. Studies in cardiac surgical patients revealed good correlation intraoperatively with a mean bias of -0.28 litre/min. Presence of sternal wires, orarrythmia may lead to inaccurate readings in the postoperative period[57].Results were also not encouraging in critically ill patients. Moreover, it has been considered as trend analysis monitor rather than a diagnostic one[58].

***Thoracic bioreactance***

Thoracic bioreactance (NICOM device, Cheetah medical, Portland, OR) is a modification of TEB which avoids interferences by noise and external sources. It analyses changes in the phase of electrical voltage signal to the current applied across the thorax. Changes in electrical capacitive and inductive properties occurs secondary to change in intrathoracic volume.

The method involves placement of two dual electrodes on either side of the thorax. Sine-wave high-frequency (75 kHz) current is transmitted into the body through one electrode and other electrode is used by the voltage input amplifier. The mean of two will give final value[59].

Electrocautery also affects its accuracy however if the device receives signal for atleast 20 s over a minute the CO value can be determined. Major advantage is the ease of use in intubated patients, arrhythmias, emergency room (ER), ICU and operating room (OR).Validating studies with PAC showed good correlation between the two methodswith minimal bias[57].Moreover comparison with pulse contour devices like PiCCO and ED also showed comparable results[58,60].

***Endotracheal CO monitor***

The endotracheal CO monitor (ECOM; Con-Med, Irvine, Calif, United States) measures CO using impedance plethysmography. It is based on the principle of bioimpedance and current is passed through electrodes attached to endotracheal tube shaft and cuff. Current is passed from electrode on the shaft of ETT and change in impedance secondary to aortic blood flow is detected by electrode on the cuff of ETT. An algorithm calculates SV based on impedance changes and CO can be calculated. Impedance is affected by aortic blood flow[61].

Electrocautery affects its accuracy and coronary blood flow is not calculated. Moreover the technology is still adequately not validated in humans, is costly and has not become very popular.

***Portable doppler device***

Ultrasonic Cardiac Output Monitors (USCOM, Sydney, Australia) is a portable device which is non-invasive and uses a probe placed suprasternally to measure flow through the aorta or on the left chest to measure transpulmonaryflow[62].It uses the Doppler principle as used with ED and TEE. Main advantage is the portability of the device and it can be used with ease in ER, OR, ICU and even in wards. Since it is a non-invasive device it can be used by trained nursing staff and is an important screening tool for postoperative cardiac surgical patients as well.

Major limitations are probe positioning as misalignment of ultrasound beam with blood flow may lead to errors and estimation of proper CSA in various physiological states is also important[24].

We have used USCOM device in post cardiac surgical patients for both left and right sided CO, CI and SV measurements and found good agreement with PAC.On comparing the right-sided CO, SV, and CI with those of PAC, the mean bias was 0.03 L/min, 1.6 mL, and 0.02 L/min/m2, respectively. The comparison of left-sided CO, SV, and CI with those of TD revealed a means bias of 0.14 L/min, 1.0mL, and 0.08 L/min/m2, respectively[63].We further studied this device in OPCAB and found good correlation with PAC.The CO had a mean bias of –0.13 L/min and limits of agreement (mean bias ± 2SD) at –0.86 and 0.59 L/min[64].

***Photoelectric plethysmography***

The Nexfin HD (BMEYE B.V, Amsterdam, Netherlands) is a completely non-invasive pulse pressure analysis device that assesses pulse pressure using photoelectric plethysmography in combination with a volume-clamp technique (inflatable finger cuff). CO is derived byModelflow method. There are very few validation studies to state its efficacy[65].

**CONCLUSION**

There are various newer devices for CO monitoring available in clinical practice that are validated against the gold standard method. Newer devices have the advantage of being minimally or non-invasive and portable. Hence, a few of them can be used outside the OR and ICU. Validation with PAC and other limitations may still be an obstacle for their usein different clinical scenarios. The criteria for selection of newer devices should be based on the institutional protocol and clinical condition of the patients. More RCT’s are needed to prove their efficacy and cost benefit. PAC will remain a gold standard for CO monitoring, however, use of newer devices based on pulse contour analysis, pulse pressure analysis and Doppler methods should be encouraged.

**REFERENCES**

1 **de Waal EE**, De Boeck BW, Kruitwagen CL, Cramer MJ, Buhre WF. Effects of on-pump and off-pump coronary artery bypass grafting on left ventricular relaxation and compliance: a comprehensive perioperative echocardiography study. *Eur J Echocardiogr* 2010; **11**: 732-737 [PMID: 20421229 DOI: 10.1097/ACO.0b013e32831f44d0]

2 **Dexter L**. Cardiac catheterization in the diagnosis of congenital heart disease. *Bull N Y Acad Med* 1950; **26**: 93-102 [PMID: 15409417]

3 **Swan HJ**, Ganz W, Forrester J, Marcus H, Diamond G, Chonette D. Catheterization of the heart in man with use of a flow-directed balloon-tipped catheter. *N Engl J Med* 1970; **283**: 447-451 [PMID: 5434111]

4 **Lee AJ**, Cohn JH, Ranasinghe JS. Cardiac output assessed by invasive and minimally invasive techniques. *Anesthesiol Res Pract* 2011; **2011**: 475151 [PMID: 21776254 DOI: 10.1155/2011/475151]

5 **Mehta Y**, Sharma KK. Double knot with formation of a double loop of pulmonary artery catheter. *J Cardiothorac Anesth* 1990; **4**: 149-150 [PMID: 2131846]

6 **Domino KB**, Bowdle TA, Posner KL, Spitellie PH, Lee LA, Cheney FW. Injuries and liability related to central vascular catheters: a closed claims analysis. *Anesthesiology* 2004; **100**: 1411-1418 [PMID: 15166560]

7 **Nishikawa T**, Dohi S. Errors in the measurement of cardiac output by thermodilution. *Can J Anaesth* 1993; **40**: 142-153 [PMID: 8443853]

8 **Elkayam U**, Berkley R, Azen S, Weber L, Geva B, Henry WL. Cardiac output by thermodilution technique. Effect of injectate's volume and temperature on accuracy and reproducibility in the critically Ill patient. *Chest* 1983; **84**: 418-422 [PMID: 6352195]

9 **Singh A**, Juneja R, Mehta Y, Trehan N. Comparison of continuous, stat, and intermittent cardiac output measurements in patients undergoing minimally invasive direct coronary artery bypass surgery. *J Cardiothorac Vasc Anesth* 2002; **16**: 186-190 [PMID: 11957168]

10 **Gore JM**, Goldberg RJ, Spodick DH, Alpert JS, Dalen JE. A community-wide assessment of the use of pulmonary artery catheters in patients with acute myocardial infarction. *Chest* 1987; **92**: 721-727 [PMID: 3652758]

11 **Connors AF**, Speroff T, Dawson NV, Thomas C, Harrell FE, Wagner D, Desbiens N, Goldman L, Wu AW, Califf RM, Fulkerson WJ, Vidaillet H, Broste S, Bellamy P, Lynn J, Knaus WA. The effectiveness of right heart catheterization in the initial care of critically ill patients. SUPPORT Investigators. *JAMA* 1996; **276**: 889-897 [PMID: 8782638]

12 **Robin ED**. Death by pulmonary artery flow-directed catheter. Time for a moratorium? *Chest* 1987; **92**: 727-731 [PMID: 3652759]

13 **Sandham JD**, Hull RD, Brant RF, Knox L, Pineo GF, Doig CJ, Laporta DP, Viner S, Passerini L, Devitt H, Kirby A, Jacka M. A randomized, controlled trial of the use of pulmonary-artery catheters in high-risk surgical patients. *N Engl J Med* 2003; **348**: 5-14 [PMID: 12510037]

14 **Harvey S**, Harrison DA, Singer M, Ashcroft J, Jones CM, Elbourne D, Brampton W, Williams D, Young D, Rowan K. Assessment of the clinical effectiveness of pulmonary artery catheters in management of patients in intensive care (PAC-Man): a randomised controlled trial. *Lancet* 2005; **366**: 472-477 [PMID: 16084255]

15 **Resano FG**, Kapetanakis EI, Hill PC, Haile E, Corso PJ. Clinical outcomes of low-risk patients undergoing beating-heart surgery with or without pulmonary artery catheterization. *J Cardiothorac Vasc Anesth* 2006; **20**: 300-306 [PMID: 16750726]

16 **Binanay C**, Califf RM, Hasselblad V, O'Connor CM, Shah MR, Sopko G, Stevenson LW, Francis GS, Leier CV, Miller LW. Evaluation study of congestive heart failure and pulmonary artery catheterization effectiveness: the ESCAPE trial. *JAMA* 2005; **294**: 1625-1633 [PMID: 16204662]

17 **Bland JM**, Altman DG. Statistical methods for assessing agreement between two methods of clinical measurement. *Lancet* 1986; **1**: 307-310 [PMID: 2868172]

18 **Linton RA**, Band DM, Haire KM. A new method of measuring cardiac output in man using lithium dilution. *Br J Anaesth* 1993; **71**: 262-266 [PMID: 8123404]

19 **Montenij LJ**, de Waal EE, Buhre WF. Arterial waveform analysis in anesthesia and critical care. *Curr Opin Anaesthesiol* 2011; **24**: 651-656 [PMID: 22036950 DOI: 10.1097/ACO.0b013e32834cd2d9]

20 **Garcia-Rodriguez C**, Pittman J, Cassell CH, Sum-Ping J, El-Moalem H, Young C, Mark JB. Lithium dilution cardiac output measurement: a clinical assessment of central venous and peripheral venous indicator injection. *Crit Care Med* 2002; **30**: 2199-2204 [PMID: 12394944]

21 **Costa MG**, Della Rocca G, Chiarandini P, Mattelig S, Pompei L, Barriga MS, Reynolds T, Cecconi M, Pietropaoli P. Continuous and intermittent cardiac output measurement in hyperdynamic conditions: pulmonary artery catheter vs. lithium dilution technique. *Intensive Care Med* 2008; **34**: 257-263 [PMID: 17922106]

22 **Pearse R**, Dawson D, Fawcett J, Rhodes A, Grounds RM, Bennett ED. Early goal-directed therapy after major surgery reduces complications and duration of hospital stay. A randomised, controlled trial [ISRCTN38797445]. *Crit Care* 2005; **9**: R687-R693 [PMID: 16356219]

23 **Hofer CK**, Cecconi M, Marx G, della Rocca G. Minimally invasive haemodynamic monitoring. *Eur J Anaesthesiol* 2009; **26**: 996-1002 [PMID: 19916204]

24 **Funk DJ**, Moretti EW, Gan TJ. Minimally invasive cardiac output monitoring in the perioperative setting. *Anesth Analg* 2009; **108**: 887-897 [PMID: 19224798 DOI: 10.1213/ane.0b013e31818ffd99]

25 **Wesseling KH**, Purschke R, Smith NT, Wüst HJ, de Wit B, Weber HA. A computer module for the continuous monitoring of cardiac output in the operating theatre and the ICU. *Acta Anaesthesiol Belg* 1976; **27 suppl**: 327-341 [PMID: 1015235]

26 **Sakka SG**, Kozieras J, Thuemer O, van Hout N. Measurement of cardiac output: a comparison between transpulmonary thermodilution and uncalibrated pulse contour analysis. *Br J Anaesth* 2007; **99**: 337-342 [PMID: 17611251]

27 **Oren-Grinberg A**. The PiCCO Monitor. *Int Anesthesiol Clin* 2010; **48**: 57-85 [PMID: 20065727 DOI: 10.1097/AIA.0b013e3181c3dc11]

28 **Marik PE**, Cavallazzi R, Vasu T, Hirani A. Dynamic changes in arterial waveform derived variables and fluid responsiveness in mechanically ventilated patients: a systematic review of the literature. *Crit Care Med* 2009; **37**: 2642-2647 [PMID: 19602972 DOI: 10.1097/CCM.0b013e3181a590da.PMID: ]

29 PULSION Medical Inc, “Training documents—advanced hemodynamic monitoring,” August 2009, http: //www3.pulsion.de/fileadmin/pulsion share/Education/Training/TraintheTrainer/TtT MPI851405US R00 101008 Parameters.pdf.

30 **Buhre W**, Weyland A, Kazmaier S, Hanekop GG, Baryalei MM, Sydow M, Sonntag H. Comparison of cardiac output assessed by pulse-contour analysis and thermodilution in patients undergoing minimally invasive direct coronary artery bypass grafting. *J Cardiothorac Vasc Anesth* 1999; **13**: 437-440 [PMID: 10468257]

31 **Halvorsen PS**, Sokolov A, Cvancarova M, Hol PK, Lundblad R, Tønnessen TI. Continuous cardiac output during off-pump coronary artery bypass surgery: pulse-contour analyses vs pulmonary artery thermodilution. *Br J Anaesth* 2007; **99**: 484-492 [PMID: 17650518]

32 **Della Rocca G**, Costa MG, Coccia C, Pompei L, Di Marco P, Vilardi V, Pietropaoli P. Cardiac output monitoring: aortic transpulmonary thermodilution and pulse contour analysis agree with standard thermodilution methods in patients undergoing lung transplantation. *Can J Anaesth* 2003; **50**: 707-711 [PMID: 12944446]

33 **Boyle M**, Murgo M, Lawrence J, Belessis A, Shehabi Y. Assessment of the accuracy of continuous cardiac output and pulse contour cardiac output in tracking cardiac index changes induced by volume load. *Aust Crit Care* 2007; **20**: 106-112 [PMID: 17629491]

34 **Hamm JB**, Nguyen BV, Kiss G, Wargnier JP, Jauffroy A, Helaine L, Arvieux CC, Gueret G. Assessment of a cardiac output device using arterial pulse waveform analysis, Vigileo, in cardiac surgery compared to pulmonary arterial thermodilution. *Anaesth Intensive Care* 2010; **38**: 295-301 [PMID: 20369763]

35 **Langewouters GJ**, Wesseling KH, Goedhard WJ. The static elastic properties of 45 human thoracic and 20 abdominal aortas in vitro and the parameters of a new model. *J Biomech* 1984; **17**: 425-435 [PMID: 6480618]

36 **Mayer J**, Boldt J, Poland R, Peterson A, Manecke GR. Continuous arterial pressure waveform-based cardiac output using the FloTrac/Vigileo: a review and meta-analysis. *J Cardiothorac Vasc Anesth* 2009; **23**: 401-406 [PMID: 19464625 DOI: 10.1053/j.jvca.2009.03.003]

37 Edwards Lifesciences LLC, “FloTrac system 3rd generation software: The next generation in hemodynamic management,”December 2010, http: //www.edwards.com/sitecollectionimages/products/mininvasive/ar04099.pdf.

38 **De Backer D**, Marx G, Tan A, Junker C, Van Nuffelen M, Hüter L, Ching W, Michard F, Vincent JL. Arterial pressure-based cardiac output monitoring: a multicenter validation of the third-generation software in septic patients. *Intensive Care Med* 2011; **37**: 233-240 [PMID: 21153399 DOI: 10.1007/s00134-010-2098-8]

39 **Mehta Y**, Chand RK, Sawhney R, Bhise M, Singh A, Trehan N. Cardiac output monitoring: comparison of a new arterial pressure waveform analysis to the bolus thermodilution technique in patients undergoing off-pump coronary artery bypass surgery. *J Cardiothorac Vasc Anesth* 2008; **22**: 394-399 [PMID: 18503927 DOI: 10.1053/j.jvca.2008.02.015]

40 **Monnet X**, Anguel N, Naudin B, Jabot J, Richard C, Teboul JL. Arterial pressure-based cardiac output in septic patients: different accuracy of pulse contour and uncalibrated pressure waveform devices. *Crit Care* 2010; **14**: R109 [PMID: 20537159 DOI: 10.1186/cc9058]

41 **Metzelder S**, Coburn M, Fries M, Reinges M, Reich S, Rossaint R, Marx G, Rex S. Performance of cardiac output measurement derived from arterial pressure waveform analysis in patients requiring high-dose vasopressor therapy. *Br J Anaesth* 2011; **106**: 776-784 [PMID: 21441548 DOI: 10.1093/bja/aer066]

42 **Monnet X**, Anguel N, Jozwiak M, Richard C, Teboul JL. Third-generation FloTrac/Vigileo does not reliably track changes in cardiac output induced by norepinephrine in critically ill patients. *Br J Anaesth* 2012; **108**: 615-622 [PMID: 22265900 DOI: 10.1093/bja/aer491]

43 **Mayer J**, Boldt J, Mengistu AM, Röhm KD, Suttner S. Goal-directed intraoperative therapy based on autocalibrated arterial pressure waveform analysis reduces hospital stay in high-risk surgical patients: a randomized, controlled trial. *Crit Care* 2010; **14**: R18 [PMID: 20156348 DOI: 10.1186/cc8875]

44 **Giomarelli P**, Biagioli B, Scolletta S. Cardiac output monitoring by pressure recording analytical method in cardiac surgery. *Eur J Cardiothorac Surg* 2004; **26**: 515-520 [PMID: 15302045]

45 **Bendjelid K**, Marx G, Kiefer N, Simon TP, Geisen M, Hoeft A, Siegenthaler N, Hofer CK. Performance of a new pulse contour method for continuous cardiac output monitoring: validation in critically ill patients. *Br J Anaesth* 2013; **111**: 573-579 [PMID: 23625132 DOI: 10.1093/bja/aet116]

46 **DiCorte CJ**, Latham P, Greilich PE, Cooley MV, Grayburn PA, Jessen ME. Esophageal Doppler monitor determinations of cardiac output and preload during cardiac operations. *Ann Thorac Surg* 2000; **69**: 1782-1786 [PMID: 10892923]

47 **Laupland KB**, Bands CJ. Utility of esophageal Doppler as a minimally invasive hemodynamic monitor: a review. *Can J Anaesth* 2002; **49**: 393-401 [PMID: 11927480]

48 **Sinclair S**, James S, Singer M. Intraoperative intravascular volume optimisation and length of hospital stay after repair of proximal femoral fracture: randomised controlled trial. *BMJ* 1997; **315**: 909-912 [PMID: 9361539]

49 **Mythen MG**, Webb AR. Perioperative plasma volume expansion reduces the incidence of gut mucosal hypoperfusion during cardiac surgery. *Arch Surg* 1995; **130**: 423-429 [PMID: 7535996]

50 **Sharma J**, Bhise M, Singh A, Mehta Y, Trehan N. Hemodynamic measurements after cardiac surgery: transesophageal Doppler versus pulmonary artery catheter. *J Cardiothorac Vasc Anesth* 2005; **19**: 746-750 [PMID: 16326299]

51 **Poelaert J**, Schmidt C, Colardyn F. Transoesophageal echocardiography in the critically ill. *Anaesthesia* 1998; **53**: 55-68 [PMID: 9505744]

52 **Perrino AC**, Harris SN, Luther MA. Intraoperative determination of cardiac output using multiplane transesophageal echocardiography: a comparison to thermodilution. *Anesthesiology* 1998; **89**: 350-357 [PMID: 9710392]

53 **Poelaert J**, Schmidt C, Van Aken H, Hinder F, Mollhoff T, Loick HM. A comparison of transoesophageal echocardiographic Doppler across the aortic valve and the thermodilution technique for estimating cardiac output. *Anaesthesia* 1999; **54**: 128-136 [PMID: 10215707]

54 **Alhashemi JA**, Cecconi M, della Rocca G, Cannesson M, Hofer CK. Minimally invasive monitoring of cardiac output in the cardiac surgery intensive care unit. *Curr Heart Fail Rep* 2010; **7**: 116-124 [PMID: 20623210 DOI: 10.1007/s11897-010-0019-3]

55 **Gueret G**, Kiss G, Rossignol B, Bezon E, Wargnier JP, Miossec A, Corre O, Arvieux CC. Cardiac output measurements in off-pump coronary surgery: comparison between NICO and the Swan-Ganz catheter. *Eur J Anaesthesiol* 2006; **23**: 848-854 [PMID: 16953944]

56 **Kubicek WG**, Karnegis JN, Patterson RP, Witsoe DA, Mattson RH. Development and evaluation of an impedance cardiac output system. *Aerosp Med* 1966; **37**: 1208-1212 [PMID: 5339656]

57 **Spiess BD**, Patel MA, Soltow LO, Wright IH. Comparison of bioimpedance versus thermodilution cardiac output during cardiac surgery: evaluation of a second-generation bioimpedance device. *J Cardiothorac Vasc Anesth* 2001; **15**: 567-573 [PMID: 11687996]

58 **Squara P**, Rotcajg D, Denjean D, Estagnasie P, Brusset A. Comparison of monitoring performance of Bioreactance vs. pulse contour during lung recruitment maneuvers. *Crit Care* 2009; **13**: R125 [PMID: 19638227 DOI: 10.1186/cc7981]

59 **Keren H**, Burkhoff D, Squara P. Evaluation of a noninvasive continuous cardiac output monitoring system based on thoracic bioreactance. *Am J Physiol Heart Circ Physiol* 2007; **293**: H583-H589 [PMID: 17384132]

60 **Marqué S**, Cariou A, Chiche JD, Squara P. Comparison between Flotrac-Vigileo and Bioreactance, a totally noninvasive method for cardiac output monitoring. *Crit Care* 2009; **13**: R73 [PMID: 19454009 DOI: 10.1186/cc7884]

61 CONMED Corporation, “ECOM endotracheal cardiac output monitor,” 2010, http: //www.conmed.com/products ECOM.php

62 **Meyer S**, Todd D, Wright I, Gortner L, Reynolds G. Review article: Non-invasive assessment of cardiac output with portable continuous-wave Doppler ultrasound. *Emerg Med Australas* 2008; **20**: 201-208 [PMID: 18400002 DOI: 10.1111/j.1742-6723.2008.01078.x]

63 **Chand R**, Mehta Y, Trehan N. Cardiac output estimation with a new Doppler device after off-pump coronary artery bypass surgery. *J Cardiothorac Vasc Anesth* 2006; **20**: 315-319 [PMID: 16750729]

64 **Arora D**, Chand R, Mehta Y, Trehan N. Cardiac output estimation after off-pump coronary artery bypass: a comparison of two different techniques. *Ann Card Anaesth* 2007; **10**: 132-136 [PMID: 17644886]

65 **Alhashemi JA**, Cecconi M, Hofer CK. Cardiac output monitoring: an integrative perspective. *Crit Care* 2011; **15**: 214 [PMID: 21457508 DOI: 10.1186/cc9996]

**P-Reviewers:** Hung MJ, Huang JP, Sicari R **S-Editor:** Wen LL **L-Editor: E-Editor:**

**Table 1 Advantages and disadvantages of methods of cardiac output monitoring**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **No** | **Device** | **Type** | **Advantages** | **Disadvantages** |
| 1 | PAC | Invasive | Gold standard | Catheter related complications |
| 2 | Continous CO by PAC | Invasive | Continous CO measurement | Catheter related complications  Cost |
| 3 | LiDCO | Minimally invasive | Only one arterial line  Continuous CO measurements  Measure SV and SVV | Requires good arterial waveform  Requires Calibration  Contraindicated in Lithium therapy |
| 4 | PiCCO | Minimally invasive | Continuous CO measurement  Effective during hemodynamic instability | Requires good arterial waveform  Requires calibration |
| 5 | FloTrac | Minimally invasive | Continuous CO measurement  No calibration | Requires good arterial waveform |
| 6 | PRAM | Minimally invasive | No Calibration | Still not validated |
| 7 | Esophgeal Doppler | Minimally invasive | Simple to use  Reliable  Useful in GDT | Measure flow only in descending thoracic aorta  Assumptions about aortic size may not be accurate |
| 8 | TEE | Minimally invasive | Evaluate cardiac anatomy preload and ventricular function | Cost  Skilled personnel |
| 9 | Partial non-rebreathing  systems | Non invasive | Ease of use  Continuous CO measurement | Affected by changes in dead space or V/Q matching |
| 10 | Thoracic  bioimpedance | Non invasive | Continuous CO measurement | Affected by electrical noise, movement, temperature and humidity  Requires hemodynamic stability  Not useful in dysrhythmias |
| 11 | ECOM | Non invasive | Continuous CO measurement | Coronary blood flow not recorded  Electrocautery produces interference |

CO: Cardiac output; LiDCO: Lithium dilution; PiCCO and FloTrac: Pulse contour analysis; PRAM: Pressure recording analytic method; ED: Esophgeal Doppler; TEE: Transesophgeal echocardiography; ECOM: Endotracheal cardiac output monitor; PAC: Pulmonary artery catheter.