| Impact of dose calculation algorithm on radiation therapy Wen-Zhou Chen, Ying Xiao, Jun Li |
| --- |
| CITATION | Chen WZ, Xiao Y, Li J. Impact of dose calculation algorithm on radiation therapy. World J Radiol 2014; 6(11): 874-880 |
| URL | http://www.wjgnet.com/1949-8470/full/v6/i11/874.htm |
| DOI | http://dx.doi.org/10.4329/wjr.v6.i11.874 |
| OPEN ACCESS | Articles published by this Open-Access journal are distributed under the terms of the Creative Commons Attribution Non-commercial License, which permits use, distribution, and reproduction in any medium, provided the original work is properly cited, the use is non commercial and is otherwise in compliance with the license. |
| CORE TIP | This paper is a review of the impact of current commercial dose calculation algorithms on radiation therapy, with a focus on discussing the impact on tumor control probability and normal tissue complication probability. |
| KEY WORDS | Dose calculation; Algorithm; Radiation therapy; Tumor control probability; Normal tissue complication probability |
| COPYRIGHT | © 2014 Baishideng Publishing Group Inc. All rights reserved. |
| COPYRIGHT LICENSE | Order reprints or request permissions: bpgoffice@wjgnet.com |
| NAME OF JOURNAL | World Journal of Radiology |
| ISSN | 1949-8470 (online) |
| PUBLISHER | Published by Baishideng Publishing Group Inc, 8226 Regency Drive, Pleasanton, CA 94588, USA |
| WEBSITE | http://www.wjgnet.com |

ESPS Manuscript NO: 11212

Columns: MINIREVIEWS

Impact of dose calculation algorithm on radiation therapy

Wen-Zhou Chen, Ying Xiao, Jun Li

Wen-Zhou Chen, Ying Xiao, Jun Li, Department of Radiation Oncology, Jefferson Medical College, Philadelphia, PA 19107, United States

Author contributions: All three authors contributed to manuscript preparation

Supported by In part, under a grant with the Pennsylvania Department of Health

Correspondence to: Jun Li, PhD, Department of Radiation Oncology, Jefferson Medical College, 111 South 11th St., Philadelphia, PA 19107, United States. jun.li@jefferson.edu

Telephone: +1-215-9557945  Fax: +1-215-5032386

Received: May 8, 2014           Revised: August 4, 2014

Accepted: September 23, 2014

Published online: November 28, 2014

Abstract

The quality of radiation therapy depends on the ability to maximize the tumor control probability while minimizing the normal tissue complication probability. Both of these two quantities are directly related to the accuracy of dose distributions calculated by treatment planning systems. The commonly used dose calculation algorithms in the treatment planning systems are reviewed in this work. The accuracy comparisons among these algorithms are illustrated by summarizing the highly cited research papers on this topic. Further, the correlation between the algorithms and tumor control probability/normal tissue complication probability values are manifested by several recent studies from different groups. All the cases demonstrate that dose calculation algorithms play a vital role in radiation therapy.

© 2014 Baishideng Publishing Group Inc. All rights reserved.

Key words: Dose calculation; Algorithm; Radiation therapy; Tumor control probability; Normal tissue complication probability

Core tip: This paper is a review of the impact of current commercial dose calculation algorithms on radiation therapy, with a focus on discussing the impact on tumor control probability and normal tissue complication probability.

Chen WZ, Xiao Y, Li J. Impact of dose calculation algorithm on radiation therapy. World J Radiol 2014; 6(11): 874-880  Available from: URL: http://www.wjgnet.com/1949-8470/full/v6/i11/874.htm  DOI: http://dx.doi.org/10.4329/wjr.v6.i11.874

INTRODUCTION

The quality of radiation therapy depends on the ability to maximize the tumor control probability (TCP) while minimizing the normal tissue complication probability (NTCP) at the same time. Since these two quantities are directly dependent on the absorbed dose in the targets and in the organs at risk (OARs) respectively, accurate knowledge of dose distribution within the patient are crucial in radiation therapy. International Commission on Radiation Units and Measurements (ICRU)[1] has recommended an overall dose accuracy within 5%. Considering the uncertainties resulting from patient setup, machine calibration and dose calculation from treatment planning systems, it is necessary to have a dose calculation algorithm that can predict dose distribution within 3% accuracy.

Accurate calculation of dose distribution in an inhomogeneous medium such as human body is a complicated task, especially for tumors located in the lung. To date, only the Monte Carlo method is considered to be the most accurate algorithm for dose calculation but it requires the greatest processing time. Apart from Monte Carlo method, all other methods make different degrees of approximation and simplification which lead to much faster calculation speed but also result in less accurate dose distribution comparing with the Monte Carlo simulation.

The purpose of this work is to review the effect of dose calculation algorithms on the radiation therapy for different disease sites and special focus is given for the lung region. As mentioned in the American Association of Physicists in Medicine (AAPM) Report No. 85[2], the level of dose differences can be detected clinically. In order to quantify the clinical effects, we review the works on the correlation of dose calculation algorithms with computed values of tumor control probability and normal tissue complication probability. The impact of the accuracy of the algorithms is directly related to the quality of radiation therapy.

DOSE CALCULATION ALGORITHMS

The Monte Carlo dose calculation method is considered to be the most accurate algorithm and has always been used as the generation of benchmark dose distribution with which to compare the results of other less-computer-intensive dose calculation methods[3]. The Monte Carlo method uses photon and electron transport physics to consider the trajectories of individual particles and thus the pattern of dose deposition. Each particle’s history is determined by the random number generator and millions of particles’ histories are traced. The dose distribution is built by summing the energy deposition in each particle’s history.

Apart from the Monte Carlo simulation, all other commonly used dose calculation algorithms can be categorized into two groups[2,4,5]: (1) Methods based on equivalent path length (EPL)[6] scaling or equivalent tissue-air ratio (ETAR)[7] for inhomogeneity corrections. In these methods the changes in lateral transport of electrons are not modeled; and (2) Methods based on convolution techniques, in which the inhomogeneities are handled either by an equivalent path length correction or scaled kernels and the lateral electron transport is considered in an approximate way. In this work, these two types of algorithms are referred to as type (1) and type (2) methods. In type (1) methods, the equivalent path length correction is a one-dimensional method that takes into account of electron density information along a ray path from the source to the point in question. There are two methods: ratio of tissue-air ratios (RTAR) method[2] and power law method which is also referred as modified Batho method[8]. These methods correctly account for the change in the attenuation of the primary dose but not in the scatter contribution, thus result in an overestimation of dose when the electron density is less than unity and an underestimation when the electron density is greater than unity. The equivalent tissue-air ratio method is a three-dimensional correction method which is based on full three-dimensional density information acquired from CT images. This method applies a ray trace to determine the change in the primary dose and calculate the scatter dose based on the three-dimensional density data. Although methods in type (1) do not perform an accurate dose distribution calculation in patients, they are still used by some treatment planning systems for a quick dose calculation to give the planner a rough idea about the absorbed dose and by some dose verification systems to perform a second independent check to catch the gross errors.

In type (2) methods, the model-based convolution/superposition algorithms[9-13] are widely used in commercial radiotherapy treatment planning systems (TPSs), which perform dose calculations with accuracies close to the results of Monte Carlo simulation while taking much less time. All convolution algorithms have two essential components: one representing the energy imparted to the medium by the interactions of primary photons, called Terma (total energy released per unit mass) and one representing the energy deposited about a primary photon interaction site, the kernel. The kernel can be further separated into two parts: the primary kernel which calculates the primary dose and the scatter kernel which calculates the first and multiple scatter doses. The dose at any point can be calculated from the convolution of the Terma with the kernel. In order to account for tissue heterogeneities in a patient, kernel is scaled by radiological distances which are calculated from the material densities defined by CT images. Rigorously speaking, when the scaled kernel is used, the process is not a convolution any more since the kernel is not invariant in space and it is in fact a superposition of varying kernels with the Terma. The treatment planning systems that use the superposition algorithm include, for example, XiO (Elekta, Inc.). Several variations of the convolution/superposition algorithms exist today and two typical and mostly used ones are collapsed cone convolution (CCC) and pencil beam convolution (PBC) techniques[4]. The collapsed cone convolution method uses a polyenergetic Terma and kernel, where the kernel is represented analytically and expressed in polar coordinates. There are a finite number of polar angles with respect to the primary beam. The interaction site can be considered to be at the apex of a set of radially directed lines spreading out in three dimensions. Each line is considered to be the axis of a cone. The kernel along each line is actually the energy deposited within the entire cone collapsed onto the line. The advantage of the CCC method over standard convolution is that the computation time increase with MN3 as opposed to N6, where M is the number of cones and N is the number of voxels along one side of the calculation volume. The treatment planning systems that use the CCC method include, for example, Pinnacle (Philips, Inc.) and Oncentra MasterPlan (Nucletron, Inc.). In the pencil beam convolution method, the dose deposited at a point is calculated as a convolution of Terma with a pencil-shape-like kernel which is derived from the measured beam data. The pencil-beam kernel describes the dose distribution of a very narrow beam entering a water phantom along the beam’s central axis. Inhomogeneity correction is performed with an equivalent path length correction for the primary dose contribution and a one-dimensional convolution along fan lines for scattered radiation[14,15]. The anisotropic analytical algorithm (AAA)[16,17] used by Eclipse TPS (Varian Medical Systems) is based on the pencil beam convolution technique. The AAA uses spatially variant convolution scatter kernels which are derived from Monte Carlo simulation, and separate modeling for primary photons, scattered photons, and contaminant electrons. Inhomogeneity is handled with radiological scaling of the dose deposition functions in the beamlet direction and electron-density-based scaling of the photon scatter kernels in 16 lateral directions. The final doses are obtained by superposing the doses from the photon and electron convolutions[18,19]. The anisotropic analytical algorithm is an attractive option for routine clinical use because of its relatively short computation time and accuracy comparing with the Monte Carlo method.

COMPARISON OF DOSE CALCULATION ALGORITHMS AND THEIR CLINICAL IMPACT

Comparisons of dose calculation algorithms for clinical treatment disease sites have been studied in many references[4,19-21]. In this review, we first summarize the comparisons of dose calculation algorithms for four commonly treated disease sites, which demonstrate that dose calculation algorithms that can calculate dose accurately in inhomogeneous environment are essential for lung tumor treatment. Then we focus on the dose calculation algorithm for lung tumor treatment planning. Different treatment techniques are discussed. Finally we show the correlation of the algorithms with TCP/NTCP.

In Knöös et al[4]’s paper, the authors studied the performance of different dose calculation algorithms from five commercial radiotherapy treatment planning systems for four common treatment disease sites: prostate, head and neck, breast and lung. The Monte Carlo algorithm was used as a benchmark for comparison between different algorithms. Increasing the complexity from the relatively homogeneous pelvic region to the very inhomogeneous lung region resulted in less accurate dose distributions. Improvements in the accuracy of dose calculation were observed when the methods taking into account of volume scatter and changes in electron transport were used, that is, when type (2) algorithms were used. That was especially important when the extension of the irradiated volume was limited such as in the breast case and when low densities were presented such as in the lung case. In the prostate case, no significant differences were found in the results calculated with different algorithms. For instance, when 6 MV was used, the dose to 95% of the PTV was in a range of 96.2% to 100.3% for all studied systems, with an average value of 98.2%. Qualitatively, all the plans which were calculated with different methods, were very similar. The similar situation existed in the head and neck case. The average dose per monitor unit (MU) to the PTV was decreased by 1% for the low energy if more accurate methods, i.e., type (2) methods, were used. This difference was not presented for the higher energy, due to less scatter in the high energy beam. The dose to 95% of the PTV showed no significant change when moving from type (1) methods to type (2) methods for both low and high energies. The ETAR method of type (1) resulted in doses closer to that calculated with type (2) methods, due to the improved scatter integration which took into account the 3D extension of the volume more accurately. In the breast case, two equally weighted opposed tangential beams were used. The average PTV doses were decreased by 0.7% and 1.6% for low and high energies, respectively, when comparing type (1) with type (2) methods. In general, larger differences in dose calculation were found in high energy treatment due to the longer range of electrons, especially in the low density lung tissues. In the pulmonary case, for 6 MV, the average dose per MU to the PTV was decreased by 2.5% when the type (2) methods were used, compared with that calculated with type (1) methods. Changing the energy to high energies increased the difference to 3.7%. The high dose volume within the PTV was decreased by 3.4% and 4.6%, moving from type (1) methods to type (2) methods for low and high energies, respectively. This implies that accurate tumor doses are different from the doses predicted with those methods, and accurate tumor doses needs to be predicted with advanced dose calculation algorithms, i.e., Monte Carlo algorithm. Thus the algorithm directly affects the local control of tumors in lung cancer. That is, less coverage for tumor is presented when more realistic and accurate methods is used. This paper and many other references[19-22] showed that the dose calculation algorithms have a significant impact on radiation therapy for lung cancer treatment.

Remarkable impact of dose calculation algorithms on radiation therapy has been observed in the treatment of lung cancer, when tissue density correction was taken into account. Differences between dose calculations with and without density corrections in the thoracic region have been reported[23-28]. In Xiao et al[27]’s paper, a retrospective dosimetric study was carried out based on the treatment plans submitted to Radiation Therapy Oncology Group (RTOG) 0236 clinical trials of non-small-cell lung cancer (NSCLC) treatment with stereotactic body radiotherapy (SBRT). The protocol required each institution to submit two plans: one plan without heterogeneity correction and one plan with heterogeneity correction, with identical MUs. In Xiao et al[27]’s study, the authors found that the planning target volume receiving greater than 60 Gy was decreased, on average, by 10.1% when heterogeneity corrections were applied. The maximal dose to any point greater than 2 cm away from the planning target volume increased from 35.2 Gy to 38.5 Gy.

The impact of heterogeneity corrections of dose algorithms on target coverage in the SBRT lung treatment was studied in more details in Ding et al[22]’s paper. The dose calculations using four different algorithms were compared with experimental measurements. The pencil beam algorithm with no heterogeneity corrections (PB-NC) and with modified Batho heterogeneity corrections (PB-MB), the anisotropic analytical algorithm (AAA) and Monte Carlo simulation were investigated in ten patients’ treatment planning. The plans included 8-10 non-opposed photon beams and 2-4 of the beams were non-coplanar. The field sizes ranged from 3.5 cm × 3.5 cm to 6 cm × 6 cm with the mean value close to 4 cm × 4 cm. The mixed 6 and 10 MV energies were used. The authors found that the differences in calculated doses to 95% or 99% of the PTV, between calculations using the PB-NC and the AAA, were within 10% of prescribed dose. Compared to that calculated with the AAA, the minimum doses to 95% of PTV calculated using the PB-MB were overestimated by up to 40% of the prescribed dose. The calculated maximum doses were underestimated by up to 27% using the PB-NC and overestimated by 19% using the PB-MB. The dose distributions near the interface calculated with the AAA agreed with those from Monte Carlo calculations and the measurements.

The above publications demonstrated the impact of dose calculation algorithms on the lung cancer treatment. These comparisons were mainly between type (1) and type (2) methods. The direct comparisons between type (2) algorithms and Monte Carlo simulation have also been done extensively. For instance, in Vanderstraeten et al[20]’s study[21], the authors compared the accuracy between Monte Carlo, convolution/superposition, and pencil beam dose calculations for intensity modulated radiation therapy (IMRT) of lung cancer, and they found that the convolution/superposition methods showed an excellent agreement with Monte Carlo method for dose calculation within the target structures, whereas the best agreement in OAR doses was found between collapsed cone convolution model and Monte Carlo simulation. Results from pencil beam algorithm were unsatisfying for both target and OARs. In Li et al[28]’s paper, the authors compared superposition algorithm with Monte Carlo method for SBRT non-small-cell lung cancer treatment and they found that the important dosimetric parameter R50 (ratio of 50% prescription isodose volume to PTV) recommended by RTOG 0813 protocol had 12% difference on average between superposition and Monte Carlo calculations.

All these research studies have demonstrated that for dose calculation in lung region the advanced type (2) methods are necessary, and the collapsed cone convolution algorithm and anisotropic analytical algorithm are appropriate options for their relative accurate calculation results compared with the Monte Carlo method.

In the above, we have discussed that the different dose calculation algorithms could give different levels of dose distribution accuracy. Further we will discuss that this different levels of accuracy could be detected clinically, which affect the quality of radiotherapy. The American Association of Physicists in Medicine (AAPM) Report No. 85[2] on tissue inhomogeneity corrections mentioned that a 5% change in dose may result in a significant change in tumor control probability (TCP) and normal tissue complication probabilities (NTCP). In this report, the authors mentioned two examples[29]: A 7% difference in dose delivered to different groups of patients was discovered by a radiation oncologist; and two experiences from the Institut Gustave Roussy, which were related to tumor regression and normal tissue reactions, respectively.

Although it is still a relative new topic, the correlation between dose algorithms and local control, TCP and NTCP, has already been investigated by several groups and more research is expected to be done in the future. In Chetty et al[30]’s study, 133 NSCLC patients with stereotactic ablative radiotherapy (SABR)-based treatment were chosen for the correlation study. The correction-based pencil-beam algorithm, model-based convolution/superposition algorithm, and Monte Carlo algorithm were applied for dose calculation. TCP was computed using the Marsden[31,32] model and associations between dose and outcome were inferred. The authors found that model-based mean TCP’s were approximately 8%-9%, 6%-8%, and 3%-5% lower than those of correction-based algorithms for volumes < 60, 60-100, and > 100 cm3, respectively, when the same treatment arrangement were applied. This was because that the advanced type (2) methods simulated the dose deposition physics in a more realistic way than that type (1) methods. Further, the maximum decrement in Monte Carlo-based TCP was about 50% for volumes < 30 cm3. Variation in TCP ranges among model-based algorithms is due to the differences in the PTV minimum doses observed in the dose-volume histograms which were the direct products of the calculation algorithms. Though these differences did not have a significant effect on the PTV D95, they had a strong impact on the TCP. The results implied that more advanced algorithms are essential to assess the quality of the treatment clinically in the more realistic way. The detailed results of the percent mean tumor control probability (TCP) values for all algorithms as a function of PTV volume are cited and listed in Table 1.

In Bufacchi et al[33]’s study, the focus was shifted to the clinical implication of algorithms on NTCP models for four tumor sites: prostate, head and neck, breast and lung. The pencil beam convolution and anisotropic analytical algorithm were used for 80 treatment plans. The authors found that when the original PBC treatment plans were recalculated using AAA with the same number of monitor units, the NTCP became lower, except for the breast treatments. Further the authors concluded that this difference in NTCP between PBC and AAA treatment plans could be clinically significant. In Nielsen et al[34]’s paper, the study was specifically focused on the influence of dose calculation algorithms on NTCP in NSCLC patients. Six dose algorithms from four different treatment planning systems were investigated: Eclipse AAA, Oncentra MasterPlan Collapsed Cone and Pencil Beam, Pinnacle Collapsed Cone, and XiO Multigrid Superposition and Fast Fourier Transform Convolution. NTCP values for heart and lungs were calculated using the relative seriality model[35] and the LKB model[36,37], respectively. The authors found that the influence of density correction on the NTCP values depended on the dose calculation algorithms and the NTCP model parameter set. Compared to mean lung dose (MLD) and V20, the calculated NTCP values for pneumonitis were more sensitive to the calculation algorithms. All these implied that for plan evaluation the algorithms play an extremely important role and the dosimetric parameters such as MLD and V20 might not be sensitive enough for the assessment. The differences of the quantities calculated with and without density correction using each algorithm are cited and listed in Table 2.

To summarize, we list the clinical impact of dose calculation algorithms in Table 3. Five references[30,33,34,38,39] with their results and conclusions are summarized.

CONCLUSION

In this study we reviewed the commonly used dose calculation algorithms: correction-based type (1) methods and model-based type (2) methods. The calculation accuracy of different algorithms illustrated by several studies was summarized. Special focus was given to dose calculation comparison in the lung region. All the research studies demonstrated that for dose calculation in lung region, the advanced type (2) methods are necessary. Further, the accuracy of dose calculation algorithms was correlated to the quantities of TCP/NTCP, and the connection between the algorithms and clinical impact was established. The clinically related TCP/NTCP values are sensitive to the accuracy of dose algorithms. In conclusion, dose calculation algorithms play a vital role in radiation therapy.

REFERENCES

1	International Commission on Radiation Units and Measurements (ICRU). Determination of absorbed dose in a patient irradiated by beams of X or gamma rays in radiotherapy procedures. ICRU Report 24. Washington (DC): ICRU, 1976: 67

2	AAPM. The American Association of Physicists in Medicine (AAPM) Report 85 Tissue inhomogeneity corrections for MV photon beams. USA: Report of Task Group No. 65 of AAPM, 2004

3	Metcalfe P, Kron T, Hoban P. The physics of radiotherapy x-rays and electrons. Madison, WI: Medical Physics Publishing, 2007

4	Knöös T, Wieslander E, Cozzi L, Brink C, Fogliata A, Albers D, Nyström H, Lassen S. Comparison of dose calculation algorithms for treatment planning in external photon beam therapy for clinical situations. Phys Med Biol 2006; 51: 5785-5807 [PMID: 17068365 DOI: 10.1088/0031-9155/51/22/005]

5	Fogliata A, Vanetti E, Albers D, Brink C, Clivio A, Knöös T, Nicolini G, Cozzi L. On the dosimetric behaviour of photon dose calculation algorithms in the presence of simple geometric heterogeneities: comparison with Monte Carlo calculations. Phys Med Biol 2007; 52: 1363-1385 [PMID: 17301460 DOI: 10.1088/0031-9155/52/5/011]

6	Milan J, Bentley RE. The storage and manipulation of radiation dose data in a small digital computer. Br J Radiol 1974; 47: 115-121 [PMID: 4206210 DOI: 10.1259/0007-1285-47-554-115]

7	Sontag MR, Cunningham JR. The equivalent tissue-air ratio method for making absorbed dose calculations in a heterogeneous medium. Radiology 1978; 129: 787-794 [PMID: 725060]

8	Webb S, Fox RA. Verification by Monte Carlo methods of a power law tissue-air ratio algorithm for inhomogeneity corrections in photon beam dose calculations. Phys Med Biol 1980; 25: 225-240 [PMID: 7384209 DOI: 10.1088/0031-9155/25/2/003]

9	Boyer A, Mok E. A photon dose distribution model employing convolution calculations. Med Phys 1985; 12: 169-177 [PMID: 4000072 DOI: 10.1118/1.595772]

10	Mackie TR, Scrimger JW, Battista JJ. A convolution method of calculating dose for 15-MV x rays. Med Phys 1985; 12: 188-196 [PMID: 4000075 DOI: 10.1118/1.595774]

11	Mohan R, Chui C, Lidofsky L. Differential pencil beam dose computation model for photons. Med Phys 1986; 13: 64-73 [PMID: 3951411 DOI: 10.1118/1.595924]

12	Ahnesjö A. Collapsed cone convolution of radiant energy for photon dose calculation in heterogeneous media. Med Phys 1989; 16: 577-592 [PMID: 2770632 DOI: 10.1118/1.596360]

13	Fogliata A, Nicolini G, Vanetti E, Clivio A, Cozzi L. Dosimetric validation of the anisotropic analytical algorithm for photon dose calculation: fundamental characterization in water. Phys Med Biol 2006; 51: 1421-1438 [PMID: 16510953 DOI: 10.1088/0031-9155/51/6/004]

14	Ahnesjo A, Trepp A. Acquisition of the effective lateral energy fluence distribution for photon beam dose calculations by convolution models. Phys Med Biol 1991; 36: 973-985 [DOI: 10.1088/0031-9155/36/7/006]

15	Ahnesjö A, Saxner M, Trepp A. A pencil beam model for photon dose calculation. Med Phys 1992; 19: 263-273 [PMID: 1584117 DOI: 10.1118/1.596856]

16	Ulmer W, Kaissl W. The inverse problem of a Gaussian convolution and its application to the finite size of the measurement chambers/detectors in photon and proton dosimetry. Phys Med Biol 2003; 48: 707-727 [PMID: 12699190 DOI: 10.1088/0031-9155/48/6/302]

17	Ulmer W, Pyyry J, Kaissl W. A 3D photon superposition/convolution algorithm and its foundation on results of Monte Carlo calculations. Phys Med Biol 2005; 50: 1767-1790 [PMID: 15815095 DOI: 10.1088/0031-9155/50/8/010]

18	Sievinen J, Ulmer W, Kaissl W. AAA photon dose calculation model in Eclipse. Palo Alto (CA): Varian Medical Systems, 2005

19	Gagné IM, Zavgorodni S. Evaluation of the analytical anisotropic algorithm in an extreme water-lung interface phantom using Monte Carlo dose calculations. J Appl Clin Med Phys 2007; 8: 33-46 [PMID: 17592451]

20	Vanderstraeten B, Reynaert N, Paelinck L, Madani I, De Wagter C, De Gersem W, De Neve W, Thierens H. Accuracy of patient dose calculation for lung IMRT: A comparison of Monte Carlo, convolution/superposition, and pencil beam computations. Med Phys 2006; 33: 3149-3158 [PMID: 17022207 DOI: 10.1118/1.2241992]

21	Hasenbalg F, Neuenschwander H, Mini R, Born EJ. Collapsed cone convolution and analytical anisotropic algorithm dose calculations compared to VMC++ Monte Carlo simulations in clinical cases. Phys Med Biol 2007; 52: 3679-3691 [PMID: 17664570 DOI: 10.1088/0031-9155/52/13/002]

22	Ding GX, Duggan DM, Lu B, Hallahan DE, Cmelak A, Malcolm A, Newton J, Deeley M, Coffey CW. Impact of inhomogeneity corrections on dose coverage in the treatment of lung cancer using stereotactic body radiation therapy. Med Phys 2007; 34: 2985-2994 [PMID: 17822007 DOI: 10.1118/1.2745923]

23	Matsuo Y, Takayama K, Nagata Y, Kunieda E, Tateoka K, Ishizuka N, Mizowaki T, Norihisa Y, Sakamoto M, Narita Y, Ishikura S, Hiraoka M. Interinstitutional variations in planning for stereotactic body radiation therapy for lung cancer. Int J Radiat Oncol Biol Phys 2007; 68: 416-425 [PMID: 17363190 DOI: 10.1016/j.ijrobp.2006.12.012]

24	Panettieri V, Wennberg B, Gagliardi G, Duch MA, Ginjaume M, Lax I. SBRT of lung tumours: Monte Carlo simulation with PENELOPE of dose distributions including respiratory motion and comparison with different treatment planning systems. Phys Med Biol 2007; 52: 4265-4281 [PMID: 17664607 DOI: 10.1088/0031-9155/52/14/016]

25	Lax I, Panettieri V, Wennberg B, Amor Duch M, Näslund I, Baumann P, Gagliardi G. Dose distributions in SBRT of lung tumors: Comparison between two different treatment planning algorithms and Monte-Carlo simulation including breathing motions. Acta Oncol 2006; 45: 978-988 [PMID: 16982567 DOI: 10.1080/02841860600900050]

26	Schuring D, Hurkmans CW. Developing and evaluating stereotactic lung RT trials: what we should know about the influence of inhomogeneity corrections on dose. Radiat Oncol 2008; 3: 21 [PMID: 18662379 DOI: 10.1186/1748-717X-3-21]

27	Xiao Y, Papiez L, Paulus R, Timmerman R, Straube WL, Bosch WR, Michalski J, Galvin JM. Dosimetric evaluation of heterogeneity corrections for RTOG 0236: stereotactic body radiotherapy of inoperable stage I-II non-small-cell lung cancer. Int J Radiat Oncol Biol Phys 2009; 73: 1235-1242 [PMID: 19251095 DOI: 10.1016/j.ijrobp.2008.11.019]

28	Li J, Galvin J, Harrison A, Timmerman R, Yu Y, Xiao Y. Dosimetric verification using monte carlo calculations for tissue heterogeneity-corrected conformal treatment plans following RTOG 0813 dosimetric criteria for lung cancer stereotactic body radiotherapy. Int J Radiat Oncol Biol Phys 2012; 84: 508-513 [PMID: 22365630 DOI: 10.1016/j.ijrobp.2011.12.005]

29	Dutreix A. When and how can we improve precision in radiotherapy? Radiother Oncol 1984; 2: 275-292 [PMID: 6522641 DOI: 10.1016/S0167-8140(84)80070-5]

30	Chetty IJ, Devpura S, Liu D, Chen D, Li H, Wen NW, Kumar S, Fraser C, Siddiqui MS, Ajlouni M, Movsas B. Correlation of dose computed using different algorithms with local control following stereotactic ablative radiotherapy (SABR)-based treatment of non-small-cell lung cancer. Radiother Oncol 2013; 109: 498-504 [PMID: 24231237 DOI: 10.1016/j.radonc.2013.10.012]

31	Nahum AE, Sanchez-Nieto B. Tumour control probability modelling: basic principles and applications in treatment planning. Phys Med 2001; 17: 13-23

32	Webb S, Nahum AE. A model for calculating tumour control probability in radiotherapy including the effects of inhomogeneous distributions of dose and clonogenic cell density. Phys Med Biol 1993; 38: 653-666 [PMID: 8346278 DOI: 10.1088/0031-9155/38/6/001]

33	Bufacchi A, Nardiello B, Capparella R, Begnozzi L. Clinical implications in the use of the PBC algorithm versus the AAA by comparison of different NTCP models/parameters. Radiat Oncol 2013; 8: 164 [PMID: 23826854 DOI: 10.1186/1748-717X-8-164]

34	Nielsen TB, Wieslander E, Fogliata A, Nielsen M, Hansen O, Brink C. Influence of dose calculation algorithms on the predicted dose distribution and NTCP values for NSCLC patients. Med Phys 2011; 38: 2412-2418 [PMID: 21776775 DOI: 10.1118/1.3575418]

35	Källman P, Agren A, Brahme A. Tumour and normal tissue responses to fractionated non-uniform dose delivery. Int J Radiat Biol 1992; 62: 249-262 [PMID: 1355519 DOI: 10.1080/09553009214552071]

36	Lyman JT. Complication probability as assessed from dose-volume histograms. Radiat Res Suppl 1985; 8: S13-S19 [PMID: 3867079 DOI: 10.2307/3576626]

37	Kutcher GJ, Burman C. Calculation of complication probability factors for non-uniform normal tissue irradiation: the effective volume method. Int J Radiat Oncol Biol Phys 1989; 16: 1623-1630 [PMID: 2722599 DOI: 10.1016/0360-3016(89)90972-3]

38	Chandrasekaran M, Panettieri V, Baker C, and Nahum AE. The clinical impact of differences in photon dose algorithms: 3DCRT and SBRT of lung tumors. Radiother Oncol 2011; 99: S427-S428 [DOI: 10.1016/S0167-8140(11)71269-5]

39	Liu MB, Eclov NC, Trakul N, Murphy J, Diehn M, Le QT, Dieterich S, Maxim PG, Loo BW. Clinical impact of dose overestimation by effective path length calculation in stereotactic ablative radiation therapy of lung tumors. Pract Radiat Oncol 2013; 3: 294-300 [PMID: 24674401 DOI: 10.1016/J.PRRO.2012.09.003]

P- Reviewer: Xiao Y    S- Editor: Ji FF    L- Editor: A    E- Editor: Lu YJ

# Text boxes

Table 3  Clinical impact of dose calculation algorithms

| Ref. | Tumor site/technique | Algorithms studied | Results/conclusion |
| --- | --- | --- | --- |
| Nielsen et al[34], 2011 | NSCLC | Eclipse AAA OTP CC Pinnacle CC XiO Sup OTP PB XiO FFT | Differences in dose to target predicted by the different algorithms are of a magnitude. Calculated NTCP values for pneumonitis are more sensitive to the choice of algorithm than mean lung dose and V20 |
| Chandrasekaran et al[38], 2011 | Lung/3DCRT,SBRT | PBC, Eclipse AAA, Pinnacle CCC, Masterplan PBC and CCC | PBC yielded higher TCP in comparison with other algorithms. For small tumor, TCP was overestimated by 4%-13% by PBC; for large tumor, there was an increase of up to 6%-22% |
| Liu et al[39], 2013 | Lung/SABR | EPL, MC | EPL overestimates dose by amounts that substantially decrease TCP in a large proportion. Compared with MC, prescribing based on EPL translated to a median TCP decrement of 4.3% (range, 1.2%-37%) and a > 5% decrement in 46% of tumors |
| Bufacchi et al[33], 2013 | Prostate, HN, Lung, Breast /3DCRT | PBC, AAA | NTCP calculated with AAA was lower than the NTCP calculated with PBC, except for the breast treatments |
| Chetty et al[30], 2013 | NSCLC/SABR | EPL-1D, EPL-3D, AAA, CCC, Acuros, MC | Average TCP decrements (5%-10%, ranging up to approximately 50%) were observed with model-based algorithms relative to the EPL-based methods |

Eclipse AAA: Eclipse Anisotropic Analytical Algorithm; OTP CC: Oncentra MasterPlan Collapsed Cone algorithm; Pinnacle CC: Pinnacle Collapsed Cone algorithm; XiO Sup: XiO Multigrid Superposition algorithm; OTP PB: Oncentra MasterPlan Pencil Beam algorithm; XiO FFT: XiO Fast Fourier Transform Convolution algorithm; EPL: Equivalent path length; MC: Monte Carlo.

Table 2  Relative differences calculated as (without-with)/with density corrections using each algorithm

|  | Eclipse AAA | OTP CC | Pinnacle CC | XiO Sup | OTP PB | XiO FFT |
| --- | --- | --- | --- | --- | --- | --- |
| Combined lungs |  |  |  |  |  |  |
| NTCPBurman | -0.29 | -0.2 | -0.22 | -0.25 | -0.36 | -0.45 |
| NTCPSeppenwoolde | -0.19 | -0.13 | -0.12 | -0.15 | -0.23 | -0.3 |
| Mean dose | -0.08 | -0.05 | -0.05 | -0.06 | -0.09 | -0.13 |
| V20 | -0.06 | -0.06 | -0.04 | -0.03 | -0.05 | -0.07 |
| Heart |  |  |  |  |  |  |
| NTCP | -0.19 | -0.15 | -0.15 | -0.13 | -0.17 | -0.21 |
| Mean dose | -0.06 | -0.05 | -0.05 | -0.05 | -0.06 | -0.09 |
| V50 | -0.11 | -0.1 | -0.1 | -0.08 | -0.08 | -0.13 |
| PTV |  |  |  |  |  |  |
| Mean dose | -0.06 | -0.05 | -0.05 | -0.05 | -0.13 | -0.1 |
| D01 | -0.05 | -0.05 | -0.04 | -0.04 | -0.1 | -0.11 |
| D99 | -0.07 | -0.04 | -0.05 | -0.05 | -0.14 | -0.09 |
| GTV |  |  |  |  |  |  |
| Mean dose | -0.07 | -0.06 | -0.06 | -0.06 | -0.08 | -0.1 |
| D01 | -0.07 | -0.07 | -0.06 | -0.06 | -0.09 | -0.11 |
| D99 | -0.07 | -0.06 | -0.06 | -0.06 | -0.07 | -0.1 |

Negative results indicate lower values when no density corrections are included. Eclipse AAA: Eclipse Anisotropic Analytical Algorithm; OTP CC: Oncentra MasterPlan Collapsed Cone algorithm; Pinnacle CC: Pinnacle Collapsed Cone algorithm; XiO Sup: XiO Multigrid Superposition algorithm; OTP PB: Oncentra MasterPlan Pencil Beam algorithm; XiO FFT: XiO Fast Fourier Transform Convolution algorithm. (Cited from Nielsen et al[34] 2011).

Table 1  Calculated percent mean tumor control probability values (ranges in parentheses) for all algorithms as a function of planning target volume volume

| PTV bins (cm3) | Mean PTV volume (range, cm3) | n | EPL-1D | EPL-3D | AAA | CCC | Acuros | MC |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 4 ≤ v < 10 | 7.8 (4.8-9.9) | 15 | 100.0 (100-100) | 99.9 (99.6-100) | 93.1 (76.3-99.8) | 91.3 (63.0-99.9) | 91.8 (60.8-99.8) | 90.5 (51.1-99.9) |
| 10 ≤ v < 20 | 15.0 (10.4-19.8) | 27 | 100.0 (99.8-100) | 99.9 (99.5-100) | 91.3 (61.7-100) | 91.3 (50.4-100) | 91.4 (65.4-99.9) | 91.1 (53.2-100) |
| 20 ≤ v < 30 | 24.3 (20.4-29.6) | 29 | 98.5 (99.8-100) | 98.9 (77.6-100) | 92.7 (74.9-99.9) | 90.5 (46.4-99.9) | 90.9 (65.1-99.9) | 91.1 (48.4-99.9) |
| 30 ≤ v < 40 | 34.9 (30.2-39.8) | 18 | 99.8 (97.4-100) | 99.6 (98.2-100) | 92.0 (63.4-99.9) | 92.1 (69.7-99.9) | 90.9 (61.6-99.8) | 92.4 (56.3-99.9) |
| 40 ≤ v < 60 | 47.3 (40.2-58.4) | 17 | 99.5 (93.1-100) | 99.1 (95.6-100) | 92.6 (78.6-99.9) | 91.4 (64.4-99.9) | 93.6 (77.6-99.9) | 92.3 (63.6-99.9) |
| 60 ≤ v < 100 | 78.0 (60.4-95.9) | 16 | 99.5 (95.6-100) | 99.0 (95.8-100) | 92.7 (70.7-99.8) | 92.8 (66.2-99.9) | 93.4 (70.4-99.8) | 94.7 (74.6-99.9) |
| V ≥ 100 | 162.4 (100.5-360.2) | 11 | 99.2 (96.1-99.9) | 98.7 (95.0-100) | 96.3 (89.9-100) | 95.6 (91.6-99.8) | 95.3 (83.0-99.9) | 97.1 (88.8-99.9) |

PTV: Planning target volume; EPL-1D: 1-D equivalent path-length (pencil beam-type); EPL-3D: 3-D equivalent-path-length (pencil beam-type); AAA: Anisotropic analytical algorithm; CCC: Collapsed cone convolution-superposition; Acuros: Acuros AXB; MC: Monte Carlo. (Cited from Chetty et al[30] 2013).