| Dual mobility cups in total hip arthroplasty Ivan De Martino, Georgios Konstantinos Triantafyllopoulos, Peter Keyes Sculco, Thomas Peter Sculco |
| --- |
| CITATION | De Martino I, Triantafyllopoulos GK, Sculco PK, Sculco TP. Dual mobility cups in total hip arthroplasty. World J Orthop 2014; 5(3): 180-187 |
| URL | http://www.wjgnet.com/2218-5836/full/v5/i3/180.htm |
| DOI | http://dx.doi.org/10.5312/wjo.v5.i3.180 |
| OPEN ACCESS | Articles published by this Open-Access journal are distributed under the terms of the Creative Commons Attribution Non-commercial License, which permits use, distribution, and reproduction in any medium, provided the original work is properly cited, the use is non commercial and is otherwise in compliance with the license. |
| CORE TIP | Instability remains a significant issue after both primary and revision total hip arthroplasty. Dual mobility or tripolar unconstrained acetabular components can provide a viable alternative in preventing and treating instability. Reported outcomes of several European studies using dual mobility cups with mid- to long-term follow up support their effectiveness. Concerns such as intra-prosthetic dislocation and accelerated wear have been emphasized, although they seem to be less significant in older, low-demand patients. The use of dual mobility cups in younger patients should be viewed with caution based on a lack of current data concerning this high demand patient population. |
| KEY WORDS | Dual mobility; Total hip arthroplasty; Dislocation; Revision total hip arthroplasty; Instability |
| COPYRIGHT | © 2014 Baishideng Publishing Group Inc. All rights reserved. |
| COPYRIGHT LICENSE | Order reprints or request permissions: bpgoffice@wjgnet.com |
| NAME OF JOURNAL | World Journal of Orthopedics |
| ISSN | 2218-5836 ( online) |
| PUBLISHER | Baishideng Publishing Group Co., Limited, 8226 Regency Drive, Pleasanton, CA 94588, USA |
| WEBSITE | http://www.wjgnet.com |

ESPS Manuscript NO: 8773

Columns: TOPIC HIGHLIGHT

Dual mobility cups in total hip arthroplasty

Ivan De Martino, Georgios Konstantinos Triantafyllopoulos, Peter Keyes Sculco, Thomas Peter Sculco

Ivan De Martino, Georgios Konstantinos Triantafyllopoulos, Peter Keyes Sculco, Thomas Peter Sculco, Department of Orthopaedic Surgery, Division of Adult Reconstruction and Joint Replacement, Hospital for Special Surgery, New York, NY 10021, United States

Author contributions: De Martino I, Triantafyllopoulos GK and Sculco PK performed the literature review and drafted the article; Sculco TP revised the article critically for important intellectual content; De Martino I, Triantafyllopoulos GK, Sculco PK and Sculco TP finally approved for print.

Correspondence to: Thomas Peter Sculco, MD, Department of Orthopedic Surgery, Division of Adult Reconstruction and Joint Replacement, Hospital for Special Surgery, 535 East 70th Street, New York, NY 10021, United States. sculcot@hss.edu

Telephone: +1-212-6061475  Fax: +1-212-7349572

Received: January 6, 2014      Revised: February 21, 2014   Accepted: April 17, 2014

Published online: July 18, 2014

Abstract

Total hip arthroplasty (THA) is considered one of the most successful surgical procedures in orthopaedics. With the increase in the number of THAs performed in the world in the next decades, reducing or preventing medical and mechanical complications such as post-operative THA instability will be of paramount importance, particularly in an emerging health care environment based on quality control and patient outcome. Dual mobility acetabular component (also known as unconstrained tripolar implant) was introduced in France at the end of the 1970s as an alternative to standard sockets, to reduce the risk of THA dislocation in patients undergoing primary THA in France. Dual mobility cups have recently gained wider attention in the United States as an alternative option in the prevention and treatment of instability in both primary and revision THA and offer the benefit of increased stability without compromising clinical outcomes and implant longevity. In this article, we review the use of dual mobility cup in total hip arthroplasty in terms of its history, biomechanics, outcomes and complications based on more than 20 years of medical literature.

© 2014 Baishideng Publishing Group Inc. All rights reserved.

Key words: Dual mobility; Total hip arthroplasty; Dislocation; Revision total hip arthroplasty; Instability

Core tip: Instability remains a significant issue after both primary and revision total hip arthroplasty. Dual mobility or tripolar unconstrained acetabular components can provide a viable alternative in preventing and treating instability. Reported outcomes of several European studies using dual mobility cups with mid- to long-term follow up support their effectiveness. Concerns such as intra-prosthetic dislocation and accelerated wear have been emphasized, although they seem to be less significant in older, low-demand patients. The use of dual mobility cups in younger patients should be viewed with caution based on a lack of current data concerning this high demand patient population.

De Martino I, Triantafyllopoulos GK, Sculco PK, Sculco TP. Dual mobility cups in total hip arthroplasty. World J Orthop 2014; 5(3): 180-187  Available from: URL: http://www.wjgnet.com/2218-5836/full/v5/i3/180.htm  DOI: http://dx.doi.org/10.5312/wjo.v5.i3.180

INTRODUCTION

Total hip arthroplasty (THA) is considered one of the most successful surgical procedures providing pain relief and improvement of function in patients with end-stage hip arthritis that is non-responsive to non-operative treatments[1,2]. As health care continues to improve and life expectancy increases, the demand for total joint replacement will grow to reflect this more active, aging population. The number of THAs performed in the United States is projected to reach 572000 by 2030, an increase of 174% compared to 2005[3].

Reducing or preventing medical and mechanical complications such as post-operative THA instability will be of paramount importance, particularly in an emerging health care environment based on quality control and patient outcome. The incidence of instability after THA in the primary and revision setting has been reported as high as 7% and 25% respectively[4]. Risk factors for instability after THA are multifactorial and may be patient-specific (gender, age, abductor deficiency) or related to operative variables (surgical approach, component malposition, femoral head diameter)[5]. Instability after THA remains one of the major causes of readmission and revision surgery accounting for 32.4% of THA readmissions and 22.5% of all THA revisions in the United States[6,7]. Readmission and revision surgery carry considerable economic cost as the surgical treatment of a dislocating THA can raise cost 148%[8]. Modifications in surgical technique (e.g., anterior surgical approach, repair of posterior soft-tissues, increased offset and restoration of abductor tension) and the incorporation of larger femoral heads with greater inherent stability decrease the risk of instability after THA. Conversion to a bipolar arthroplasty and a constrained liner are salvage procedures for recurrent instability that provide stability but reduce functional outcome and implant longevity. Dual mobility acetabular components (also known as unconstrained tripolar implants) have recently gained wider attention in the United States as an alternative option in the prevention and treatment of instability in both primary and revision THA and offer the benefit of increased stability without compromising clinical outcomes and implant longevity.

HISTORY OF AND EVOLUTION OF DUAL MOBILITY CUPS

The dual articulation cup was developed by Professor Gilles Bousquet and André Rambert (engineer) in 1974 and combined the “low friction” principle of THA popularized by Charnley[9] with the McKee-Farrar concept of using a larger diameter femoral head to enhance implant stability[10]. The goal of the dual articulation was to achieve the greatest possible range of motion in a stable environment in addition to reducing wear. The original design (Novae-1®, Serf, Décines, France) incorporated a 22.2 mm metallic head articulating with a polyethylene liner, which in turn articulated with the acetabular shell. The shell was manufactured from stainless steel, coated with a porous plasma sprayed alumina (AL2O3) and had a cylindrical/spherical configuration. A three-point fixation system consisted of two Morse taper pegs, for impaction into the ischiopubic ramus and the ischium, and a bicortical iliac screw designed to enhance press-fit cup fixation. The liner was made from ultra-high molecular weight polyethylene (UHMWPE), gamma sterilized in air.

Modifications and improvements were made to the mechanics, metallurgy, and materials of the original design: titanium and hydroxyapatite replaced alumina coating[11], flanges and modular shells were added for screw fixation[12], highly cross-linked UHMWPE enriched with vitamin-E improved wear[13], larger femoral heads added stability[14], and anatomic designs decreased anterior overhang[15]. Advances in polyethylene manufacturing and sterilization decreased risk of catastrophic volumetric wear and allowed for the use of larger femoral heads and the use of a 10/12 Morse taper and a highly polished neck reduced liner impingement[16]. While the dual mobility was intended for primary and revision THA with minimal bone loss, cemented designs with concomitant impaction grafting were introduced for cases with more significant bone loss[17].

Dual mobility cups have been in clinical use for many years in Europe, but did not receive U.S. Food and Drug Administration approval until 2009. The designs currently available include the POLARCUP® (Smith and Nephew Orthopaedics AG, Rotkreuz, Switzerland), Anatomic Dual Mobility (ADM®) (Stryker, Mahwah, NJ), Active Articulation E1® (Biomet, Warsaw, IN) (Stick/K-Arm) and uncemented variations (SunFit TH, Coptos TH, Evolution TH) of the original Novae® cup (Serf, Décines, France). The POLARCUP® offers both cemented and press-fit options with the use of pegs and screws. The shell consists of a plasma sprayed titanium fixation surface and a stainless steel bearing surface. The Anatomic Dual Mobility (ADM®) (Stryker, Mahwah, NJ) also includes a titanium plasma sprayed fixation surface, but has a cobalt-chrome bearing surface and features an anatomic design with a recess in the shell to accommodate the iliopsoas tendon and reduce impingement symptoms. The Active Articulation E1® (Biomet, Warsaw, IN) integrates a vitamin-E impregnated polyethylene with a cobalt-chrome bearing surface. Again, fixation is promoted through osseointegration with a plasma sprayed titanium surface. The Modular Dual Mobility X3® (MDM®) (Stryker, Mahwah, NJ) uses a shell with screw holes for additional fixation and a modular highly polished cobalt-chrome liner which articulates with polyethylene. The MDM offers the advantage of screw fixation and the use of a standard shell which is available in hospital inventories and can be implanted with familiar instrumentation. The MDM acetabular shell can be converted to a dual mobility component by placing the metallic insert[18].

BIOMECHANICS

The dual mobility component increases hip range of motion (ROM) until impingement occurs through its two articulations design. In the first articulation the head is “engaged” but mobile within the polyethylene (PE) liner and follows the typical mechanical behavior of a hard-on-soft bearing in a standard THA. However, if the femoral neck and the rim of the PE liner come into contact, a second articulation begins to function and consists of the back of the PE liner and the metallic acetabular shell. As the PE liner articulates, effective ROM is increased until impingement of the femoral neck against the rim of the shell ultimately occurs (Figure 1). In this way, the head-liner complex theoretically functions as a large femoral head, increasing the head-neck ratio and subsequently the jump distance before dislocation. In an experimental setting, dual mobility cups with 22.2 mm and 28 mm femoral heads demonstrated significantly greater ROM compared to conventional implants with similar head sizes[19]. While there was no statistically significant difference between the two dual mobility head sizes and ROM, a larger head increases the range of motion before impingement of the neck against the PE liner, theoretically reducing the risk of intraprosthetic dislocation (IPD)[16].

OUTCOMES

Dual mobility in primary THA

Several studies on DM cups in primary THA have reported a low rate of postoperative implant instability[19-22]. Farizon[23], Philippot[12], Lautridou[24], Vielpau[21] and Boyer[25] reported their experience on the use of first generation Bousquet cups (Novae®, Serf, Décines, France) with a 22.2 mm metal head and conventional PE. At 15 years follow-up, survivorship ranged from 81.4% to 96.3% with a dislocation rate (with large articulation) between 0% and 1%. However, these authors did not include the dislocation rate of the femoral head and the mobile PE bearing (the small articulation) that ranged from 0% to 5.2%. Causes of cup failure included aseptic loosening (1.8%-3.4%), excessive PE wear (1%-2%) and acetabular screw fracture (1%). Guyen[19], Leclercq[20] and Vielpeau[21] published series of 167, 200, and 231 primary THA patients using current DM designs with a follow-up time period of 3 to 6 years and reported a 0% dislocation rate.

Dual mobility in revision THA

Dislocation rate after revision THA ranges from 5% to 30%[26-29]. Many factors have been implicated in postoperative instability including muscular insufficiency, aggressive capsulectomy, bone loss and implant positioning problems[30]. Leiber-Wackenheim[14], Hamadouche[31], Langlais[17] and Guyen[16] reported their results on the use of DM cups in patients revised for instability after primary THA. Survivorship of the cups at a mean follow-up period of 5 years was between 94.5% and 98% with a dislocation rate of 1.1% to 5.5%. These studies suggest DM cups are a reliable treatment option for patients revised for instability after primary THA.

Dual mobility in femoral neck fractures

Femoral neck fractures (FNF) treated with osteosynthesis have an increased risk of reoperation when compared to hip arthroplasty[32]. Although THA showed better functional results than osteosynthesis in FNF treatments[33], prosthetic dislocation remains a serious problem. In a recent meta-analysis by Iorio et al[34] the mean dislocation rate was 10.7% in patients with FNF treated with THA, five times higher than THA for osteoarthritis. Adam et al[35] reported 3 dislocations (1.4%) at 9 mo follow-up in a series of 214 patients with FNF treated with DM implants. Tarasevicius et al[36] compared dislocation rates of DM cups with that of conventional cups in patients with FNF treated with THA through a posterior approach. At 1 year follow-up, there were 8 dislocations (14.3%) in the conventional THA group and no dislocations in DM group. DM cups may also be considered as an option to prevent postoperative dislocation when treating FNFs in elderly patients who are candidates for THA.

Dual mobility in tumor resection

THA after tumor resection has also been associated with a high risk of dislocation due to bone loss and soft tissue compromise. Philippeau et al[37] retrospectively analysed 71 patients with bony lesions of the hip treated with a THA and DM cups. They reported 7 postoperative dislocations (9.8%). Dislocation rate was lower when abductors were preserved (3.5%) and higher when abductors were sectioned/reattached (9.5%) and when the gluteus medius muscle or nerve were resected (18%). They also reported acetabular loosening in 4 cases (5.6%).

Dual mobility in spastic disorders

Several studies on THAs in patients with cerebral palsy (CP) showed good results on pain relief and function outcome; however the dislocation rate in this challenging patient group is reported to be as high as 14%[38-43]. Sanders et al[44] reported on 10 hips (8 patients) with CP treated[41-43,45] with THA and a DM cup and had no dislocations after a mean follow-up of 39 mo.

Indications and contraindications

The original goal of the DM cup, introduced at the end of the 1970s as an alternative to standard sockets, was to reduce the risk of THA dislocation in patients undergoing primary THA. Currently, DM cups are a well-accepted treatment option for any patient at an elevated risk for instability after primary or revision THA and in the treatment of recurrent dislocation[12,14,16,17,20,23-25,30,31,46-48]. Patients at higher risk of dislocation include patients with neuromuscular diseases, cognitive dysfunction, an American Society of Anesthesiologists score of 3 or more, and all patients older than 75 years with a history of prior hip surgery[28,49-53]. In addition, the use of DM cups is indicated in revision THA for any cause[17,30], primary THA after femoral neck fracture[35,36], and primary THA after tumor resection[37]. While several studies demonstrate a reduction in the dislocation rate in patients over 60 years, limited data is currently available on active patients younger than 50 years old and care should be taken when using DM cups in a population more prone to develop wear and osteolysis[20,21,24,48,54].

Special considerations and complications

Dual mobility acetabular components are associated with some specific complications secondary to its dual articulating design. For example, intra-prosthetic dislocation or retentive failure is a complication observed exclusively with this type of implant and involves failure of the articulation between the femoral head and the PE liner. The proposed mechanism for dissociation is a result of wear of the PE liner’s retentive chamfer[25]. After dissociation, the head articulates directly with the metallic bearing surface of the acetabular shell, producing acute limb shortening and limp. Furthermore, as the shell is not designed for a metal-on-metal articulation, friction between its bearing surface and the femoral head results in rapid wear, metal ion release and surrounding soft-tissue metallosis[55]. In plain X-rays, the asymmetric position of the femoral head within the cup can be visualized which may be mistakenly attributed to “polyethylene wear”. However, the characteristic “bubble sign” which corresponds to the dislocated liner is pathognomonic of retentive failure (Figure 2). Management is dependent on the time interval from dislocation to diagnosis. If IPD is diagnosed early, before significant wear of the femoral head and acetabular shell occurs, it can be treated with simple liner exchange. In cases of late diagnosis, revision of the acetabular component, as well as femoral head exchange may be necessary due to femoral head and acetabular shell damage. Boyer et al[25] in a series of 240 hips followed for 9 years and 11 months reported a 4.1% incidence of IPD. A similar incidence (4%) was reported by Philippot et al[15] among 1960 primary THAs with a mean follow-up of 14 years. The authors recognized three distinct types of IPD: type 1, which was typically due to liner wear; type 2, which was related with arthrofibrosis blocking the liner; and type 3, which was associated with cup loosening. When comparing the two stems with different neck diameters and incidence of retentive failure, no statistically significant difference was observed. The authors note that the narrower neck was unpolished titanium (which is rougher than stainless steel used in the larger neck diameter stem) and could have counteracted the positive effects of the smaller neck size. In their series of 231 primary THAs where a second generation dual mobility cup was used, Vielpeau et al[21] reported 0% retentive failure rate at 5.2 years. In 437 hips using the original Bousquet implant design, intra-prosthetic dislocation after a mean of 16.2 years was observed in 3 hips. The authors attributed the low incidence of intra-prosthetic dislocation to the smooth, polished, and narrow femoral neck. In other studies with mid- to long-term follow-up, the incidence of retentive failure ranges from 0% to 5.2%[12,14,19,20,47,48,56]. Table 1 summarizes reported retentive failure rates in the literature.

The configuration of the dual mobility cup with its two articulations and thinner liner has raised concern for accelerated PE wear and associated osteolysis. In a retrieval study of liners removed after revision surgery for infection or aseptic loosening, no difference in total volumetric polyethylene wear was noted between tripolar unconstrained cups and conventional cups with 22.2 mm heads[57]. However, greater wear was noted at the convex bearing surface of the liner. Failure rate due to accelerated polyethylene wear was 2% in a series of patients with first generation PE[25]. In the same series, age < 50 years was associated with significantly greater wear rates, apparently due to the higher activity level of these patients. Similarly, in a series of Philippot et al[48] revision rate due to PE wear was 1.6%. Combes et al[22] reported a 7% rate of osteolysis (with first-generation PE), especially in patients of younger age and those treated for sequelae of childhood hip disease.

Radiographic evaluation of PE wear can be difficult in the setting of a dual mobility component, because of the deep position of the head within the cup and the cylindrical-spherical shape of the shell itself. An eccentric femoral head implies concomitant wear of both the concave and the convex bearing surface of the liner[58]. Highly cross-linked UHMWPE and vitamin-E impregnated polyethylene has reduced volumetric wear in standard implants[13,59] and have been integrated into dual mobility implants in an effort to deal with accelerated wear issues. Bauchu et al[60] in a retrospective series of primary THAs with the POLARCUP® component (Smith and Nephew Orthopaedics AG, Rotkreuz, Switzerland), which incorporates a highly cross-linked UHMWPE liner, reported no incidents of wear-related osteolysis. However, there are currently no independent studies of tripolar cups further supporting these findings. For all these reasons, the use of dual mobility cups should be used with caution in younger patients with high demands and increased risk of wear-related osteolysis.

Another issue with the DM cup is aseptic loosening. Loss of fixation of the original design was attributed to delamination of the plasma sprayed alumina layer[23] which led to design modifications. Some authors propose that fixation of tripolar cups, particularly second-generation implants, should always be supplemented with screws[11]. Nonetheless, as noted earlier, current shell options include monoblock cups, which take advantage of modern porous coated surfaces for enhanced osseointegration. A potential drawback of these monoblock cups is the difficulty assessing proper seating of the cup within the acetabulum which may contribute to the reported rates of aseptic loosening ranging from 0% to 8.3%[12,14,19-22,47,48,56] (Table 1).

CONCLUSION

Instability remains a significant issue after both primary and revision THA. Dual mobility or tripolar unconstrained acetabular components can provide a viable alternative in preventing and treating instability. Reported outcomes of studies using DM cups with mid- to long-term follow up support their effectiveness. Concerns such as intra-prosthetic dislocation and accelerated wear have been emphasized, although they seem to be less significant in older, low-demand patients. The use of dual mobility cups in younger patients should be viewed with caution based on a lack of current data concerning this high demand patient population.

REFERENCES

1	Austin MS, Higuera CA, Rothman RH. Total hip arthroplasty at the rothman institute. HSS J 2012; 8: 146-150 [PMID: 23874255 DOI: 10.1007/s11420-012-9268-x]

2	NIH consensus conference: Total hip replacement. NIH Consensus Development Panel on Total Hip Replacement. JAMA 1995; 273: 1950-1956 [PMID: 7783307 DOI: 10.1001/jama.1995.03520480070043]

3	Nho SJ, Kymes SM, Callaghan JJ, Felson DT. The burden of hip osteoarthritis in the United States: epidemiologic and economic considerations. J Am Acad Orthop Surg 2013; 21 Suppl 1: S1-S6 [PMID: 23818185 DOI: 10.5435/JAAOS-21-07-S1]

4	Patel PD, Potts A, Froimson MI. The dislocating hip arthroplasty: prevention and treatment. J Arthroplasty 2007; 22: 86-90 [PMID: 17570285 DOI: 10.1016/j.arth.2006.12.111]

5	Prudhon JL, Ferreira A, Verdier R. Dual mobility cup: dislocation rate and survivorship at ten years of follow-up. Int Orthop 2013; 37: 2345-2350 [PMID: 24026216 DOI: 10.1007/s00264-013-2067-2]

6	Bozic KJ, Kurtz SM, Lau E, Ong K, Vail TP, Berry DJ. The epidemiology of revision total hip arthroplasty in the United States. J Bone Joint Surg Am 2009; 91: 128-133 [PMID: 19122087 DOI: 10.2106/JBJS.H.00155]

7	Schairer WW, Sing DC, Vail TP, Bozic KJ. Causes and frequency of unplanned hospital readmission after total hip arthroplasty. Clin Orthop Relat Res 2014; 472: 464-470 [PMID: 23801061 DOI: 10.1007/s11999-013-3121-5]

8	Sanchez-Sotelo J, Haidukewych GJ, Boberg CJ. Hospital cost of dislocation after primary total hip arthroplasty. J Bone Joint Surg Am 2006; 88: 290-294 [PMID: 16452739 DOI: 10.2106/JBJS.D.02799]

9	Charnley J. The long-term results of low-friction arthroplasty of the hip performed as a primary intervention. J Bone Joint Surg Br 1972; 54: 61-76 [PMID: 5011747]

10	McKee GK, Watson-Farrar J. Replacement of arthritic hips by the McKee-Farrar prosthesis. J Bone Joint Surg Br 1966; 48: 245-259 [PMID: 5937593]

11	Massin P, Orain V, Philippot R, Farizon F, Fessy MH. Fixation failures of dual mobility cups: a mid-term study of 2601 hip replacements. Clin Orthop Relat Res 2012; 470: 1932-1940 [PMID: 22161085 DOI: 10.1007/s11999-011-2213-3]

12	Philippot R, Camilleri JP, Boyer B, Adam P, Farizon F. The use of a dual-articulation acetabular cup system to prevent dislocation after primary total hip arthroplasty: analysis of 384 cases at a mean follow-up of 15 years. Int Orthop 2009; 33: 927-932 [PMID: 18521598 DOI: 10.1007/s00264-008-0589-9]

13	Oral E, Christensen SD, Malhi AS, Wannomae KK, Muratoglu OK. Wear resistance and mechanical properties of highly cross-linked, ultrahigh-molecular weight polyethylene doped with vitamin E. J Arthroplasty 2006; 21: 580-591 [PMID: 16781413 DOI: 10.1016/j.arth.2005.07.009]

14	Leiber-Wackenheim F, Brunschweiler B, Ehlinger M, Gabrion A, Mertl P. Treatment of recurrent THR dislocation using of a cementless dual-mobility cup: a 59 cases series with a mean 8 years’ follow-up. Orthop Traumatol Surg Res 2011; 97: 8-13 [PMID: 21273156 DOI: 10.1016/j.otsr.2010.08.003]

15	Philippot R, Boyer B, Farizon F. Intraprosthetic dislocation: a specific complication of the dual-mobility system. Clin Orthop Relat Res 2013; 471: 965-970 [PMID: 23054529 DOI: 10.1007/s11999-012-2639-2]

16	Guyen O, Pibarot V, Vaz G, Chevillotte C, Béjui-Hugues J. Use of a dual mobility socket to manage total hip arthroplasty instability. Clin Orthop Relat Res 2009; 467: 465-472 [PMID: 18780135 DOI: 10.1007/s11999-008-0476-0]

17	Langlais FL, Ropars M, Gaucher F, Musset T, Chaix O. Dual mobility cemented cups have low dislocation rates in THA revisions. Clin Orthop Relat Res 2008; 466: 389-395 [PMID: 18196422 DOI: 10.1007/s11999-007-0047-9]

18	Stulberg SD. Dual poly liner mobility optimizes wear and stability in THA: affirms. Orthopedics 2011; 34: e445-e448 [PMID: 21902126 DOI: 10.3928/01477447-20110714-23]

19	Guyen O, Chen QS, Bejui-Hugues J, Berry DJ, An KN. Unconstrained tripolar hip implants: effect on hip stability. Clin Orthop Relat Res 2007; 455: 202-208 [PMID: 17279045 DOI: 10.1097/01.blo.0000238796.59596.1f]

20	Leclercq S, Benoit JY, de Rosa JP, Euvrard P, Leteurtre C, Girardin P. Results of the Evora dual-mobility socket after a minimum follow-up of five years. Rev Chir Orthop Reparatrice Appar Mot 2008; 94: e17-e22 [PMID: 19070709 DOI: 10.1016/j.rco.2007.10.015]

21	Vielpeau C, Lebel B, Ardouin L, Burdin G, Lautridou C. The dual mobility socket concept: experience with 668 cases. Int Orthop 2011; 35: 225-230 [PMID: 21184223 DOI: 10.1007/s00264-010-1156-8]

22	Combes A, Migaud H, Girard J, Duhamel A, Fessy MH. Low rate of dislocation of dual-mobility cups in primary total hip arthroplasty. Clin Orthop Relat Res 2013; 471: 3891-3900 [PMID: 23516032 DOI: 10.1007/s11999-013-2929-3]

23	Farizon F, de Lavison R, Azoulai JJ, Bousquet G. Results with a cementless alumina-coated cup with dual mobility. A twelve-year follow-up study. Int Orthop 1998; 22: 219-224 [PMID: 9795807 DOI: 10.1007/s002640050246]

24	Lautridou C, Lebel B, Burdin G, Vielpeau C. [Survival of the cementless Bousquet dual mobility cup: Minimum 15-year follow-up of 437 total hip arthroplasties]. Rev Chir Orthop Reparatrice Appar Mot 2008; 94: 731-739 [PMID: 19070715 DOI: 10.1016/j.rco.2008.06.001]

25	Boyer B, Philippot R, Geringer J, Farizon F. Primary total hip arthroplasty with dual mobility socket to prevent dislocation: a 22-year follow-up of 240 hips. Int Orthop 2012; 36: 511-518 [PMID: 21698430 DOI: 10.1007/s00264-011-1289-4]

26	Lewinnek GE, Lewis JL, Tarr R, Compere CL, Zimmerman JR. Dislocations after total hip-replacement arthroplasties. J Bone Joint Surg Am 1978; 60: 217-220 [PMID: 641088]

27	Gioe TJ. Dislocation following revision total hip arthroplasty. Am J Orthop (Belle Mead NJ) 2002; 31: 225-227 [PMID: 12008856]

28	Alberton GM, High WA, Morrey BF. Dislocation after revision total hip arthroplasty : an analysis of risk factors and treatment options. J Bone Joint Surg Am 2002; 84-A: 1788-1792 [PMID: 12377909]

29	Ballard WT, Callaghan JJ, Johnston RC. Revision of total hip arthroplasty in octogenarians. J Bone Joint Surg Am 1995; 77: 585-589 [PMID: 7713976]

30	Philippot R, Adam P, Reckhaus M, Delangle F, Verdot F-, Curvale G, Farizon F. Prevention of dislocation in total hip revision surgery using a dual mobility design. Orthop Traumatol Surg Res 2009; 95: 407-413 [PMID: 19656750 DOI: 10.1016/j.otsr.2009.04.016]

31	Hamadouche M, Biau DJ, Huten D, Musset T, Gaucher F. The use of a cemented dual mobility socket to treat recurrent dislocation. Clin Orthop Relat Res 2010; 468: 3248-3254 [PMID: 20532718 DOI: 10.1007/s11999-010-1404-7]

32	Blomfeldt R, Törnkvist H, Ponzer S, Söderqvist A, Tidermark J. Comparison of internal fixation with total hip replacement for displaced femoral neck fractures. Randomized, controlled trial performed at four years. J Bone Joint Surg Am 2005; 87: 1680-1688 [PMID: 16085605 DOI: 10.2106/JBJS.D.02655]

33	Wang J, Jiang B, Marshall RJ, Zhang P. Arthroplasty or internal fixation for displaced femoral neck fractures: which is the optimal alternative for elderly patients? A meta-analysis. Int Orthop 2009; 33: 1179-1187 [PMID: 19343344 DOI: 10.1007/s00264-009-0763-8]

34	Iorio R, Healy WL, Lemos DW, Appleby D, Lucchesi CA, Saleh KJ. Displaced femoral neck fractures in the elderly: outcomes and cost effectiveness. Clin Orthop Relat Res 2001; (383): 229-242 [PMID: 11210960]

35	Adam P, Philippe R, Ehlinger M, Roche O, Bonnomet F, Molé D, Fessy MH. Dual mobility cups hip arthroplasty as a treatment for displaced fracture of the femoral neck in the elderly. A prospective, systematic, multicenter study with specific focus on postoperative dislocation. Orthop Traumatol Surg Res 2012; 98: 296-300 [PMID: 22463868 DOI: 10.1016/j.otsr.2012.01.005]

36	Tarasevicius S, Busevicius M, Robertsson O, Wingstrand H. Dual mobility cup reduces dislocation rate after arthroplasty for femoral neck fracture. BMC Musculoskelet Disord 2010; 11: 175 [PMID: 20691056 DOI: 10.1186/1471-2474-11-175]

37	Philippeau JM, Durand JM, Carret JP, Leclercq S, Waast D, Gouin F. Dual mobility design use in preventing total hip replacement dislocation following tumor resection. Orthop Traumatol Surg Res 2010; 96: 2-8 [PMID: 20170850 DOI: 10.1016/j.rcot.2009.12.011]

38	Blake SM, Kitson J, Howell JR, Gie GA, Cox PJ. Constrained total hip arthroplasty in a paediatric patient with cerebral palsy and painful dislocation of the hip. A case report. J Bone Joint Surg Br 2006; 88: 655-657 [PMID: 16645115 DOI: 10.1302/0301-620X.88B5.17206]

39	Buly RL, Huo M, Root L, Binzer T, Wilson PD. Total hip arthroplasty in cerebral palsy. Long-term follow-up results. Clin Orthop Relat Res 1993; (296): 148-153 [PMID: 8222418]

40	Raphael BS, Dines JS, Akerman M, Root L. Long-term followup of total hip arthroplasty in patients with cerebral palsy. Clin Orthop Relat Res 2010; 468: 1845-1854 [PMID: 19924492 DOI: 10.1007/s11999-009-1167-1]

41	Root L, Goss JR, Mendes J. The treatment of the painful hip in cerebral palsy by total hip replacement or hip arthrodesis. J Bone Joint Surg Am 1986; 68: 590-598 [PMID: 3957984]

42	Schroeder K, Hauck C, Wiedenhöfer B, Braatz F, Aldinger PR. Long-term results of hip arthroplasty in ambulatory patients with cerebral palsy. Int Orthop 2010; 34: 335-339 [PMID: 19387644 DOI: 10.1007/s00264-009-0771-8]

43	Weber M, Cabanela ME. Total hip arthroplasty in patients with cerebral palsy. Orthopedics 1999; 22: 425-427 [PMID: 10220058]

44	Sanders RJ, Swierstra BA, Goosen JH. The use of a dual-mobility concept in total hip arthroplasty patients with spastic disorders: no dislocations in a series of ten cases at midterm follow-up. Arch Orthop Trauma Surg 2013; 133: 1011-1016 [PMID: 23632783 DOI: 10.1007/s00402-013-1759-9]

45	Schörle CM, Fuchs G, Manolikakis G. [Total hip arthroplasty in cerebral palsy]. Orthopade 2006; 35: 823-833 [PMID: 16816962 DOI: 10.1007/s00132-006-0988-9]

46	Aubriot JH, Lesimple P, Leclercq S. [Study of Bousquet’s non-cemented acetabular implant in 100 hybrid total hip prostheses (Charnley type cemented femoral component). Average 5-year follow-up]. Acta Orthop Belg 1993; 59 Suppl 1: 267-271 [PMID: 8116409]

47	Philippot R, Adam P, Farizon F, Fessy MH, Bousquet G. [Survival of cementless dual mobility sockets: ten-year follow-up]. Rev Chir Orthop Reparatrice Appar Mot 2006; 92: 326-331 [PMID: 16948459]

48	Philippot R, Farizon F, Camilleri JP, Boyer B, Derhi G, Bonnan J, Fessy MH, Lecuire F. Survival of cementless dual mobility socket with a mean 17 years follow-up. Rev Chir Orthop Reparatrice Appar Mot 2008; 94: e23-e27 [PMID: 19070710 DOI: 10.1016/j.rco.2007.10.013]

49	Berry DJ. Unstable total hip arthroplasty: detailed overview. Instr Course Lect 2001; 50: 265-274 [PMID: 11372323]

50	Sanchez-Sotelo J, Berry DJ. Epidemiology of instability after total hip replacement. Orthop Clin North Am 2001; 32: 543-552, vii [PMID: 11689368]

51	Jolles BM, Zangger P, Leyvraz PF. Factors predisposing to dislocation after primary total hip arthroplasty: a multivariate analysis. J Arthroplasty 2002; 17: 282-288 [PMID: 11938502]

52	Mahoney CR, Pellicci PM. Complications in primary total hip arthroplasty: avoidance and management of dislocations. Instr Course Lect 2003; 52: 247-255 [PMID: 12690852]

53	Prokopetz JJ, Losina E, Bliss RL, Wright J, Baron JA, Katz JN. Risk factors for revision of primary total hip arthroplasty: a systematic review. BMC Musculoskelet Disord 2012; 13: 251 [PMID: 23241396 DOI: 10.1186/1471-2474-13-251]

54	Bouchet R, Mercier N, Saragaglia D. Posterior approach and dislocation rate: a 213 total hip replacements case-control study comparing the dual mobility cup with a conventional 28-mm metal head/polyethylene prosthesis. Orthop Traumatol Surg Res 2011; 97: 2-7 [PMID: 21177149 DOI: 10.1016/j.otsr.2010.07.008]

55	Mohammed R, Cnudde P. Severe metallosis owing to intraprosthetic dislocation in a failed dual-mobility cup primary total hip arthroplasty. J Arthroplasty 2012; 27: 493.e1-493.e3 [PMID: 21752581 DOI: 10.1016/j.arth.2010.11.019]

56	Hamadouche M, Arnould H, Bouxin B. Is a cementless dual mobility socket in primary THA a reasonable option? Clin Orthop Relat Res 2012; 470: 3048-3053 [PMID: 22639330 DOI: 10.1007/s11999-012-2395-3]

57	Adam P, Farizon F, Fessy MH. [Dual articulation retentive acetabular liners and wear: surface analysis of 40 retrieved polyethylene implants]. Rev Chir Orthop Reparatrice Appar Mot 2005; 91: 627-636 [PMID: 16327667]

58	Grazioli A, Ek ET, Rüdiger HA. Biomechanical concept and clinical outcome of dual mobility cups. Int Orthop 2012; 36: 2411-2418 [PMID: 23073926 DOI: 10.1007/s00264-012-1678-3]

59	Digas G, Kärrholm J, Thanner J, Herberts P. 5-year experience of highly cross-linked polyethylene in cemented and uncemented sockets: two randomized studies using radiostereometric analysis. Acta Orthop 2007; 78: 746-754 [PMID: 18236180 DOI: 10.1080/17453670710014518]

60	Bauchu P, Bonnard O, Cyprès A, Fiquet A, Girardin P, Noyer D. The dual-mobility POLARCUP: first results from a multicenter study. Orthopedics 2008; 31 [PMID: 19298026]

61	Hailer NP, Weiss RJ, Stark A, Kärrholm J. Dual-mobility cups for revision due to instability are associated with a low rate of re-revisions due to dislocation: 228 patients from the Swedish Hip Arthroplasty Register. Acta Orthop 2012; 83: 566-571 [PMID: 23116439 DOI: 10.3109/17453674.2012.742395]

P- Reviewers: Buttaro MA, Fisher DA    S- Editor: Wen LL    L- Editor: A    E- Editor: Lu YJ

FIGURE LEGENDS

Figure 1  Standard cup vs dual mobility cup. Standard metal-on-polyethylene implants (A) include one articulation between the femoral head and the acetabular liner (dashed line). A dual mobility cup (B) consists of two distinct articulations, one between the femoral head and the liner, and another one between the liner and the shell. This configuration allows for greater range of motion before impingement of the femoral neck occurs (C and D, angle  > angle ).

Figure 2  Bubble sign. AP pelvis radiograph of a patient with acute onset of left hip pain and limp. On the left, eccentric position of the femoral head within the dual mobility cup can be noted. Careful scrutiny reveals a circular radiolucent area superior to the acetabular component (arrows), which represents the dislocated polyethylene liner (“bubble sign”).

# Text boxes

Table 1  Main published results of dual mobility cups in total hip arthroplasty

| Ref | N. of hips | Indication | Mean FU | Implant design (cup) | Head size (mm) | Intraprosthetic dislocation (%) | Dislocation rate (%) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Boyer et al[25], 2012 | 240 | Primary THA | 22 yr | Novae®1 | 22.2 | 4.1 | 0 |
| Farizon et al[23], 1998 | 135 | Primary THA | 12 yr | Novae®1 | 22.2 | 2 | 0 |
| Lautridou et al[24], 2008 | 437 | Primary THA | 16.5 yr | Novae-1®1 | 22.2 | 0.7 | 1.1 |
| Philippot et al[47], 2006 | 106 | Primary THA | 10 yr | Novae-1®1 | 22.2 | 1.9 | 0 |
| Philippot et al[12], 2009 | 384 | Primary THA | 15.3 yr | Novae-1®1 | 22.2 | 3.6 | 0 |
| Philippot et al[48], 2008 | 438 | Primary THA | 17 yr | Novae-1®1 | 22.2 | 5.2 | 0 |
| Guyen et al[19], 2007 | 167 | Primary THA | 3 yr | Saturne®2 | n/a | 0 | 0 |
| Leclercq et al[20], 2008 | 200 | Primary THA | 6 yr | Evora®3 | 22.2 (n = 175) 26 (n = 18) 28 (n = 7) | 0 | 0 |
| Hamadouche et al[56], 2012 | 168 | Primary THA | 6 yr | Tregor®4 | 22.2 | 2.4 | 0 |
| Vielpeau et al[21], 2011 | 437 (Group A) 231 (Group B) | Primary THA | 16.5 yr 5.2 yr | Original Bousquet Novae-E®1 | 22.2 | 0.7 0 | 0 0 |
| Bouchet et al[54], 2011 | 105 | Primary THA | 2.3 yr | Novae®1, Statfit®5, Avantage®6, Gyros®7 | 28 | n/a | 0 |
| Bauchu et al[60], 2008 | 150 | Primary THA | 6.2 yr | Polarcup®8 3rd gen | n/a | 0 | 0 |
| Combes et al[22], 2013 | 2480 | Primary THA | 7 yr | Novae®1, Avantage®6, Collegia®9, EOL®10, Gyros®7, Tregor®4, Polarcup®8, Saturne®2, Evora®3 | 28 (n = 1484) 22 (n = 956) | 0.1 0.6 | 0.7 0.5 |
| Tarasevicius et al[36], 2010 | 42 | Neck Fractures | 1 yr | Avantage®6 | 28 | n/a | 0 |
| Adam et al[35], 2012 | 214 | Neck Fractures | 3-9 mo | Saturne®2 | 28 (n = 182) 22.2 (n = 32) | 0 | 1.4 |
| Sanders et al[44], 2013 | 10 | Spastic disorders | 3.2 yr | Avantage®6 | n/a | 0 | 0 |
| Philippeau et al[37], 2010 | 71 | Tumor resection | 3.3 yr | Avantage®6, Saturn®2, Novae®1, other | n/a | n/a | 9.8 |
| Langlais et al[17], 2008 | 85 | Revision THA | 3.2 yr | Tregor®4 | 22 | n/a | 1.1 |
| Leiber-Wackenheim et al[14], 2011 | 59 | Revision THA | 8 yr | Novae-1®1 Novae-E®1 | 28 | 0 | 1.7 |
| Hamadouche et al[31], 2010 | 51 | Revision THA | 4.3 yr | Tregor®4 | 22.2 | 2 | 2 |
| Guyen et al[16], 2009 | 54 | Revision THA | 3.9 yr | Saturne®2 | n/a | 3.7 | 1.8 |
| Hailer et al[61], 2012 | 228 | Revision THA | 2 yr | Avantage®6 | n/a | n/a | 2 |
| Philippot et al[30], 2009 | 163 | Revision THA | 5 yr | Novae®1 | 22.2 | 0 | 3.7 |

1Serf, Décines, France; 2Amplitude, Valence, France; 3Science et Médecine, Créteil, France; 4Aston, St Etienne, France; 5Zimmer, Etupes, France; 6Biomet, Valence, France; 7DePuy, St Priest, France; 8Smith and Nephew Orthopaedics AG, Rotkreuz, Switzerland; 9Cremascoli-Wright, Paris, France; 10Norton-Ceramconcept, Paris, France. n/a: Not available.